16 Channel High Voltage Board
Hardware Design and Test Document
Revision History

	Date
	Revision
	Author
	Comment

	July 15, 2003
	New
	Fred Keske
	Initial Release

	August 8, 2003
	A
	Fred Keske
	Cleaned-up test procedures

	August 22, 2003
	B
	Fred Keske
	Added Table 3

	November 7, 2003
	C
	Fred Keske
	Fixed typo in 2.1.2, item 2

	
	
	
	



41
High Voltage Functional Overview


41.1
Functional Requirements


41.2
Electrical Characteristics


41.2.1
Input Power


41.2.2
Power Supply


41.2.3
Bus Control


51.2.4
Digital to Analog Converter


51.2.5
High Voltage Amplifiers


71.3
Board Connections


71.3.1
Edge connectors


71.3.2
Headers


81.4
Channel Sequence


91.4.1
Board to DM Connections


92
Test Procedure


102.1
High Voltage Amplifier Channel Low Voltage Bench Test


102.1.1
Test Requirements


102.1.2
Procedure


112.2
CPLD Programming


112.2.1
Requirements


112.2.2
Procedure


122.3
DAC Address Test


122.3.1
Requirements


132.3.2
Procedure


142.4
High Voltage Amplifier Test


142.4.1
Requirements


142.4.2
Procedure


17Appendix A: Schematics




4Figure 1


7Figure 4


8Figure 5


8Figure 6


10Figure 7


17Figure 8


17Figure 9


18Figure 10


18Figure 11


19Figure 12


19Figure 13




1 High Voltage Functional Overview

The High Voltage Board is comprised of 16 channels that can swing +/- 450 volts that is used to drive piezo material as part of an Adaptive Optics System.  Each channel is individually addressable through one of two Digital to Analog Converters (DAC) on board.  Dipswitches set the upper three address bits for the board.
1.1 Functional Requirements

The following is a list of guidelines and requirements for the High Voltage Board design:
· Load = 9nF (mirror segment) + 300pF (cabling) = 9.3nF
· Output voltage range = ±450 volts

· 250 volt maximum swing for an update period (T0)

· Update period (T0) = 1/10th of 2kHz = 50μs
· Slew rate = 250V/50μS = 5V/μs
1.2 Electrical Characteristics

A block diagram showing the implementation of the High Voltage Board is shown below in Figure 1.


[image: image1.emf]Bus

Interface

VME

Backplane

Interface

CPLD

(x2) 8 channel

13 Bit DAC

x16

High Voltage

Connector

Interface

16 Channels

High Voltage

Amplifiers

Power

Supply

+5V -5V


Figure 1
1.2.1 Input Power

The High Voltage Board is designed to operate from +5VDC ±5%, 21mA typical operating current.
1.2.2 Power Supply

The power supply block provides the -5 volts to the DACs.  This function is performed with a MAX 764 which is an inverting switching regulator using a 5 volt input.  The part requires a maximum of supply current of 120uA and will provide a minimum output current of 150mA @ -5 volts.
1.2.3 Bus Control
Control of the interfaces to the VME backplane and to the Digital to Analog Converter (DAC) is with a Cypress CPLD (CY37256).  This part controls the bus transceivers (74ACTQ245) which provides the address (ADDR[0..7]), data (D[0..15]), and control (CTRL[0..15])to the transceivers which in turn controls the data from the VME backplane.  The CPLD also provides the data (DO[0..12]) and control to the two DACs.
Shown below is the board address range select:
	Switch
	Address Range

	5
	6
	7
	

	ON
	ON
	ON
	0-11

	OFF
	ON
	ON
	12-23

	ON
	OFF
	ON
	24-35

	OFF
	OFF
	ON
	36-47

	ON
	ON
	OFF
	48-59

	OFF
	ON
	OFF
	60-71

	ON
	OFF
	OFF
	72-83

	OFF
	OFF
	OFF
	84-95


Table 1
Note: switches 1 through 4 and 8 are not used and should remain in the ‘ON’ position.

1.2.4 Digital to Analog Converter
The Digital to Analog (DAC) function is performed with two 13-bit DACs (MAX547).  Each of these DACs provides eight channels to the high voltage amplifiers.  The digital data interface is offset binary in which 0 to 8191 is scaled to the -4.5V to +4.5V output respectively.
1.2.5 High Voltage Amplifiers
A high voltage amplifier (PA95) that can swing from -450V to +450V is used to drive the piezo material.  The gain is set to 99 to amplify the -4.5V to +4.5V input signal originating from the DAC.
1.2.5.1 Current Limit

The current limit is set to 45.7mA with the three-5.11Ω resistors in series.  The limit value is derived from using a 1.7A Sorensen power supply driving the higher capacitive 24 outer actuators with a 0.6 de-rating factor.
1.2.5.2 Loop Stability

C47, R27, and C48 have been added to ensure loop stability.  In addition, three 3.3Ω resistors in series have been added to the output of the amplifier to isolate the load somewhat.
As can be seen by the Bode Plot in figure 2, we have a good rate of closure of 20dB/decade.  The phase shift plot shown in figure 3 indicates a good phase margin of 45°.
	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	[image: image16.png]Gain, dB

140

Bode Plot

120

100

il

Aol

60

0

Eil

20

X

1.E+00

1E01

16402

1EW3 1EsD4

Frequency, Hz

16405

1E408

1E07

16408

Aol/Cload

- - 1/Beta

— - —Ad

- - Signal at
Cload
—x—Fel





	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	[image: image17.png]Phase Shift, Degrees

=]
\
\

-180

Phase Shift

45

o I

I

1EWD 1ED!

16402

1E03 1E0s 18405
Frequency, Hz

1E408

1E07

16408

Open Loop

Closed 1
Closed'10

Closed'100
Closed*1000

—x—Fel





	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	

	
	
	
	
	
	
	
	
	


1.2.5.3 Slew Rate
The slew rate is based on the maximum current the amplifier is able to source, which in this case is determined by the current limit resistor and the total capacitance of the load.  45.6mA/9.3nF = 4.9V/μs.

1.3 Board Connections

1.3.1 Edge connectors

1.3.1.1 Backplane Connector

The Backplane connector labeled P1 is used to connect the High Voltage Board to the VME backplane.  The signals available on the connector are:  R_CTRL[0..6,8,10,11,13] used for control, RB_D[0..13] used for data, and R_ADDR[0..7] used for address.  
1.3.1.2 High Voltage Connector

The High Voltage Connector is used to connect the High Voltage Board to the piezo element material.  Pin 1 is located on the right side of the connector.  The connector encompasses the original 12 channels on the top and bottom rows along with the addition of the four new channels on the middle row.  See sheet 1 of the schematic for connection details and also the .
1.3.2 Headers

1.3.2.1 JTAG Header

The 10-pin JTAG header designated as JP1 is used to program the CPLD.  

1.3.2.2 Test/Debug Header
JP2 is a 10-pin header that is used as a test and debug connector for the CPLD which is not used in normal operation.  Pin 1 is located at the bottom left corner of the header.

1.3.2.3 VCC Header
JP3 is used to connect +5 volts to the board.  Pin 1 is located at the top of the header.  See Pin-out below.

[image: image2]
Figure 4
1.3.2.4 High Voltage Header
JP4 is used to connect ±350 volts to the board.  Pin 1 is located at the top of the header.  See Pin-out below.


[image: image3]
Figure 5
1.3.2.5 Channel Header

There is a 3-pin header for each of the 16 channels.  The header is designated as JP5 with an appended alpha character specifying the respective channel.  Pin 1 is the Analog ground, pin 2 is to monitor the output of the DAC, and pin 3 is to drive the input of the power amplifier with a ±4.5 volt signal.  Pin 1 is located at the left side of the header.  See Pin-out below.

[image: image4]
Figure 6
1.4 Channel Sequence

The High Voltage Board was designed with an additional four channels over the previous version of the board for a total of 16 channels that are now available to the backplane.  The table below lists the board channel designation, the on-board circuit components designated with an alpha subscript, and the backplane channel designation, which also corresponds to the present software.

	Board

Channel

Sequence
	Channel

Component

Subscript
	Backplane

Channel

Sequence

	1
	A
	0

	2
	B
	1

	3
	C
	2

	4
	D
	3

	5
	E
	4

	6
	F
	5

	7
	G
	N/A

	8
	H
	N/A

	9
	I
	6

	10
	J
	7

	11
	K
	8

	12
	L
	9

	13
	M
	10

	14
	N
	11

	15
	O
	N/A

	16
	P
	N/A


Table 2
1.4.1 Board to DM Connections

The table below lists the board channel pin to MS27467T19 connector pin to the pad number on the DM.
	Backplane Channel
	P2A Board Pin
	MS27467T19 Pin
	System Channel
	DM Pad

	0
	Z26
	46
	HV36
	17

	1
	D24
	47
	HV37
	18

	2
	Z22
	48
	HV38
	34

	3
	D20
	49
	HV39
	35

	4
	Z18
	51
	HV40
	55

	5
	D16
	52
	HV41
	56

	6
	Z14
	53
	HV42
	81

	7
	D12
	54
	HV43
	82

	8
	Z10
	56
	HV44
	83

	9
	D8
	57
	HV45
	84

	10
	Z6
	58
	HV46
	85

	11
	D4
	59
	HV47
	86


Table 3
2 Test Procedure
Shown in the figure below is the top side of the board.  The headers that are used in testing with its associated pin 1 are designated.
[image: image5.jpg]



Figure 7
2.1 High Voltage Amplifier Channel Low Voltage Bench Test

2.1.1 Test Requirements
· Function Generator

· Oscilloscope, 2 channel

· DC Power Supply with 15V and -15V outputs

· Cables to connect power supply and Function Generator to board
2.1.2 Procedure
1. Set up the function generator to create a square wave with a 4ms period and amplitude +/-100mV.

2. On JP4, connect -15V to pin 1, GND to pin 2, and +15V to pin 3.

3. Turn on the power supply.

4. On JP5A, connect output of the Function Generator to pin 3 and GND from the Function Generator to pin 1.

5. Also, connect the output and GND from the Function Generator to channel A of the oscilloscope.  Set the channel to 50mVDC and 1ms.
6. On P2, connect the second channel of the oscilloscope to pin z26 on the high voltage connector using an appropriate breakout connecter.  Set the channel to 5VDC
7. Check to see that:

a. The waveform on channel B of the oscilloscope swings from -10V to +10V.

b. The waveform on channel B of the oscilloscope is approximately 180° out of phase when compared to channel A of the oscilloscope.

c. The transitions are fast and not rounded at the top.

8. If any of the above criteria are not met, the channel fails the test.  Keep track of board serial number and channel for debugging.

9. Repeat steps 5-8 using corresponding connector in step 5 for each of the channels.

10. If all channels pass testing, board has passed High Voltage Amplifier Channel Low Voltage Bench Test.
2.2 CPLD Programming

2.2.1 Requirements
· Cypress ISR programming software

· Cypress UltraISR programming cable
· 5V power supply

· hva.jed programming file
2.2.2 Procedure
1. Plug Cypress UltraISR programming cable into the parallel port of the PC and start the Cypress ISR programming software.

2. Select ‘New’ from the ‘File’ menu

[image: image6.png][=[ofx]

Fle Edt Operations Options Help

NEEIEEEEL]

Device Operation Filename Browse

[New Device Chain ]

e
5746 i Flrne. | TN

‘Watking Ditectory: Browse.

Feady 7





3. Type ‘1’ in the ‘Number of devices in JTAG Chain’ text box.

4. Type any filename in the ‘JTAG Chain Filename’ text box.

5. Browse and select or type in the directory that you would like to save the programming session.

6. Press ‘Ok’

[image: image7.png]Fle Edt Operations Options Help

NEEIEEEEE]

Device

Operation

Filename

CY37256P 160

[Program & Verf





7. In the ‘Devices’ box select ‘CY37256P160’.

8. In the ‘Operation’ box, select ‘Program & Verify’.

9. Use the ‘Browse’ button to locate and set the path and filename in the ‘filename’ text box to ‘G:\\Electronics\arc_2001_1_17\vhdl\HVA\hva.jed.
10. Press the [image: image8.bmp] button to compose the programming file.

11. On JP3, connect power supply GND to pins 1 and 2, +5V to pin 3.

12. Connect Cypress UltraISR programming cable to header JP1.  Connector should be polarized, but if not ensure that pin 1 of cable connects to pin 1 of the header.

13. Turn on the power supply.

14. Press the [image: image9.bmp] button to program the CPLD.  Programming may take several second to complete

15. Check the log that is displayed to see that the CPLD programmed and verified successfully.

16. If programming or verify was not successful:

a. Verify that the path and filename are correct in the ‘filename’ textbox.  If not, repeat from step 9.

b. Check to see that power and ground are connected correctly and that supply voltage is set to 5VDC.  If not, repeat form step 11.

c. Check to see that part is installed and soldered correctly to board.  If not, correct problem and repeat from step 11.

17. Turn off power supply and disconnect cables.

2.3 DAC Address Test

2.3.1 Requirements
· 2 Multifunction Boards

· 1 Chassis

· 2 PCs with Linux OS

· 2 FC-FC Multimode Fiber Optic Cables

· Oscilloscope

2.3.2 Procedure
1. If not already done so, log into AOUIM by typing ‘ao’ and hitting return.

2. Type the password, ‘wai!mea’ and hit return.
3. Start xwindows by typing ‘startx’
4. On AOUIM, open 3 Xterm windows

5. In first Xterm window type ‘rlogin –l ao aoicm’ and hit return.
6. Type the password, ‘wai!mea’ and hit return.

7. Change to the ‘manoa85/dio32’ directory.

8. Type ‘su aroot’ and hit return.

9. Type the password, ‘wai!mea’ and hit return.

10. Type ‘start_dio’ and hit return.  This should start dio32.

11. In second Xterm window type ‘rlogin –l ao aoicm’ and hit return.

12. Type the password, ‘wai!mea’, and hit return.

13. Change to the ‘manoa85/dio32’ directory.

14. Type ‘cat /dev/rtf0’ and hit return.

15. In the third Xterm window, type ‘rlogin –l ao aoicm’ and hit return.

16. Type the password, ‘wai!mea’ and hit return.

17. Change to the ‘manoa85/dio32’ directory.

18. Type ‘cat > /dev/rtf1’ and hit return.

19. You should now be able to see the commands that you type in Xterm 3 show up in Xterm window 2.  To test this, type ‘test’ in Xterm 3.  If you do not see ‘test’ in Xterm 2:
a. Hit ‘ctrl-c’ in Xterm 3.

b. In Xterm 3, type ‘exit’ and hit return.

c. Hit ‘ctrl-c’ in Xterm 2.

d. In Xterm 2, type ‘exit’ and hit return.

e. In Xterm 1, type ‘stop_dio’ and hit return.

f. Repeat from step 7.

20. Make sure that all dip switches on SW1 are in the ‘ON’ position

21. Plug board that into the chassis.

22. On JP5A, connect oscilloscope GND to pin 1 and oscilloscope channel A to pin 2.  Set the voltage scale to 5VDC with the time base at 20ms/division.
23. Turn on power by flipping switch on front of chassis

24. In window #3, type ‘a i’ and hit return.

25. Type ‘f m 1’ and hit return to set the test pattern to saw tooth.

26. Type ‘f p 10000’ and hit return to set the period of the pulse to 80 ms.

27. Type ‘f a 0’ and hit return to direct commands to HV0.

28. Type ‘f b’ and hit return to start pattern generation.

29. Check the oscilloscope to see a saw tooth pattern is being generated.  If output stays constant on oscilloscope:
a. Turn off chassis power by flipping switch on front of chassis.

b. Wait 5-10 seconds and turn power back on.

c. Check if pattern is being generated.

d. If still no output, move oscilloscope probe to corresponding connector on channel A and type ‘f a 1’ and hit return, then repeat steps 21through 24.
e. If not, repeat the above steps a through c four times.

f. If still no output, there may be a problem with one of the DACs.  Turn off chassis and remove board.  Check for assembly errors around U13 and U14.

30. Move switch #5 on SW1 to the ‘OFF’ position.

31. Type ‘f a 12’ and hit return.

32. Type ‘f b’ and hit return.

33. Check oscilloscope to make sure that the saw tooth pattern is being generated.

34. Type ‘f a 1’ and hit return.

35. Check to make sure that the saw tooth pattern is NOT being generated.

36. Type ‘f a 24’ and hit return.

37. Check to make sure that the saw tooth pattern is NOT being generated.

38. Move switch #5 on SW1 to the ‘ON’ position and switch #6 on SW1 to the ‘OFF’ position.

39. Check oscilloscope to make sure that the saw tooth pattern is being generated.

40. Type ‘f a 0’ and hit return.

41. Check to make sure that the saw tooth pattern is NOT being generated.

42. Type ‘f a 12’ and hit return.

43. Check to make sure that the saw tooth pattern is NOT being generated.

44. Turn off power.

2.4 High Voltage Amplifier Test

2.4.1 Requirements
· Multifunction Boards

· 1 Chassis

· 2 PCs with Linux OS

· 2 FC-FC Multimode Fiber Optic Cables

· High Voltage Oscilloscope, 2 channel
· High Voltage Cable

· Piezo Material with two leads connected to it

NOTE: You will be working with +/-450 volts with this test.  Make sure to follow all the steps in the order that they are written.  Failure to do so could result in serious electrical shock.

2.4.2 Procedure

2.4.2.1 High Voltage Test
1. If not already done so, perform steps 1 through 21 from section 2.3.2.1 to enable the test program.
2. On JP5A, connect oscilloscope GND to pin 1 and oscilloscope channel A to pin 2.  Set the voltage scale to 5VDC with the time base at 20ms/division.
3. Turn on power by flipping switch on front of chassis

4. In window #3, type ‘a i’ and hit return.

5. Type ‘f m 1’ and hit return to set the test pattern to saw tooth.

6. Type ‘f p 10000’ and hit return to set the period of the pulse to 80ms.

7. Type ‘f a 0’ and hit return to direct commands to HV0.

8. Type ‘f b’ and hit return to start pattern generation.

9. Check the oscilloscope to see a saw tooth pattern is being generated.  If output stays constant on the oscilloscope:
a. Turn off chassis power by flipping the switch on the front of the chassis.

b. Wait 5-10 seconds and turn power back on.

c. Check if pattern is being generated.

d. If not, repeat the above steps a through c four times.

10. Install the high voltage connector assembly in the J2 connector on the PCB.
11. Insert the 12 individual shrink-wrapped pins into the PCB side of the high voltage connector corresponding to the 12 channel pins inside the connector.

12. Connect the high voltage supply to JP4.

13. Turn on the high voltage power supply from the power strip.  Make sure the voltage is at 450 volts on both supplies.
14. Connect probe of channel B on the oscilloscope to the shrink-wrapped pin corresponding to high voltage channel 0. Set the voltage scale to 200VDC.
15. Check to see that the voltage from oscilloscope channel B roughly follows the saw tooth pattern on oscilloscope channel A and that the voltage is approximately +/-450V.
16. Turn off the high voltage power supply from the power strip
NOTE: This is a very important step as +/-450V is being carried on the high voltage cable.  If you do not turn off the high voltage supply before doing this step, there is a high risk for electrical shock.
17. Type ‘f a 6’ and hit return to direct commands to HV6.
18. Check that the channel being tested is NOT outputting the saw tooth pattern (unless HV6 is the channel being tested, which is on JP5I).  If output remains flat, channel has passed the test.

19. Move oscilloscope channel A probe to the corresponding JP5 connector on the next channel with the oscilloscope channel B probe to the corresponding socket on the high voltage cable.
20. Turn on the high voltage power supply from the power strip.
21. Type ‘f a 0’ and hit return to direct commands to HV0.
22. Check that the channel is NOT outputting the saw tooth pattern.  If output remains flat, channel has passed the test.

23. Type ‘f a x’ with ‘x’ being the corresponding channel number that is being tested, then hit return.
24. Check the oscilloscope to see a saw tooth pattern is being generated.
25. epeat steps 13 through 24 until all the channels have all been tested, noting the pass/fail status of the channels.R
26. Turn off all power supplies.

27. If not planning to test another board:

a. Press ‘ctrl-c’ in Xterm 3.

b. Type ‘exit’ in Xterm 3 and hit return.

c. Press ‘ctrl-c’ in Xterm 2.

d. Type ‘exit’ in Xterm 2 and hit return.

e. In Xterm 1 type ‘stop_dio’ and hit return.
f. Type ‘exit’ and hit return.

2.4.2.2 Filter response test

1. Setup the function generator by setting the waveform to a square wave, the frequency to 2kHz, and the voltage to 0.35Vpp.  Make sure the offset is set to zero.

2. Connect function generator probe to JP5A pin 3 and ground to pin 1.  Set the voltage scale to 200mVDC with the time base at 50us/division.
3. Turn on the cursor data function dt and the dV on the oscilloscope.

4. Make sure the piezo material is not connected to the high voltage connector assembly. Turn on the high voltage power supply from the power strip.  Adjust the voltage to 50 volts on both supplies.
5. Turn off the high voltage power supply from the power strip.

6. Connect the piezo material to the high voltage connector assembly.

7. Turn on the high voltage power supply from the power strip. 
8. Connect probe of channel B on the oscilloscope element corresponding to high voltage channel 0.  Set the voltage scale to 20VDC.

9. Place the vertical bar of cursor 1 on the oscilloscope on the edge of channel A and the horizontal bar of cursor 2 on the settled voltage of channel B.
10. Check to see that the dt value is less than 36us.

11. Move the horizontal bar of cursor 1 to the peak of channel B.
12. Check to see that the dV value is less than 7V.

13. Turn off the high voltage power supply from the power strip.

14. Move oscilloscope channel A probe to the corresponding JP5 connector on the next channel with the oscilloscope channel B probe to the corresponding element on the piezo material.

15. Turn on the high voltage power supply from the power strip.
16. Connect probe of channel B on the oscilloscope element corresponding to the next high voltage channel.  

17. epeat steps 9 through 16 until all the channels have all been tested, noting the pass/fail status of the channels.
R
18. Turn off the high voltage power supply from the power strip.
19. Remove board from chassis.
Appendix A: Schematics
[image: image10.emf]1

1

2

2

3

3

4

4

5

5

6

6

D D

C C

B B

A A

1

UH - Institute for Astonomy

640 N. Aohoku Place

Hilo, HI 96720

USA

6

External I/O

IRTF - AO A

6/3/2003 4:41:29 PM

C:\Documents and Settings\Fred keske\My Documents\Projects\High Voltage\Connectors.SCHDOC

Title

Size: Number:

Date:

File:

Revision:

Sheet of Time:

B

Engineer

Fred Keske

REPEAT(IN)REPEAT(HV_OUTPUT)

REPEAT(HVC,1,16)

High Voltage Channel.SCHDOC

C1

C1

C2

C2

C3

C3

C4

C4

C5

C5

C6

C6

C7

C7

C8

C8

C9

C9

C10

C10

C11

C11

C12

C12

C13

C13

C14

C14

C15

C15

C16

C16

C17

C17

C18

C18

C19

C19

C20

C20

C21

C21

C22

C22

C23

C23

C24

C24

C25

C25

C26

C26

C27

C27

C28

C28

C29

C29

C30

C30

C31

C31

C32

C32

B1

B1

B2

B2

B3

B3

B4

B4

B5

B5

B6

B6

B7

B7

B8

B8

B9

B9

B10

B10

B11

B11

B12

B12

B13

B13

B14

B14

B15

B15

B16

B16

B17

B17

B18

B18

B19

B19

B20

B20

B21

B21

B22

B22

B23

B23

B24

B24

B25

B25

B26

B26

B27

B27

B28

B28

B29

B29

B30

B30

B31

B31

B32

B32

A1

A1

A2

A2

A3

A3

A4

A4

A5

A5

A6

A6

A7

A7

A8

A8

A9

A9

A10

A10

A11

A11

A12

A12

A13

A13

A14

A14

A15

A15

A16

A16

A17

A17

A18

A18

A19

A19

A20

A20

A21

A21

A22

A22

A23

A23

A24

A24

A25

A25

A26

A26

A27

A27

A28

A28

A29

A29

A30

A30

A31

A31

A32

A32

A B C

DIN96

P1

RB_D5

RB_D10

R_ADDR7

R_ADDR6

R_ADDR5

R_ADDR4

R_ADDR3

R_ADDR2

R_ADDR1

RB_D13

RB_D12

RB_D11

RB_D9

RB_D8

RB_D7

RB_D6

RB_D4

RB_D3

RB_D2

RB_D1

RB_D0

VCC VCC VCC

d2

d2

d4

d4

d6

d6

d8

d8

d10

d10

d12

d12

d14

d14

d16

d16

d18

d18

d20

d20

d22

d22

d24

d24

d26

d26

d28

d28

d30

d30

d32

d32

TYPE F CONNECTOR

P2A

b2

b2

b4

b4

b6

b6

b8

b8

b10

b10

b12

b12

b14

b14

b16

b16

b18

b18

b20

b20

b22

b22

b24

b24

b26

b26

b28

b28

b30

b30

b32

b32

TYPE F CONNECTOR

P2B

z2

z2

z4

z4

z6

z6

z8

z8

z10

z10

z12

z12

z14

z14

z16

z16

z18

z18

z20

z20

z22

z22

z24

z24

z26

z26

z28

z28

z30

z30

z32

z32

TYPE F CONNECTOR

P2C

+HV

-HV

HV OUTPUT[1..16]

HV OUTPUT12

HV OUTPUT10

HV OUTPUT8

HV OUTPUT4

HV OUTPUT2

HV OUTPUT16

HV OUTPUT15

HV OUTPUT14

HV OUTPUT13

HV OUTPUT11

HV OUTPUT9

HV OUTPUT5

HV OUTPUT3

HV OUTPUT1

HV OUTPUT7

HV OUTPUT6

HV OUTPUT IN IN[1..16]

+HV

-HV

VCC

1

2

3

HEADER 3

JP4

1

2

3

4

5

HEADER 5

JP3

RB_D[0..13]

R_ADDR[0..7]

R_CTRL[0..13] VOUT[1..16]

Bus IO

Bus.SCHDOC; CPLD.SCHDOC; DAC.SCHDOC

R_ADDR[0..7]

RB_D[0..13]

R_CTRL[0..13]

VCC -5V

Power Supply

Power Supply.SCHDOC -5V VCC

R_CTRL0 R_CTRL1

R_CTRL2

R_CTRL3

R_CTRL4

R_CTRL5

R_CTRL6

R_CTRL8

R_CTRL10

R_CTRL11

R_CTRL13

0

1/8 W

R18

R_ADDR0

HV0

HV1

HV2

HV3

HV4

HV5

HV6

HV7

HV8

HV9

HV10

HV11


Figure 8
[image: image11.emf]1

1

2

2

3

3

4

4

5

5

6

6

D D

C C

B B

A A

2

UH - Institute for Astonomy

640 N. Aohoku Place

Hilo, HI 96720

USA

6

Bus Interface

IRTF - AO A

5/9/2003 4:41:01 PM

C:\Documents and Settings\Fred keske\My Documents\Projects\High Voltage\Bus.SCHDOC

Title

Size: Number:

Date:

File:

Revision:

Sheet of Time:

B

Engineer

Fred Keske

LAYOUT

PETER ONAKA

1 OF XX

CTRL8

CTRL9

CTRL10

CTRL11

CTRL12

CTRL13

CTRL14

CTRL15

CTRL0

CTRL1

CTRL2

CTRL3

CTRL4

CTRL5

CTRL6

CTRL7

G

19

DIR

1

A1

2

B1

18

A2

3

B2

17

A3

4

B3

16

A4

5

B4

15

A5

6

B5

14

A6

7

B6

13

A7

8

B7

12

A8

9

B8

11

74ACTQ245

U1

RB_D0

RB_D1

RB_D2

RB_D3

RB_D4

RB_D5

RB_D6

RB_D7

RB_D8

RB_D9

RB_D10

RB_D11

RB_D12

RB_D13

D0

D1

D2

D3

D4

D5

D6

D7

D8

D9

D10

D11

D12

D13

D14

D15

ADDR7

ADDR6

ADDR5

ADDR4

ADDR3

ADDR2

ADDR1

ADDR0

R_ADDR7

R_ADDR6

R_ADDR5

R_ADDR4

R_ADDR3

R_ADDR2

R_ADDR1

G

19

DIR

1

A1

2

B1

18

A2

3

B2

17

A3

4

B3

16

A4

5

B4

15

A5

6

B5

14

A6

7

B6

13

A7

8

B7

12

A8

9

B8

11

74ACTQ245

U4

G

19

DIR

1

A1

2

B1

18

A2

3

B2

17

A3

4

B3

16

A4

5

B4

15

A5

6

B5

14

A6

7

B6

13

A7

8

B7

12

A8

9

B8

11

74ACTQ245

U2

G

19

DIR

1

A1

2

B1

18

A2

3

B2

17

A3

4

B3

16

A4

5

B4

15

A5

6

B5

14

A6

7

B6

13

A7

8

B7

12

A8

9

B8

11

74ACTQ245

U3

G

19

DIR

1

A1

2

B1

18

A2

3

B2

17

A3

4

B3

16

A4

5

B4

15

A5

6

B5

14

A6

7

B6

13

A7

8

B7

12

A8

9

B8

11

74ACTQ245

U5

VCC

VCC

VCC VCC

VCC

49.9k

1/8 W

R1

49.9k

1/8 W

R32

RB_D[0..13] R_ADDR[0..7] R_CTRL[0..13]

0.1uF

50 V

C1

0.1uF

50 V

C2

0.1uF

50 V

C5

0.1uF

50 V

C4

0.1uF

50 V

C3

ADDR_ENABLE

ADDR_DIR

CTRL_H_ENABLE

CTRL_H_DIR

CTRL_L_ENABLE

CTRL_L_DIR

D[0..15]

ADDR[0..7]

CTRL[0..15] DATA_L_ENABLE

DATA_L_DIR

DATA_H_ENABLE

DATA_H_DIR

DATA_L_ENABLE

DATA_L_DIR

DATA_H_ENABLE

DATA_H_DIR

ADDR_ENABLE

ADDR_DIR

CTRL_H_ENABLE

CTRL_H_DIR

CTRL_L_ENABLE

CTRL_L_DIR

D[0..15] ADDR[0..7] CTRL[0..15]

RB_D[0..13]

D[0..15]

R_ADDR[0..7]

ADDR[0..7]

R_CTRL[0..13]

CTRL[0..15]

49.9k

1/8 W

R31

49.9k

1/8 W

R3

49.9k

1/8 W

R33

49.9k

1/8 W

R50

49.9k

1/8 W

R51

R_CTRL0

R_CTRL1

R_CTRL2

R_CTRL3

R_CTRL4

R_CTRL5

R_CTRL6

R_CTRL8

R_CTRL10

R_CTRL11

R_CTRL13

R_ADDR0


Figure 9
[image: image12.emf]1

1

2

2

3

3

4

4

5

5

6

6

D D

C C

B B

A A

3

UH - Institute for Astonomy

640 N. Aohoku Place

Hilo, HI 96720

USA

6

Bus Control

IRTF - AO A

5/12/2003 4:44:08 PM

C:\Documents and Settings\Fred keske\My Documents\Projects\High Voltage\CPLD.SCHDOC

Title

Size: Number:

Date:

File:

Revision:

Sheet of Time:

B

Engineer

Fred Keske

GND

1

I/0 16

2

I/0 17

3

I/0 18

4

I/0 19

5

TCLK

6

I/0 21

7

I/0 22

8

I/0 23

9

GND

10

I/0 24

11

I/0 25

12

I/0 26

13

I/0 27

14

I/0 28

15

I/0 29

16

I/0 30

17

I/0 31

18

CLK 0

19

VCCO

20

GND

21

CLK 1

22

I/0 32

23

I/0 33

24

I/0 34

25

I/0 35

26

I/0 36

27

I/0 37

28

I/0 38

29

I/0 39

30

GND

31

I/O 40

32

I/O 41

33

I/O 42

34

I/O 43

35

I/O 44

36

I/O 45

37

I/O 46

38

I/O 47

39

VCCO

40

GND

41

I/O 48

42

I/O 49

43

I/O 50

44

I/O 51

45

TMS

46

I/O 53

47

I/O 54

48

I/O 55

49

GND

50

I/O 56

51

I/O 57

52

I/O 58

53

I/O 59

54

I/O 60

55

I/O 61

56

I/O 62

57

I/O 63

58

I 2

59

VCCO

60

GND

61

VCC

62

I/O 64

63

I/O 65

64

I/O 66

65

I/O 67

66

I/O 68

67

I/O 69

68

I/O 70

69

I/O 71

70

GND

71

I/O 72

72

I/O 73

73

I/O 74

74

I/O 75

75

TDO

76

I/O 77

77

I/O 78

78

I/O 79

79

VCCO

80

GND

81

I/O 80

82

I/O 81

83

I/O 82

84

I/O 83

85

I/O 84

86

I/O 85

87

I/O 86

88

GND

90

I/O 88

91

I/O 89

92

I/O 90

93

I/O 91

94

I/O 92

95

I/O 93

96

I/O 94

97

I/O 95

98

I/O 87

89

CLK 2

99

VCCO

100

GND

101

CLK3

102

I/O 96

103

I/O 97

104

I/O 98

105

I/O 99

106

I/O 100

107

I/O 101

108

I/O 102

109

I/O 103

110

GND

111

I/O 104

112

I/O 105

113

I/O 106

114

I/O 107

115

TDI

116

I/O 109

117

I/O 110

118

I/O 111

119

VCCO

120

GND

121

I/O 112

122

I/O 113

123

I/O 114

124

I/O 115

125

I/O 116

126

I/O 117

127

I/O 118

128

I/O 119

129

GND

130

I/O 120

131

I/O 121

132

I/O 122

133

I/O 123

134

I/O 124

135

I/O 125

136

I/O 126

137

I/O 127

138

JTAG EN

139

VCC

140

GND

141

VCCO

142

I/O 0

143

I/O 1

144

I/O 2

145

I/O 3

146

I/O 4

147

I/O 5

148

I/O 6

149

I/O 7

150

GND

151

I/O 8

152

I/O 9

153

I/O 10

154

I/O 11

155

I/O 12

156

I/O 13

157

I/O 14

158

I/O 15

159

VCCO

160

CY37256

U6

RES0

1

RES1

2

RES2

3

RES3

4

COM0

5

RES4

6

RES5

7

RES6

8

RES7

9

COM1

10

4.7k

1/16 WR10

VCC

VCC

7.32k

1/8 W

R5

499k

1/8 W

R4

1k

1/8 W

R9

1k

1/8 W

R8

CTRL0

CTRL1

CTRL2

CTRL3

CTRL4

CTRL5

CTRL6

CTRL7

CTRL8

CTRL9

CTRL10

CTRL11

CTRL12

CTRL13

CTRL14

CTRL15

ADDR0

ADDR1

ADDR2

ADDR3

ADDR4

ADDR5

ADDR6

ADDR7

TMS

TMS

1

2

3

4

5

6

7

8

16

15

14

13

12

11

10

9

SW DIP-8

SW1

TDO

TDO

DAILD0

DAILD1

DAILD2

DAILD3

DAALD0

DAALD1

DAALD2

DAALD3

ADDRI0

ADDRI1

ADDRI2

TDI

DO0 DO1 DO2 DO3 DO4 DO5 DO6 DO7 DO8 DO9

DO10 DO11 DO12

D0 D1 D2 D3 D4 D5 D6 D7 D8 D9

D10 D11 D12 D13 D14 D15

TCK

TCK JTAGEN

VCC

VCC VCC VCC VCC

VCC VCC VCC

SHDN0

SHDN1

SHDN2

1k

1/8 W

R6

VCC

VCC

VCC

VCC

VCC

VCC

VCC

0.1uF

50 V

C8

0.1uF

50 V

C9

0.1uF

50 V

C10

0.1uF

50 V

C11

0.1uF

50 V

C17

0.1uF

50 V

C16

0.1uF

50 V

C15

0.1uF

50 V

C14

0.1uF

50 V

C6

0.1uF

50 V

C7

VCC

0.1uF

50 V

C13

1 2

3 4

5 6

7 8

9 10

HEADER5x2

JP1

NCOUTPUT

8

Oscillator

U8

LXH103ID

LED1

1k

1/8 W

R7

DO[0..12]

WR

CLR

ADDRI[0..2]

D[0..15]

CTRL_H_ENABLE

CTRL_H_DIR

CTRL_L_ENABLE

CTRL_L_DIR

DATA_L_DIR

DATA_L_ENABLE

CTRL[0..15]

ADDR[0..7]

DATA_H_DIR

DATA_H_ENABLE

ADDR_ENABLE

ADDR_DIR

DAALD[0..3]

DACS1

SHDN[0..2]

DAILD[0..3]

DACS0

WR

CLR

ADDRI[0..2]

DAALD[0..3]

DACS1

SHDN[0..2]

DAILD[0..3]

DACS0

DO[0..12]

ADDRI[0..2]

DAALD[0..3]

SHDN[0..2]

DAILD[0..3]

CTRL_H_ENABLE

CTRL_H_DIR

CTRL_L_ENABLE

CTRL_L_DIR

ADDR[0..7]

CTRL[0..15]

CTRL[0..15]

ADDR[0..7]

DATA_L_ENABLE

DATA_L_DIR

DATA_H_ENABLE

DATA_H_DIR

ADDR_ENABLE

ADDR_DIR

D[0..15]

D[0..15] DO[0..12]

1 2

74AC14SC

U7A

3 4

74AC14SC

U7B

5 6

74AC14SC

U7C

9 8

74AC14SC

U7D

11 10

74AC14SC

U7E

13 12

74AC14SC

U7F

LXH103ID

LED2

LXH103ID

LED3

1 2

3 4

5 6

7 8

9 10

HEADER5x2

JP2 VCC


Figure 10
[image: image13.emf]1

1

2

2

3

3

4

4

5

5

6

6

D D

C C

B B

A A

4

UH - Institute for Astonomy

640 N. Aohoku Place

Hilo, HI 96720

USA

6

Analog Output

IRTF - AO A

5/9/2003 4:44:49 PM

C:\Documents and Settings\Fred keske\My Documents\Projects\High Voltage\DAC.SCHDOC

Title

Size: Number:

Date:

File:

Revision:

Sheet of Time:

B

Engineer

Fred Keske

VOUTB

1

VOUTA

2

VDD

3

REFAB

4

AGNDAB

5

LDAB

6

LDCD

7

CS

8

WR

9

A2

10

A1

11

A0

12

D12

13

D11

14

D10

15

D9

16

D8

17

D7

18

D6

19

D5

20

D4

21

D3

22

D2

23

D1

24

D0

25

LDEF

26

LDGH

27

GND

28

AGNDGH

29

REFGH

30

VDD

31

VOUTH

32

VOUTG

33

VOUTF

34

VOUTE

35

VSS

36

REFF

37

AGNDF

38

CLR

39

AGNDCD

40

REFCD

41

VSS

42

VOUTD

43

VOUTC

44

MAX547

U13

VOUTB

1

VOUTA

2

VDD

3

REFAB

4

AGNDAB

5

LDAB

6

LDCD

7

CS

8

WR

9

A2

10

A1

11

A0

12

D12

13

D11

14

D10

15

D9

16

D8

17

D7

18

D6

19

D5

20

D4

21

D3

22

D2

23

D1

24

D0

25

LDEF

26

LDGH

27

GND

28

AGNDGH

29

REFGH

30

VDD

31

VOUTH

32

VOUTG

33

VOUTF

34

VOUTE

35

VSS

36

REFF

37

AGNDF

38

CLR

39

AGNDCD

40

REFCD

41

VSS

42

VOUTD

43

VOUTC

44

MAX547

U14

ADDRI0

ADDRI1

ADDRI2

VCC

49.9k

1/8 W

R11

VCC

14k

1/8 W

R17

35.7k

1/8 W

R16

33pF

50 V

C41

7.32k

1/8 W

R14

7.32k

1/8 W

R15

7.32k

1/8 W

R13

DAALD0

DAALD1

DAALD2

DAALD3

VCC

VCC

VCC

SHDN0

SHDN1

SHDN2

VCC

VCC

-5V

3

2

6

7

4

1, 5

8

MAX4131ESA

U9

3

2

6

7

4

1, 5

8

MAX4131ESA

U11

VCC

7.32k

1/8 W

R12

VCC

3

2

6

7

4

1, 5

8

MAX4131ESA

U10

-5V -5V

VCC

VCC -5V

-5V -5V

VCC

ADDRI1

ADDRI2

DAILD0

DAILD1

VCC

DO0 DO1 DO2 DO3 DO4 DO5 DO6 DO7 DO8 DO9

DO10 DO11 DO12

ADDRI0

ADDRI1

ADDRI2

ADDRI0

DO0 DO1 DO2 DO3 DO4 DO5 DO6 DO7 DO8 DO9

DO10 DO11 DO12

DAALD0

DAALD1

DAALD2

DAALD3

DAILD0

DAILD1

DAILD2

DAILD3

VOUT1

DAILD2

DAILD3

4

NC NC NC

NC NC

1, 2, 3

5, 7

6, 8

REF1004C1.2

U12

VOUT[1..16]

0.1uF

50 V

C21

0.1uF

50 V

C22

0.1uF

50 V

C24

0.1uF

50 V

C23

0.1uF

50 V

C25

0.1uF

50 V

C26

0.1uF

50 V

C27

0.1uF

50 V

C28

0.1uF

50 V

C29

0.1uF

50 V

C30

0.1uF

50 V

C31

0.1uF

50 V

C32

0.1uF

50 V

C37

0.1uF

50 V

C38

0.1uF

50 V

C39

0.1uF

50 V

C40

0.1uF

50 V

C33

0.1uF

50 V

C34

0.1uF

50 V

C35

0.1uF

50 V

C36

33pF

50 V

C43

33pF

50 V

C42

VOUT1

VOUT2

VOUT3VOUT4 VOUT5VOUT6

VOUT7

VOUT8 VOUT9

VOUT10

VOUT11VOUT12 VOUT13VOUT14

VOUT15

VOUT16

VOUT2

VOUT3

VOUT4

VOUT5

VOUT6

VOUT7

VOUT8

VOUT9

VOUT10

VOUT11

VOUT12

VOUT13

VOUT14

VOUT15

VOUT16

CLR

DACS0

WR

ADDRI[0..2]

DAALD[0..3]

DO[0..12]

DACS1

DAILD[0..3]

SHDN[0..2]

SHDN[0..2]

ADDRI[0..2]

DAALD[0..3]

DAILD[0..3]

DO[0..12]

WR

SHDN[0..2]

CLR

DACS0 DACS1

CLR

WR

ADDRI[0..2]

DAALD[0..3]

DO[0..12]

DAILD[0..3]

VOUT[1..16]

22uF

10 V

C12

22uF

10 V

C18

-5V VCC


Figure 11
[image: image14.emf]1

1

2

2

3

3

4

4

5

5

6

6

D D

C C

B B

A A

5

UH - Institute for Astonomy

640 N. Aohoku Place

Hilo, HI 96720

USA

6

High Voltage Channel

IRTF - AO A

5/9/2003 4:45:18 PM

C:\Documents and Settings\Fred keske\My Documents\Projects\High Voltage\High Voltage Channel.SCHDOC

Title

Size: Number:

Date:

File:

Revision:

Sheet of Time:

B

Engineer

Fred Keske

4.99k

1/8 W

R23

549

1/8 W

R27

33pF

1000 V

C46

0.033uF

50 V

C48

7

12

1

2

4

8

10

6

PA95

U15

IN

HV_OUTPUT

5pF

1000 V

C47

0.1uF

630 V

C45

0.1uF

630 V

C44

+HV

-HV

1

2

3

HEADER 3

JP5

4.99k

1/8 W

R22

165k

1/8 W

R19

165k

1/8 W

R20

165k

1/8 W

R21

3.3

1/8 W

R24

3.3

1/8 W

R25

3.3

1/8 W

R26

17.4

1/8 W

R28

17.4

1/8 W

R29

17.4

1/8 W

R30


Figure 12
[image: image15.emf]1

1

2

2

3

3

4

4

5

5

6

6

D D

C C

B B

A A

6

UH - Institute for Astonomy

640 N. Aohoku Place

Hilo, HI 96720

USA

6

5V Power Supply

IRTF - AO A

8/12/2003 10:55:05 AM

C:\Documents and Settings\Fred keske\My Documents\My Designs\High Voltage\Power Supply.SCHDOC

Title

Size: Number:

Date:

File:

Revision:

Sheet of Time:

B

Engineer

Fred Keske

DIODE SCHOTTKY

20 V

D1

220uF

10 V

C53

47uH

L1

220uF

10 V

C52

OUT

1

FB

2

SHDN

3

REF

4

GND

5

V+

6

V+

7

LX

8

MAX764CSA

U16 VCC

-5V

0.1uF

50 V

C49

0.1uF

50 V

C51

0.1uF

50 V

C50


Figure 13
AGND





GND





JP3


Pin 1 top





NC





NC





VCC





5





4





3





2





1





2





3





3





2





1





AGND





1





HV+





HV-





AGND





Amp


In





DAC


Out





Figure � SEQ Figure \* ARABIC �2�





Figure � SEQ Figure \* ARABIC �3�








JP4


Pin 1 top





JP5x


Pin 1 left








PAGE  
Page 5

_1112619027.vsd
�

�

Bus
Interface�

VME
Backplane
Interface�

CPLD�

�

(x2) 8 channel
13 Bit DAC�

High Voltage
Connector
Interface�

16 Channels�

High Voltage
Amplifiers�

x16�

Power
Supply�

+5V�

-5V�


