## Sensor ID and Status
Sensor ID | Sensor Type | Model No. | Location | Elevation | Cable Length | Monitor Device | Status & Notes | Additional Notes
TE1 | Vaisala | Vaisala HMD60YO | External Dome Shutter | 50' |  | OM1 - Omega DP472-V-C2 | Reading O.K. when sensor does not get wet
TE2 | Vaisala | Vaisala HMD60T | External Dome Shutter | 50' |  | OM1 - Omega DP472-V-C2 | TE1, TE2, and HE1 are currently being looked into
HE1 | Vaisala | Vaisala HMD60YO | External Dome Shutter | 50' |  | OM1 - Omega DP472-V-C2 | being replaced with Omega sensors
HD1 | Johnson Cntl | TCS/1200 H8 HE-2 | Dome |  |  | OM1 - Omega DP472-V-C2
HD2 | Johnson Cntl | TCS/1200 H8 HE-9 | Dome |  |  | OM1 - Omega DP472-V-C2 | HD1, HD2 and HD3 do not read below 25%
HD3 | Johnson Cntl | TCS/1200 H8 HE-5 | Dome |  |  | OM1 - Omega DP472-V-C2
TD1 | PT100 RTD | TE-6323P1 TE-2 (14) | Dome |  |  | LS3 - LakeShore Monitor 218E | O.K. (Calibrated against cal. sensor as of 1-29-04) | TD1 Subtract 2.064 Deg. C (3-22-04)
TD2 | PT100 RTD | TE-6323P1 TE-7 (19) | Dome |  |  | LS3 - LakeShore Monitor 218E | O.K. (Calibrated against cal. sensor as of 1-29-04) | TD2 Subtract 2.264 Deg. C (3-22-04)
TD3 | PT100 RTD | TE-6323P1 TE-9 (21) | Dome |  |  | LS3 - LakeShore Monitor 218E | O.K. (Cal. 6-26-03) | 3 - subtract 1.648 degC | These offsets were based on dipping the sensors in ice water and taking the average of a few data points.
TD4 | PT100 RTD | TE-6323P1 TE-3 (15) | Dome |  |  | LS3 - LakeShore Monitor 218E | O.K. (Cal. 6-26-03) | 4 - subtract 2.696 degC
TD5 | PT100 RTD | TE-6323P1 TE-6 (18) | Dome |  |  | LS3 - LakeShore Monitor 218E | O.K. (Cal. 6-26-03) | 5 - subtract 2.591 degC | http://irtfweb.ifa.hawaii.edu/~iqup/domeenv/data/2003/06/26cal.txt
TD6 | PT100 RTD | TE-6323P1 TE-8 (20) | Dome |  |  | LS3 - LakeShore Monitor 218E | O.K. (Cal. 6-26-03) | 6 - subtract 2.158 degC
TD7 | PT100 RTD | TE-6323P1 TE-1 (13) | Dome |  |  | LS3 - LakeShore Monitor 218E | O.K. (Cal. 6-26-03) | 7 - subtract 2.258 degC
TF1 | PT100 RTD | Omega RTD-805 | Floor Level - North |  | 90' (+25ft) | LS4 - LakeShore Monitor 218E | O.K. (Added 25 ft.ext. 2-10-04) | TF1 Subtract 0.892 Deg. C (3-22-04)
TF2 | PT100 RTD | Omega RTD-805 | Floor Level - South |  | 67' (+25ft) | LS4 - LakeShore Monitor 218E | O.K. (Added 25 ft.ext. 2-10-04) | TF2 Subtract 1.1845 Deg. C (3-22-04)
TF3 | PT100 RTD | Omega RTD-805 | Floor Level - East |  | 127' (+25ft) | LS4 - LakeShore Monitor 218E | O.K. (Added 25 ft.ext. 2-10-04) | TF3 Subtract .805 Deg. C (3-22-04)
TF4 | PT100 RTD | Omega RTD-805 | Floor Level - West |  | 20' (+25ft) | LS4 - LakeShore Monitor 218E | O.K. (Added 25 ft.ext. 2-10-04) | TF4 Subtract 1.3165 Deg. C (3-22-04)
TMS3 | PT100 RTD |  | Heat Exchanger Output |  |  | LS4 - LakeShore Monitor 218E | Not installed
TT1 | PT100 RTD | Omega RTD-805 | Hexapod Back - Air |  | 50' | LS1 - LakeShore Monitor 218E | O.K.
TT2 | PT100 RTD | Omega RTD-830 | Secondary Truss Surface |  | 60' | LS1 - LakeShore Monitor 218E | O.K.
TT3 | PT100 RTD | Omega RTD-830 | Secondary Truss Surface |  | 40' | LS1 - LakeShore Monitor 218E | O.K.
TT4 | PT100 RTD | Omega RTD-805 | Secondary Truss Air |  | 40' | LS1 - LakeShore Monitor 218E | O.K.
TT5 | PT100 RTD | Omega RTD-830 | Secondary Truss Surface |  | 30' | LS1 - LakeShore Monitor 218E | O.K.
TT6 | PT100 RTD | Omega RTD-805 | Secondary Truss Air |  | 30' | LS1 - LakeShore Monitor 218E | O.K.
TT7 | PT100 RTD | Omega RTD-830 | Secondary Truss Surface |  | 20' | LS1 - LakeShore Monitor 218E | O.K.
TT8 | PT100 RTD | Omega RTD-805 | Secondary Truss Air |  | 20' | LS1 - LakeShore Monitor 218E | O.K.
TM1 | PT100 RTD | Omega RTD-3-F3102-36-T | Mirror Surface |  | 20' | LS2 - LakeShore Monitor 218E | O.K.
TM2 | PT100 RTD | Omega RTD-3-F3102-36-T | Mirror Surface |  | 22' | LS2 - LakeShore Monitor 218E | O.K.
TM3 | PT100 RTD | Omega RTD-3-F3102-36-T | Mirror Surface |  | 12' | LS2 - LakeShore Monitor 218E | O.K.
TM4 | PT100 RTD | Omega RTD-3-F3102-36-T | Mirror Surface |  | 8' + 7.75' | LS2 - LakeShore Monitor 218E | O.K.
TM5 | PT100 RTD | Omega RTD-3-F3102-36-T | Mirror Surface |  | 8' + 8.2' | LS2 - LakeShore Monitor 218E | O.K.
TM6 | PT100 RTD | Omega RTD-3-F3102-36-T | Mirror Surface |  | 8' + 6.25' | LS2 - LakeShore Monitor 218E | O.K.
TMS1 | PT100 RTD | Omega RTD-805 | Mirror Cell - East |  | 20' | LS2 - LakeShore Monitor 218E | O.K.
TMS2 | PT100 RTD | Omega RTD-805 | Mirror Cell - West |  | 20' | LS2 - LakeShore Monitor 218E | O.K.
Dewpt | Vaisala | HMP243 | Mirror Cell |  |  | Serial port to Digi-Mim | O.K.
Wnd Spd1 | Vaisala | 425A | Ext Fan Room | 20' |  | Serial port to Digi-TCS | Replacement installed 3-12 Stat-O.K.
Wnd Spd2 | Vaisala | 425A | IRTF Roof |  |  | Serial port split from Amiga to | Not done
Digi-TCS
Mir Load Cells | Newport |  | Bottom of Pri Mirror |  |  | Serial ports to Digi-Mim | Lab tested - not installed
1 to 4

## Connections
LS1 - LakeShore 218E
Sensor ID | Sensor Type | Signal |  |  | 6 Pin (P) | 6 Pin (S) | Sig Name | POD1 Inpt 1-4 | Sig Name (25D)
TT1 | PT100 RTD | V+ | Red | Red | 1 | 1 | V+ | 4 | 1V+
| RTD-805 | V- | Blk | Blk | 2 | 2 | V- | 16 | 1V-
|  | I+ |  | Wht | 3 | 3 | I+ | 3 | 1I+
|  | I- | Blk | Blk | 4 | 4 | I- | 15 | 1I-
|  |  |  | SHLD | 5 | 5 | SHLD | 2 or 14 | S |  |  |  | Serial Device Interface
TT2 | PT100 RTD | V+ | Red | Red | 1 | 1 | V+ | 7 | 2V+ | [7 empty] | digimim
| RTD-830 | V- | Blk | Blk | 2 | 2 | V- | 19 | 2V- | LakeShore |  |  | 9 Pin D to RJ45 Converter |  | Cat5 Cable |  | Digi PortServer Port # 13
|  | I+ | Red | Wht | 3 | 3 | I+ | 6 | 2I+ | Serial I/O (DTE)
|  | I- | Blk | Blk | 4 | 4 | I- | 18 | 2I- | 9 pinD (P) | Sig Name | 9 Pin D | Wire Clr | RJ45 | Wire Color | Sig Name | RJ45 (8) | RJ45 (10)
|  | SHLD |  | SHLD | 5 | 5 | SHLD | 5 or 17 | S | 1 | nc
TT3 | PT100 RTD | V+ | Red | Red | 1 | 1 | V+ | 10 | 3V+ | 2 | RXD | 2 | RED | 4 | 4 (BLU) | TXD | 4 | 5
| RTD-830 | V- | Blk | Blk | 2 | 2 | V- | 22 | 3V- | 3 | TXD | 3 | GRN | 5 | 5 (BLU/WHT) | RXD | 5 | 6
|  | I+ | Red | Wht | 3 | 3 | I+ | 9 | 3I+ | 4 | DTR
|  | I- | Blk | Blk | 4 | 4 | I- | 21 | 3I- | 5 | GND | 5 | YEL | 6 | 6 (ORG) | GND(S) | 6 | 7
|  | SHLD |  | SHLD | 5 | 5 | SHLD | 8 or 20 | S | 6 | DSR
TT4 | PT100 RTD | V+ | Red | Red | 1 | 1 | V+ | 13 | 4V+ | 7 | DTR
| RTD-805 | V- | Blk | Blk | 2 | 2 | V- | 25 | 4V- | 8 | nc
|  | I+ |  | Wht | 3 | 3 | I+ | 12 | 4I+ | 9 | nc
|  | I- | Blk | Blk | 4 | 4 | I- | 24 | 4I-
|  |  |  | SHLD | 5 | 5 | SHLD | 11 or 23 | S
POD2 Inpt 5-8
TT5 | PT100 RTD | V+ | Red | Red | 1 | 1 | V+ | 4 | 5V+
| RTD-830 | V- | Blk | Blk | 2 | 2 | V- | 16 | 5V-
|  | I+ | Red | Wht | 3 | 3 | I+ | 3 | 5I+
|  | I- | Blk | Blk | 4 | 4 | I- | 15 | 5I-
|  | SHLD |  | SHLD | 5 | 5 | SHLD | 2 or 14 | S
TT6 | PT100 RTD | V+ | Red | Red | 1 | 1 | V+ | 7 | 6V+
| RTD-805 | V- | Blk | Blk | 2 | 2 | V- | 19 | 6V-
|  | I+ |  | Wht | 3 | 3 | I+ | 6 | 6I+
|  | I- | Blk | Blk | 4 | 4 | I- | 18 | 6I-
|  |  |  | SHLD | 5 | 5 | SHLD | 5 or 17 | S
TT7 | PT100 RTD | V+ | Red | Red | 1 | 1 | V+ | 10 | 7V+
| RTD-830 | V- | Blk | Blk | 2 | 2 | V- | 22 | 7V-
|  | I+ | Red | Wht | 3 | 3 | I+ | 9 | 7I+
|  | I- | Blk | Blk | 4 | 4 | I- | 21 | 7I-
|  | SHLD |  | SHLD | 5 | 5 | SHLD | 8 or 20 | S
TT8 | PT100 RTD | V+ | Red | Red | 1 | 1 | V+ | 13 | 8V+
| RTD-805 | V- | Blk | Blk | 2 | 2 | V- | 25 | 8V-
|  | I+ |  | Wht | 3 | 3 | I+ | 12 | 8I+
|  | I- | Blk | Blk | 4 | 4 | I- | 24 | 8I-
|  | SHLD |  | SHLD | 5 | 5 | SHLD | 11 or 23 | S
LS2 - LakeShore 218E
[5 empty] | 6 Pin (P) | 6 Pin (S) | Sig Name | POD3 Inpt 1-4 | Sig Name (25D)
TM1 | PT100 RTD | V+ | Red | Red | 1 | 1 | V+ | 4 | 1V+
| RTD-3-F3102 | V- | Blk | Blk | 2 | 2 | V- | 16 | 1V-
|  | I+ | Red | Wht | 3 | 3 | I+ | 3 | 1I+
|  | I- | Blk | Blk | 4 | 4 | I- | 15 | 1I-
|  |  |  | SHLD | 5 | 5 | SHLD | 2 or 14 | S |  |  |  | Serial Device Interface
TM2 | PT100 RTD | V+ | Red | Red | 1 | 1 | V+ | 7 | 2V+ | [7 empty] | digimim
| RTD-3-F3102 | V- | Blk | Blk | 2 | 2 | V- | 19 | 2V- | LakeShore |  |  | 9 Pin D to RJ45 Converter |  | Cat5 Cable |  | Digi PortServer Port # 14
|  | I+ | Red | Wht | 3 | 3 | I+ | 6 | 2I+ | Serial I/O (DTE)
|  | I- | Blk | Blk | 4 | 4 | I- | 18 | 2I- | 9 pinD (P) | Sig Name | 9 Pin D | Wire Clr | RJ45 | Wire Color | Sig Name | RJ45 (8) | RJ45 (10)
|  | SHLD |  | SHLD | 5 | 5 | SHLD | 5 or 17 | S | 1 | nc
TM3 | PT100 RTD | V+ | Red | Red | 1 | 1 | V+ | 10 | 3V+ | 2 | RXD | 2 | RED | 4 | 4 (BLU) | TXD | 4 | 5
| RTD-3-F3102 | V- | Blk | Blk | 2 | 2 | V- | 22 | 3V- | 3 | TXD | 3 | GRN | 5 | 5 (BLU/WHT) | RXD | 5 | 6
|  | I+ | Red | Wht | 3 | 3 | I+ | 9 | 3I+ | 4 | DTR
|  | I- | Blk | Blk | 4 | 4 | I- | 21 | 3I- | 5 | GND | 5 | YEL | 6 | 6 (ORG) | GND(S) | 6 | 7
|  | SHLD |  | SHLD | 5 | 5 | SHLD | 8 or 20 | S | 6 | DSR
TM4 | PT100 RTD | V+ | Red | Red | 1 | 1 | V+ | 13 | 4V+ | 7 | DTR
| RTD-3-F3102 | V- | Blk | Blk | 2 | 2 | V- | 25 | 4V- | 8 | nc
|  | I+ | Red | Wht | 3 | 3 | I+ | 12 | 4I+ | 9 | nc
|  | I- | Blk | Blk | 4 | 4 | I- | 24 | 4I-
|  |  |  | SHLD | 5 | 5 | SHLD | 11 or 23 | S
POD4 Inpt 5-8
TM5 | PT100 RTD | V+ | Red | Red | 1 | 1 | V+ | 4 | 5V+
| RTD-3-F3102 | V- | Blk | Blk | 2 | 2 | V- | 16 | 5V-
|  | I+ | Red | Wht | 3 | 3 | I+ | 3 | 5I+
|  | I- | Blk | Blk | 4 | 4 | I- | 15 | 5I-
|  | SHLD |  | SHLD | 5 | 5 | SHLD | 2 or 14 | S
TM6 | PT100 RTD | V+ | Red | Red | 1 | 1 | V+ | 7 | 6V+
| RTD-3-F3102 | V- | Blk | Blk | 2 | 2 | V- | 19 | 6V-
|  | I+ | Red | Wht | 3 | 3 | I+ | 6 | 6I+
|  | I- | Blk | Blk | 4 | 4 | I- | 18 | 6I-
|  |  |  | SHLD | 5 | 5 | SHLD | 5 or 17 | S
TMS1 | PT100 RTD |  |  | Red | 1 | 1 | V+ | 10 | 7V+
|  |  |  | Blk | 2 | 2 | V- | 22 | 7V-
|  |  |  | Wht | 3 | 3 | I+ | 9 | 7I+
|  |  |  | Blk | 4 | 4 | I- | 21 | 7I-
|  |  |  | SHLD | 5 | 5 | SHLD | 8 or 20 | S
TMS2 | PT100 RTD |  |  | Red | 1 | 1 | V+ | 13 | 8V+
|  |  |  | Blk | 2 | 2 | V- | 25 | 8V-
|  |  |  | Wht | 3 | 3 | I+ | 12 | 8I+
|  |  |  | Blk | 4 | 4 | I- | 24 | 8I-
|  |  |  | SHLD | 5 | 5 | SHLD | 11 or 23 | S
Omega DP472
[9 empty] | 6 Chan Multiple Input |  |  |  |  | Serial Device Interface
|  |  |  | JC-Cable | Wire Color | Terminal Blk | Wire Color | [10 empty] | digidome
TE1 | Vaisala | Vaisala HMD60YO Al.Elec |  | 13 (Left side) | Red | +24VDC |  |  |  |  | Omega |  |  | 9 Pin D to RJ45 Converter |  | Cat5 Cable |  | Digi PortServer Port # 16
[5 empty] | Blk |  | Org | CH 1 + | 1 |  | Serial
TE2 | Vaisala | Vaisala HMD60T Opp Elect |  | 14 (Left side) | Red | +24VDC |  | CH1 - | 2 |  | 9 pinD (P) | Sig Name | 9 Pin D | Wire Clr | RJ45 | Wire Color | Sig Name | RJ45 (8) | RJ45 (10)
[5 empty] | Blk |  | Yel | CH2 + | 3 |  | 1 | nc
HE1 | Vaisala | Vaisala HMD60YO Al.Elec |  | 15 (Right side) | Red | +24VDC |  | CH2 - | 4 |  | 2 | RXD | 2 | RED | 4 | 4 (BLU) | TXD | 4 | 5
[5 empty] | Blk |  | Grn | CH3 + | 5 |  | 3 | TXD | 3 | GRN | 5 | 5 (BLU/WHT) | RXD | 5 | 6
HD1 | Johnson Cntl | TCS/1200 H8 HE-2 |  | 2 (Left side) | Red | +24VDC |  | CH3 - | 6 |  | 4 | DTR
[5 empty] | Blk |  | Blu | CH4 + | 7 |  | 5 | GND | 5 | YEL | 6 | 6 (ORG) | GND(S) | 6 | 7
HD2 | Johnson Cntl | TCS/1200 H8 HE-9 |  | 9 (Left side) | Red | +24VDC |  | CH4 - | 8 |  | 6 | DSR
[5 empty] | Blk |  | Vio | CH5 + | 9 |  | 7 | DTR
HD3 | Johnson Cntl | TCS/1200 H8 HE-5 |  | 5 (Left side) | Red | +24VDC |  | CH5 - | 10 |  | 8 | nc
[5 empty] | Blk |  | Gry | CH6 + | 11 |  | 9 | nc
[8 empty] | CH6 - | 12
24VDC-GND
LS3 - LakeShore 218E
[6 empty] | TB2 |  | Sig Name | POD5 Inpt 1-4 | Sig Name (25D)
TD1 | PT100 RTD | TE-6323P1 TE-2 (14) |  | 2 (Right side) | Red | 1T | Brn | V+ | 4 | 1V+
[5 empty] | Blk | 1B | Wht | V- | 16 | 1V-
[8 empty] | I+ | 3 | 1I+
[8 empty] | I- | 15 | 1I-
[8 empty] | SHLD | 2 or 14 | S
TD2 | PT100 RTD | TE-6323P1 TE-7 (19) |  | 7 (Left side) | Red | 2T | Red | V+ | 7 | 2V+
[5 empty] | Blk | 2B | Wht | V- | 19 | 2V-
[8 empty] | I+ | 6 | 2I+
[8 empty] | I- | 18 | 2I-
[8 empty] | SHLD | 5 or 17 | S
TD3 | PT100 RTD | TE-6323P1 TE-9 (21) |  | 9 (Right side) | Red | 3T | Org | V+ | 10 | 3V+
[5 empty] | Blk | 3B | Wht | V- | 22 | 3V-
[8 empty] | I+ | 9 | 3I+ |  |  |  | Serial Device Interface
[8 empty] | I- | 21 | 3I- | [7 empty] | digidome
[8 empty] | SHLD | 8 or 20 | S | LakeShore |  |  | 9 Pin D to RJ45 Converter |  | Cat5 Cable |  | Digi PortServer Port # 9
TD4 | PT100 RTD | TE-6323P1 TE-3 (15) |  | 3 (Right side) | Red | 4T | Yel | V+ | 13 | 4V+ | Serial I/O (DTE)
[5 empty] | Blk | 4B | Wht | V- | 25 | 4V- | 9 pinD (P) | Sig Name | 9 Pin D | Wire Clr | RJ45 | Wire Color | Sig Name | RJ45 (8) | RJ45 (10)
[8 empty] | I+ | 12 | 4I+ | 1 | nc
[8 empty] | I- | 24 | 4I- | 2 | RXD | 2 | RED | 4 | 4 (BLU) | TXD | 4 | 5
[8 empty] | SHLD | 11 or 23 | S | 3 | TXD | 3 | GRN | 5 | 5 (BLU/WHT) | RXD | 5 | 6
[9 empty] | POD6 Inpt 5-8 |  | 4 | DTR
TD5 | PT100 RTD | TE-6323P1 TE-6 (18) |  | 6 (Leftt side) | Red | 5T | Grn | V+ | 4 | 5V+ | 5 | GND | 5 | YEL | 6 | 6 (ORG) | GND(S) | 6 | 7
[5 empty] | Blk | 5B | Blk | V- | 16 | 5V- | 6 | DSR
[8 empty] | I+ | 3 | 5I+ | 7 | DTR
[8 empty] | I- | 15 | 5I- | 8 | nc
[8 empty] | SHLD | 2 or 14 | S | 9 | nc
TD6 | PT100 RTD | TE-6323P1 TE-8 (20) |  | 8 (Left side) | Red | 6T | Blu | V+ | 7 | 6V+
[5 empty] | Blk | 6B | Blk | V- | 19 | 6V-
[8 empty] | I+ | 6 | 6I+
[8 empty] | I- | 18 | 6I-
[8 empty] | SHLD | 5 or 17 | S
TD7 | PT100 RTD | TE-6323P1 TE-1 (13) |  | 1 (Right side) | Red | 7T | Vio | V+ | 10 | 7V+
[5 empty] | Blk | 7B | Blk | V- | 22 | 7V-
[8 empty] | I+ | 9 | 7I+
[8 empty] | I- | 21 | 7I-
[8 empty] | SHLD | 8 or 20 | S
[6 empty] | 8T | Gry | V+ | 13 | 8V+
[6 empty] | 8B | Blk | V- | 25 | 8V-
[8 empty] | I+ | 12 | 8I+
[8 empty] | I- | 24 | 8I-
[8 empty] | SHLD | 11 or 23 | S
LS4 - LakeShore 218E
Sensor ID | Sensor Type | Signal |  |  | 6 Pin (P) | 6 Pin (S) | Sig Name | POD7 Inpt 1-4 | Sig Name (25D)
TF1 | PT100 RTD | V+ | Red | Red | 1 | 1 | V+ | 4 | 1V+
North Wall | RTD-805 | V- | Blk | Blk | 2 | 2 | V- | 16 | 1V-
L=90' |  | I+ |  | Wht | 3 | 3 | I+ | 3 | 1I+
|  | I- | Blk | Blk | 4 | 4 | I- | 15 | 1I-
|  |  |  | SHLD | 5 | 5 | SHLD | 2 or 14 | S |  |  |  | Serial Device Interface
TF2 | PT100 RTD | V+ | Red | Red | 1 | 1 | V+ | 7 | 2V+ | [7 empty] | Digitcs
South Wall | RTD-830 | V- | Blk | Blk | 2 | 2 | V- | 19 | 2V- | LakeShore |  |  | 9 Pin D to RJ45 Converter |  | Cat5 Cable |  | Digi PortServer Port #16
L=67' |  | I+ | Red | Wht | 3 | 3 | I+ | 6 | 2I+ | Serial I/O (DTE)
|  | I- | Blk | Blk | 4 | 4 | I- | 18 | 2I- | 9 pinD (P) | Sig Name | 9 Pin D | Wire Clr | RJ45 | Wire Color | Sig Name | RJ45 (8) | RJ45 (10)
|  | SHLD |  | SHLD | 5 | 5 | SHLD | 5 or 17 | S | 1 | nc
TF3 | PT100 RTD | V+ | Red | Red | 1 | 1 | V+ | 10 | 3V+ | 2 | RXD | 2 | RED | 4 | 4 (BLU) | TXD | 4 | 5
East Wall | RTD-830 | V- | Blk | Blk | 2 | 2 | V- | 22 | 3V- | 3 | TXD | 3 | GRN | 5 | 5 (BLU/WHT) | RXD | 5 | 6
L=127' |  | I+ | Red | Wht | 3 | 3 | I+ | 9 | 3I+ | 4 | DTR
|  | I- | Blk | Blk | 4 | 4 | I- | 21 | 3I- | 5 | GND | 5 | YEL | 6 | 6 (ORG) | GND(S) | 6 | 7
|  | SHLD |  | SHLD | 5 | 5 | SHLD | 8 or 20 | S | 6 | DSR
TF4 | PT100 RTD | V+ | Red | Red | 1 | 1 | V+ | 13 | 4V+ | 7 | DTR
West Wall | RTD-805 | V- | Blk | Blk | 2 | 2 | V- | 25 | 4V- | 8 | nc
L=20' |  | I+ |  | Wht | 3 | 3 | I+ | 12 | 4I+ | 9 | nc
|  | I- | Blk | Blk | 4 | 4 | I- | 24 | 4I-
|  |  |  | SHLD | 5 | 5 | SHLD | 11 or 23 | S

## Sensor Calibration Curves
Telescope Truss Sensors | [10 empty] | LakeShore 6819 Calibrated PT-100 Sensor
Data Point | Cal (C) | Cal (K) | TT1 Ohms | TT2 Ohms | TT3 Ohms | TT4 Ohms | TT5 Ohms | TT6 Ohms | TT7 Ohms | TT8 Ohms | Data Point | Deg K | Ohms
1 | -30.6 | 242.55 | 88.21 | 88.58 | 88.16 | 88.12 | 87.92 | 87.93 | 87.93 | 87.88 | 1 | 1.64 | 12.172
2 | -25.77 | 247.38 | 90.12 | 90.5 | 90.07 | 90.03 | 89.83 | 89.81 | 89.84 | 89.79 | 2 | 1.669 | 13.175
3 | -20.72 | 252.43 | 92.1 | 92.52 | 92.08 | 92.01 | 91.83 | 91.79 | 91.84 | 91.77 | 3 | 1.704 | 14.171
4 | -15.72 | 257.43 | 94.11 | 94.52 | 94.07 | 94.02 | 93.81 | 93.82 | 93.83 | 93.77 | 4 | 1.746 | 15.162
5 | -10.89 | 262.26 | 95.97 | 96.43 | 95.98 | 95.87 | 95.72 | 95.66 | 95.73 | 95.62 | 5 | 1.795 | 16.143
6 | -5.82 | 267.33 | 98 | 98.45 | 97.98 | 97.91 | 97.71 | 97.69 | 97.72 | 97.64 | 6 | 1.852 | 17.127
7 | -0.77 | 272.38 | 100.01 | 100.46 | 99.98 | 99.91 | 99.7 | 99.68 | 99.71 | 99.62 | 7 | 1.916 | 18.1
8 | 4.342 | 277.492 | 102 | 102.48 | 102 | 101.89 | 101.71 | 101.66 | 101.72 | 101.6 | 8 | 1.989 | 19.08
9 | 9.229 | 282.379 | 103.95 | 104.42 | 103.93 | 103.85 | 103.62 | 103.62 | 103.63 | 103.54 | 9 | 2.07 | 20.057
10 | 14.214 | 287.364 | 105.94 | 106.39 | 105.89 | 105.83 | 105.57 | 105.6 | 105.58 | 105.51 | 10 | 2.172 | 21.14
11 | 18.808 | 291.958 | 107.74 | 108.21 | 107.69 | 107.63 | 107.37 | 107.39 | 107.37 | 107.3 | 11 | 2.341 | 22.72
12 | 21.943 | 295.093 | 108.96 | 109.44 | 108.93 | 108.84 | 108.59 | 108.6 | 108.59 | 108.48 | 12 | 2.537 | 24.307
13 | 23.872 | 297.022 | 109.71 | 110.2 | 109.68 | 109.6 | 109.35 | 109.36 | 109.34 | 109.24 | 13 | 2.762 | 25.905
14 | 28.746 | 301.896 | 111.59 | 112.12 | 111.59 | 111.47 | 111.24 | 111.22 | 111.24 | 111.1 | 14 | 3.022 | 27.547
[11 empty] | 15 | 3.304 | 29.147
[11 empty] | 16 | 3.659 | 30.962
Primary Mirror Sensors | [10 empty] | 17 | 4.116 | 33.07
Data Point | Cal (C) | Cal (K) | TM1 Ohms | TM2 Ohms | TM3 Ohms | TM4 Ohms | TM5 Ohms | TM6 Ohms |  |  | 18 | 4.847 | 36.074
[11 empty] | 19 | 5.662 | 39.061
1 | -30.61 | 242.54 | 87.77 | 87.78 | 87.85 | 87.85 | 87.84 | 87.86 |  |  | 20 | 6.557 | 42.052
2 | -25.71 | 247.44 | 89.72 | 89.73 | 89.8 | 89.8 | 89.79 | 89.8 |  |  | 21 | 7.524 | 45.05
3 | -20.66 | 252.49 | 91.71 | 91.73 | 91.8 | 91.8 | 91.79 | 91.8 |  |  | 22 | 8.553 | 48.051
4 | -15.63 | 257.52 | 93.7 | 93.72 | 93.79 | 93.79 | 93.78 | 93.79 |  |  | 23 | 9.268 | 50.048
5 | -10.64 | 262.51 | 95.67 | 95.69 | 95.76 | 95.76 | 95.75 | 95.76 |  |  | 24 | 11.14 | 55.041
6 | -5.57 | 267.58 | 97.66 | 97.69 | 97.76 | 97.76 | 97.75 | 97.76 |  |  | 25 | 13.112 | 60.047
7 | -0.83 | 272.32 | 99.52 | 99.55 | 99.62 | 99.63 | 99.61 | 99.63 |  |  | 26 | 15.147 | 65.048
8 | 4.271 | 277.421 | 101.53 | 101.56 | 101.63 | 101.64 | 101.62 | 101.63 |  |  | 27 | 17.222 | 70.039
9 | 9.245 | 282.395 | 103.48 | 103.5 | 103.58 | 103.58 | 103.57 | 103.58 |  |  | 28 | 19.323 | 75.027
10 | 14.141 | 287.291 | 105.39 | 105.41 | 105.49 | 105.5 | 105.48 | 105.49 |  |  | 29 | 21.443 | 80.025
11 | 19.165 | 292.315 | 107.34 | 107.37 | 107.45 | 107.46 | 107.44 | 107.45 |  |  | 30 | 23.571 | 85.024
12 | 23.989 | 297.139 | 109.22 | 109.24 | 109.32 | 109.33 | 109.32 | 109.33 |  |  | 31 | 25.7 | 90.02
[11 empty] | 32 | 27.827 | 95.016
[11 empty] | 33 | 29.951 | 100.016
[11 empty] | 34 | 34.173 | 110.004
Dome Floor Sensors | [10 empty] | 35 | 38.372 | 120.009
[11 empty] | 36 | 42.539 | 130.007
[11 empty] | 37 | 46.677 | 140.005
[11 empty] | 38 | 50.791 | 150.004
[11 empty] | 39 | 54.881 | 160.004
[11 empty] | 40 | 58.95 | 170.005
[11 empty] | 41 | 63.008 | 180.023
[11 empty] | 42 | 67.04 | 190.02
[11 empty] | 43 | 71.06 | 200.027
[11 empty] | 44 | 75.066 | 210.038
[11 empty] | 45 | 79.053 | 220.036
[11 empty] | 46 | 83.027 | 230.039
[11 empty] | 47 | 86.988 | 240.038
[11 empty] | 48 | 90.931 | 250.029
[11 empty] | 49 | 94.87 | 260.039
[11 empty] | 50 | 98.793 | 270.036
[11 empty] | 51 | 102.703 | 280.035
[11 empty] | 52 | 106.593 | 290.01
[11 empty] | 53 | 110.489 | 300.033
[11 empty] | 54 | 114.362 | 310.031
[11 empty] | 55 | 116.297 | 315.046
[11 empty] | 56 | 118.225 | 320.053
[11 empty] | 57 | 120.525 | 326.026
[11 empty] | 58 | 122.07 | 330.041
[11 empty] | 59 | 0 | 0

## Telescope Sensor Location
Sensor ID | Sensor Type | Model No. | Location | Cable Length
TT1 | PT100 RTD | Omega RTD-805 | Hexapod Back - Air | 50'
TT2 | PT100 RTD | Omega RTD-830 | Secondary Truss Surface | 60'
TT3 | PT100 RTD | Omega RTD-830 | Secondary Truss Surface | 40'
TT4 | PT100 RTD | Omega RTD-805 | Secondary Truss Air | 40'
TT5 | PT100 RTD | Omega RTD-830 | Secondary Truss Surface | 30'
TT6 | PT100 RTD | Omega RTD-805 | Secondary Truss Air | 30'
TT7 | PT100 RTD | Omega RTD-830 | Secondary Truss Surface | 20'
TT8 | PT100 RTD | Omega RTD-805 | Secondary Truss Air | 20'
TM1 | PT100 RTD | RTD-3-F3102 | Mirror Surface | 20'
TM2 | PT100 RTD | RTD-3-F3102 | Mirror Surface | 22'
TM3 | PT100 RTD | RTD-3-F3102 | Mirror Surface | 12'
TM4 | PT100 RTD | RTD-3-F3102 | Mirror Surface | 8' + 7.75'
TM5 | PT100 RTD | RTD-3-F3102 | Mirror Surface | 8' + 8.2'
TM6 | PT100 RTD | RTD-3-F3102 | Mirror Surface | 8' + 6.25'
TMS1 | PT100 RTD |  | Mirror Supply - Air
TMS2 | PT100 RTD |  | Mirror Supply - Air
TT3 & TT4
TT5 & TT6
TT7 & TT8
TT1
TT2
N
TM2
TM1
TM3
TM4
TM5
TM6

## Dome Sensor Location
Sensor ID | Sensor Type | Model No. | Location
TE1 | Vaisala | Vaisala HMD60YO | External Dome Shutter
TE2 | Vaisala | Vaisala HMD60T | External Dome Shutter
HE1 | Vaisala | Vaisala HMD60YO | External Dome Shutter
HD1 | Johnson Cntl | TCS/1200 H8 HE-2 | Dome
HD2 | Johnson Cntl | TCS/1200 H8 HE-9 | Dome
HD3 | Johnson Cntl | TCS/1200 H8 HE-5 | Dome
TD1 | PT100 RTD | TE-6323P1 TE-2 (14) | Dome
TD2 | PT100 RTD | TE-6323P1 TE-7 (19) | Dome
TD3 | PT100 RTD | TE-6323P1 TE-9 (21) | Dome
TD4 | PT100 RTD | TE-6323P1 TE-3 (15) | Dome
TD5 | PT100 RTD | TE-6323P1 TE-6 (18) | Dome
TD6 | PT100 RTD | TE-6323P1 TE-8 (20) | Dome
TD7 | PT100 RTD | TE-6323P1 TE-1 (13) | Dome
TF1 | PT100 RTD |  | Floor Level - North
TF2 | PT100 RTD |  | Floor Level - South
TF3 | PT100 RTD |  | Floor Level - East
TF4 | PT100 RTD |  | Floor Level - West
TMS3 | PT100 RTD |  | Heat Exchanger Output
TF1
North Wall
L=90'
TF2
South Wall
L=67'
TF3
East Wall
L=127'
TF4
West Wall
L=20'
TD1
TE2
HD2
HE1
TE1
HD1
HD3
TD2
TD3
TD4
TD5
TD6
TD7
TE1, TE2, HE1
TD1, TD2, HD1
TD3, TD4, TD5, HD2, HD3
TD6, TD7
TF2
TF1
TF3
TF4
S
N
W
E

## Vaisala HMP243
HMP243 with Dewpoint Sensor Head
[6 empty] | HMP243 | G |  |  | 1 | [6 empty] | A |  |  |  | PRICE
Transmitter type | Td + x | [19 empty] | 1995
Transmitter cover | no display | [6 empty] | 1
| local display and keypad | [6 empty] | 2 | [12 empty] | 295
Dewpoint sensor | 2 m cable | [7 empty] | A
head cable length | 5 m cable | [7 empty] | B | [11 empty] | 140
| 10 m cable | [7 empty] | C | [11 empty] | 305
Temperature | no temperature sensor head | [8 empty] | 1
sensor head cable
length
Serial bus module | RS 232C | [9 empty] | A
| RS 485/RS 422 | [9 empty] | B | [9 empty] | 140
| digital current loop | [9 empty] | C | [9 empty] | 140
Sensor protection | 16452 (stainless steel sintered filter) | [10 empty] | 1 | [8 empty] | 50
| 16720 ( PPS Grid & stainless steel netting ) | [10 empty] | 2 | [8 empty] | 30
| 17230 ( PTFE membrane) | [10 empty] | 4 | [8 empty] | 25
Analog output | 4...20 mA | [11 empty] | A | A
signals | 0...20 mA | [11 empty] | B | B
(Ch1 and Ch2) | 0...1 V | [11 empty] | C | C
| 0...5 V | [11 empty] | D | D
| 0...10 V | [11 empty] | E | E
channel 1
channel 2
Parameters for the | Td | (-40...+100 °C) (-40 .. +212 °F) | [12 empty] | 3 | 3
analog outputs | x | (0 .. 500 g/kg d.a.) (0 .. 3500 gr/lb d.a.) | [12 empty] | 6 | 6
(Ch1 and Ch2) | [8 empty] | channel 1
channel 2
Temperature | [16 empty] | A
range
(measurement)
Units (local display | metric | [16 empty] | 1
and serial bus) | non-metric | [16 empty] | 2
Installation kit for | no | [17 empty] | A
duct mounting | yes | [17 empty] | B |  | 55
Accessories | No manual | [18 empty] | A
| English + PG 9 Cable Gland | [18 empty] | B
| English manual + NPT 1/2 Conduit Fitting | [18 empty] | P
| English manual + 6 Ft. A/C power cord | [18 empty] | R
| English manual + NPT 1/2 Conduit Fitting + 6 Ft. power cord | [18 empty] | S
TOTAL
| Td = Dewpoint Temperature | [18 empty] | QTY
| x = Mixing Ratio g/Kg | [18 empty] | TOTAL VALUE
Sections in bold are included in the basic price.
Items in italics are available at an additional charge.
Shaded sections indicate the factory default configuration.
Example of order code with typical settings;
|  |  | HMP 243 | G | 1 | A | 1 | A | 1 | A | A | 3 | 6 | A | 1 | A | 3 | B

## New Vaisala Quote
Click here
for HMP 243
Housing
New quote for Vaisala Dewpt/Humidity/Temperature Sensor
With local display and serial interface
Sensor Head has heater unit for external harsh environments