----- Original Message ----- 
From: "Alan Tokunaga" <tokunaga@ifa.hawaii.edu>
To: "Vern Stahlberger" <vern@ifa.hawaii.edu>
Cc: "Eric Tollestrup" <tolles@ifa.hawaii.edu>; "Douglas Toomey"
<toomey@ifa.hawaii.edu>
Sent: Thursday, February 26, 2004 8:11 PM
Subject: Re: Acquisition Camera Flexure: Conclusions & Next Steps


Alan Tokunaga wrote:

> Also, you mentioned that some screws were missing.  What is the
> likely effect of the missing screws on your results?  Should the
> flexure testing be done again (perhaps with the help of the
> day crew) in the daytime?


Hi

A) Note: For the first discussion below, all loads are applied in the E-W
direction;
deflections are measured in E-W direction at the application points of the
loads.
Refer to slides 4 and 5 of attached .ppt

A couple of measurements we've taken on the XY stage and George Mount
indicate
a relative difference in deflection of about .002 inch for the displacement.
These were
taken at about the same distance away from the pivot axis, which is located
where the
George Mount is attached to the AO Spool. For the applied load of 20 lbs in
the E-W
direction, the George Mount moved .0005 inch, whereas the XY stage moved
.0025 inch.
Hence, relative displacement between XY stage and George Mount is therefore
about .002 inch

We got about the same displacement  (.0025 inch) when measuring at the end
of the Acquisition
Box with a load of 20 lbs. A flexure of .0005 at the George Mounts
translates into a flexure of
.0024 inch (L^3  relationship) We are not supposed to see .0025 inch at the
XY stage.
One might argue that given the the .002 inch relative motion between George
Mount and XY stage,
we should at least see .002 inch (stage) +.0025 (George flexure at
Acquisition Box). But unless we apply
the force directly to the XY stage, by the time we are off-center, the
backlash doesn't show up anymore.
We are applying a torque to the XY stage, and in doing so bypass the
backlash.
This test would indicate we should check the anti-backlash of the XY stage.
The Acquisition Box itself is very stiff in this direction and even missing
or loose screws could
not affect the result in a measurable way. Also, if there were any loose
connections between Acquisition
Box and XY stage, we would get inconsistent duplicate test results. ( this
is in reference to the 4 screws
that connect the Acquisition Box to the XY stage)

How does all this relate to measurements taken by David G?        Even if we
discover that the XY stage's anti-backlash is not working right, it does not
explain the non-symmetrical
numbers that David G got, once with Nsfcam and a second time with Spex. It
is precisely this
non-symmetry that is hard to explain and makes me think something else is
happening, possibly
inside the Acquisition Camera Assembly. (for example, an optics that moves
away from its pads
when tilting in one direction, but not in the other, or, some such thing).
Or, could it be that whatever
backlash there is between the lead screw and the nut in the XY stage, the
motion is much
freer in one direction then one 180 degrees opposite? (much less of a
possibility, I think)

B) Note: For the second discussion, refer to slides 2 and 3 of .ppt
attached.

In this case, missing/loose screws could make a difference in the readings
we got. This is
because the Base Plate of the Acquisition Box depends on its stiffness on
the box type
structure of the design. The only way we could tell is by making sure that
we have the
full complement of screws in place and that they are sufficiently torqued
down and repeat
the measurements. This could be done by the daycrew and is quite easy by
just referring
to slide 2 of the .ppt. Apply a +/- force of 20 lbs  4 inch from the end of
the Acquisition Box's
Base Plate and measure the movement with an indicator the base of which is
attached to the
MIM Frame or the Modframe. We would be doing  this test  to see how much the
flexure is
reduced when all the bolts are in place and are torqued down properly.

The readings would need to be much smaller than .005 inch, otherwise we need
to assume
looseness in the XY stage. This looseness would not be a function of the
anti-backlash,
but of the stage itself. Since the stage's main components are made from the
same
material as the Acquisition Box, the problem is not due to unequal CTE's
(coeff.of thermal expansion)
For a 20 lbs concentrated load, the max displacement at the George Mount is
.00035 inch (FEA). This
translates into a flexure at the end of the Acquisition Box of less than
.002 inch. We measured
.005 inch. If the XY stage is rock solid, we should see about .002 inch due
to the flexure of
the George Mount at the end of the Acquisition Box. If we see more, it must
come either from
the Acquisition Box, or else from the XY stage. If the Acquisition Box is
properly bolted together
it will not move very much, probably in the 1/10000 inch range. So, we have
to wait
and see if the new readings are less. If they hover around .002 inch, we
just have bending
taking place, due to George Mount. Then we have to related this .002
displacement at the
Acquisition Camera to a rotation at the POM  (Pick Off Mirror), which turns
out to be about
15 arc-sec. A raytrace may be needed to evaluate these results. If, on the
other hand, we find
that the reading is still near .005 inch, this could be an indication that
the XY stage has some
looseness.

To test the anti-backlash in the XY stage's upper stage:
Apply a force of 20 lbs. This force is directed in the N-S direction and the
indicator's base is
attached to the George Mount, thereby getting only the relative motion
between XY stage
and George Mount.The force should be applied N-S and S-N to make sure the
anti-backlash
is tested on both directions and that it is the stages N-S travel direction.
Increase the force to
50 lbs, and see what happens. Make sure you apply the force the the part of
stage's translation your are
interested in, in this case again, N-S or S-N.

Alternative way of checking the Anti-backlash of the XY stage:
Remove Animatics Motor
Add a Turning knob to end of leadscrew
Measure if there is any rotational slack in the leadscrew before stage
moves. Use
either the built in Hall sensor or a dial indicator to verify this movement.

Vern