## BIAS2 (3)
BIAS VOLTAGES AS OF 3/20/98 USING some FOWLER VOLTAGES
|  |  |  | 1 LSB = | 0.0048828 | DAC INPUT BIT CODE | [11 empty] | DIP SWITCH POSITION
| DETECTOR BIAS |  |  | VOLTS | # OF LSB'S | 12 | 11 | 10 | 9 | 8 | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | JHX | JHXX
| VDDUC | VDET |  | 9.9951 | 2047 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF
| -3.7 | -3.375
signal
slot# | Board # | DAC# | *= set for programmable voltage, voltage shown is max negative
|  |  |  | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 0 | 0 | VDDCL_B_H | -4.5 | -922 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | ON | OFF | ON | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON
5 | 0 | 1 | VDDCL_B_L | -1.3 | -267 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | ON | OFF | OFF | ON | OFF | OFF | OFF | OFF | ON | OFF | ON | OFF
5 | 0 | 2 | PS1_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 0 | 3 | PS1_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
5 | 0 | 4 | PS2_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 0 | 5 | PS2_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
5 | 0 | 6 | PROE_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 0 | 7 | PROE_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
5 | 0 | 8 | PDES_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 0 | 9 | PDES_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
5 | 0 | 10 | PFS_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 0 | 11 | PFS_B_L | -4 | -820 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | ON | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON
5 | 0 | 12 | PF1_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 0 | 13 | PF1_B_L | -4 | -820 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | ON | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON
5 | 0 | 14 | PF2_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 0 | 15 | PF2_B_L | -4 | -820 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | ON | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON
6 | 1 | 0 | VGGCL_B_H* | -3.75 | -768 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ON | OFF | ON | OFF | ON | ON | ON | ON | ON | ON | ON | ON
6 | 1 | 1 | VGGCL_B_L | -4.5 | -922 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | ON | OFF | ON | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON
6 | 1 | 2 | VROWON_B_H | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
6 | 1 | 3 | VROWON_B_L | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
6 | 1 | 4 | VRSTR_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
6 | 1 | 5 | VRSTR_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
6 | 1 | 6 | VRSTG_B_H | -2 | -410 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON
6 | 1 | 7 | VRSTG_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
6 | 1 | 8 | VIREF_B | -1.8 | -369 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | ON | OFF | OFF | ON | OFF | ON | ON | ON | OFF | OFF | OFF | OFF
6 | 1 | 9 | VDDOUT_B | -1.3 | -267 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | ON | OFF | OFF | ON | OFF | OFF | OFF | OFF | ON | OFF | ON | OFF
6 | 1 | 10 | VDETCOM_B* | -3.75 | -768 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ON | OFF | ON | OFF | ON | ON | ON | ON | ON | ON | ON | ON
6 | 1 | 11 | VDDUC_B* | -3.75 | -768 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ON | OFF | ON | OFF | ON | ON | ON | ON | ON | ON | ON | ON
6 | 1 | 12 | VNCOL_B | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
6 | 1 | 13 | VNROW_B | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
6 | 1 | 14 | PSS_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
6 | 1 | 15 | PSS_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
13 | 0 | 0 | VDDCL_T_H | -4.5 | -922 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | ON | OFF | ON | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON
13 | 0 | 1 | VDDCL_T_L | -1.3 | -267 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | ON | OFF | OFF | ON | OFF | OFF | OFF | OFF | ON | OFF | ON | OFF
13 | 0 | 2 | PS1_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
13 | 0 | 3 | PS1_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
13 | 0 | 4 | PS2_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
13 | 0 | 5 | PS2_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
13 | 0 | 6 | PROE_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
13 | 0 | 7 | PROE_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
13 | 0 | 8 | PDES_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
13 | 0 | 9 | PDES_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
13 | 0 | 10 | PFS_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
13 | 0 | 11 | PFS_T_L | -4 | -820 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | ON | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON
13 | 0 | 12 | PF1_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
13 | 0 | 13 | PF1_T_L | -4 | -820 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | ON | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON
13 | 0 | 14 | PF2_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
13 | 0 | 15 | PF2_T_L | -4 | -820 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | ON | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON
14 | 1 | 0 | VGGCL_T_H* | -3.75 | -768 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ON | OFF | ON | OFF | ON | ON | ON | ON | ON | ON | ON | ON
14 | 1 | 1 | VGGCL_T_L | -4.5 | -922 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | ON | OFF | ON | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON
14 | 1 | 2 | VROWON_T_H | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
14 | 1 | 3 | VROWON_T_L | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
14 | 1 | 4 | VRSTR_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
14 | 1 | 5 | VRSTR_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
14 | 1 | 6 | VRSTG_T_H | -2 | -410 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON
14 | 1 | 7 | VRSTG_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
14 | 1 | 8 | VIREF_T | -1.8 | -369 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | ON | OFF | OFF | ON | OFF | ON | ON | ON | OFF | OFF | OFF | OFF
14 | 1 | 9 | VDDOUT_T | -1.3 | -267 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | ON | OFF | OFF | ON | OFF | OFF | OFF | OFF | ON | OFF | ON | OFF
14 | 1 | 10 | VDETCOM_T* | -3.75 | -768 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ON | OFF | ON | OFF | ON | ON | ON | ON | ON | ON | ON | ON
14 | 1 | 11 | VDDUC_T* | -3.75 | -768 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ON | OFF | ON | OFF | ON | ON | ON | ON | ON | ON | ON | ON
14 | 1 | 12 | VNCOL_T | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
14 | 1 | 13 | VNROW_T | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
14 | 1 | 14 | PSS_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
14 | 1 | 15 | PSS_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF

## PRTSEL2 (2)
COMPONENT SELECTION TABLES
D TO A OFFSET | Bits | Gain Drift | Vout | B zero | Iout | Tsettle | Isup +/- | INdigital | Pdiss | Cost | Package
| # | ppm/C | V | %FSR | mA | us | mA | ttl/cmos | mW |  | pins/size/suffix
ANA. DEV. AD667 | 12 | +/-30 | B 2.5/5/10 | .05/.1 | 5 +/- | 3/4us | 12/25 mA | TTL latched | 1W | 11.95 | 28/600mil/J-KN,A-BD
ANA. DEV. AD767 | 12 | +/-30 | B 2.5/5/10 | .05/.1 | 5 +/- | 3/4us | 13/23 mA | TTL latched | 600 | 9.85 | 28/300mil/J-KN,A-BD
BURR BRN DAC811 | 12 | +/-30 | B 5/10+10 | 10ppmFSR | 5 +/- | 3/4us | 15/35 mA | TTL latched | 800 | 20 | 28/600mil/J-KP,A-BH
BURR BRN DAC1201 | 12 | +/-30 | B 5/10+10 | 6ppmFSR | 5 +/- | 7us | 25/35 mA | TTL latched | 1W | 11.2 | 28/600mil/KPV
DATEL DAC-HK12B | 12 | +/-20 | B 2.5/5/10 | 5ppmFSR | 5 +/- | 3/4us | 10/25 mA | TTL latched | 700 |  | 24/600mil/G-MC-2,MM
DATEL DAC-HZ12B | 12 | +/-20 | B 2.5/5/10 | 0.05 | 1 | 3/4us | 10/16 mA | TTL nl | 390 |  | 24/600mil/G-MC,MM
PMI DAC-210 | 10+SIGN | +/-40 | B 10 | 0.05 | 5 | 1.5/-us | 9/12 mA | TTL nl | 300 |  | 18/300mil/E-F-GX
POWER BUFFER | Iout | Slew Rate | Bandwidth | Voffset | Rout | Noise out | Isup +/- |  |  | Cost | Package
| mA | V/us | Mhz | mV | Ohms | nV/[Hz]-1/2 | mA |  |  |  | pins/size/suffix
LINEAR TECH LT1010 | 150 +/- | 75 | 20 | 20/110 | 10 | 100to25 | 40max | BUFFER |  | 3.25 | 8/std/CN8
NATIONAL LM6361 | 25 | 300 | 50 | 22 |  | 15 | 6.9 | OPAMP
CLK LVL SWITCH | Vrange | Ron | Ton | Toff | Topen | Ids max | Isup +/- | IN digital | ID lk out | Cost | Package
Dual SPDT | V | Ohms | nsec | ns | ns | cont mA | mA | levels | nA |  | pins/size/suffix
ANA. DEV. AD741 | 5+ | 2 | 12 | 14 | [6 empty] | 14/std/KN
ANA. DEV. AD7512DI | 25+/- | 75/100max | > | > | 300typ | 50/150max | 0.8 | TTL | 1500max | 5.8 | 14/std/KN
ANA. DEV. AD7592DI | 25+/- | 60/120max | > | > | 450max | 50/150max | 1.5 | TTL latched | 400max | 4.95 | 14/std/KN
HARRIS HI303 | 15+/- | 35/50max | 210/300max | 160/250max | 60 BBM | ? | 1 | CMOS | 1/100max |  | 14/?/-5
MAXIM DG303A | 15+/- | 35/50max | 210/300max | 160/250max | 60 BBM | 30/100max | 0.1 | TTL | 0.25 | ? | 14/?/-5
MAXIM MXDG403 | 15+/- | 35 | 150 | 100 | 60 BBM | ? | 0.001 | CMOS | 1/100max | 2.78per1k | 14/?/-5
SILICONIX DG303AA
HARRIS HI201HS | 15+/- | 30/50max | 30/50max | 150 |  | 30/80max | 4.5 | TTL | 10 |  | 16/std/
DG613?
D TO A CLK LVL | Bits | Gain Drift | Vout | B zero | Iout | Tsettle | Isup +/- | INdigital | Pdiss | Cost | Package
| # | ppm/C | V | %FSR | mA | us | mA | ttl/cmos | mW |  | pins/size/suffix
ANA. DEV. AD8426 | 8 quad | +/-30 | B 10 | .05/.1 | 10 | 3/4us | 12/25 mA | TTL latched | ? | ? | 28/600mil/J-KN,A-BD
PMI DAC8229 | 8 dual | +/-40 | B 10 | 0.05 | 5 | 1.5/-us | 9/12 mA | TTL nl | 300 |  | 18/300mil/E-F-GX
AD839 too slow 20nsec?
AD811 opamp drvr?
LOW NOISE OPAMP | Vos | Drift Vos | Ib | Vnoise | Inoise | Tsettle | Gain stable | GBWP | SLEW | Vin diff | Iout | Notes
| uV | uV/C | nA | nv/rtHz(1k) | pA/rtHz | nsec(.1%) | Min | MHz | V/usmin | V
Lin Tech LT1028AC | 40 | 0.8 | 90 | .85/1.1 | 1.0/1.6 |  | 1* | 75 | 11/15 typ | 1.8 +/- 25ma
Lin Tech LT1128AC | 40 | 0.8 |  |  |  |  | 1 | 20 | 4.5/6 typ | 1.8 +/- 25ma
OP27E | 25 | 0.6 | 40 | 3.8 | .4/.6 |  | 1 | 5/8 typ | 1.7/2.8 typ | .7 +/-
Lin Tech LT1055A | 150 | 4 | 0.05 | 14/20 | .0018/.004 | 1200 | 1 | 5 | 10/13 yp | 40 +/-
HARRIS HA7-5127 | 100 | 0.6 | 80 | 3.2/4.5 | .4/.6 | 1500 | 1 | 8.5 | 7/10 typ | .7 +/- 25ma
HARRIS HA7-5127A | 25 | 0.2 | 40 | 3.0/3.8 | .4/.6 | 1500 | 1 | 8.5 | 7/10 typ | .7 +/- 25ma
HARRIS HA7-5137 | 100 | 0.4 | 80 | 3.2/4.5 | .4/.6 | 1000 | 3 | 63 | 14/20 typ | .7 +/- 25ma
HARRIS HA7-5137A | 25 | 0.2 | 40 | 3.0/3.8 | .4/.6 | 1000 | 3 | 63 | 14/20 typ | .7 +/- 25ma
HARRIS HA7-5147 | 100 | 0.4 | 80 | 3.2/4.5 | .4/.6 | 400 | 10 | 140 | 28/35 typ | .7 +/- 25ma
HARRIS HA7-5147A | 25 | 0.2 | 40 | 3.0/3.8 | .4/.6 | 400 | 10 | 120 | 28/35 typ | .7 +/- 25ma
Ana Dev AD846AN or Q | 200 | .8/5.5 | 450 | 2 | 20 | 80 | 1 | 450 | 450typ | 1 +/-
Ana Dev AD829 | 200 | 0.3 | 3.3 | 1.7 | 1.5 | 65 | 1 | 600 | 230 | 6 +/- | ?
Ana Dev AD841 | 1mV | 35 | 3.5uA | 15 | 1.5 | 65 | 1 | 600 | 230 | ? | ?
Ana Dev AD797BN | 40 | 0.6 | 250 | 0.9 | 2 | 800(.001%) | 1 | 110 | 20 | .7 +/- 25ma | 50mA (200K)
Burr Brown OPA627 Difet | 40 | 0.4 | 1 | 4.5 | 0.0015 | 550(.01%) | 1 | 16 | 40 | ? | 45mA (1K)
Burr Brown OPA642 | +/-1.5mV | 4 | 18uA | 2.5 | 2.4 | 11.5 | 1 | 450 | 380 | ? | 40mA | +/-5V rails
opa655 jfet
Ana Dev AD745
Ana Dev AD8138
Ana Dev AD8057 | ? | ? | ? | 7 | ? | ? | 1 | 325 | 1000V/us | ? | 1Kohmsok | 3-12V rail+
OPA627
AD8057
MAX427 | 15 | 0.8 |  | 3.8 |  |  | 1 | 60 | 15
Nat Semi CLC580 | ? | ? | ? | 4 | ? | 18 | 1 | 140 | 450 |  | 80mA
Current Source | Vbrgss | IGSS | VGS | IDSS | VDSon | Rdson | Ciss | note | note | cost | Package
JFET | Vdc | nAdc | Vdc | mAdc | Vdc | Ohms | pF
2N4393 JFET ndepl | 40min | .2max | -0.5min/-3max | 5min/30max | 0.4 | 100max | 14max | MPF4391 |  |  | TO-18=TO-206AA
MBF4393L | 30min | .2max | -0.5min/-3max | 5min/30max | 0.4 | 100max | 14max |  |  |  | SOT-23
MPF4393 | 30min | .2max | -0.5min/-3max | 5min/30max | 1.4 | 100max | 14max |  |  |  | TO-92
MNBF5486L | -25min | -1max | -2min/-6max | 8min/20max |  |  | 5max |  |  |  | SOT-23
16BIT ADC | Freq | Ana Input | Input R | Tconvert | Ts&h | S/N 100khz | S/N 500khz | S/N 900khz | THD 100Khz | ref stab | Integral nonlin | Diff nonIin | Stability | Cost | Size | Power
Analogic ADC4322 | 2Mhz | +/-2.5/5/10 | 750/1.5k | 300ns | 200ns | 86dBmin | 76dBmin | 75dBmin | -86dBmax | 15ppm/C | +/- .003% | +/- .75LSB | +/-15ppm/C | 1080 | 2.4"x1.5" | 2.3W
Datel ADS-932MC | 2Mhz | +/-2.75 | 500 | ? | 200ns | 81dBmin | 81dBmin | 80dBmin | -88dBmax | +/-30ppm/C | +/- 1LSB | +1LSB | +/-15ppm/C | 845 | 2.4"x1.5" | 1.85W
Linear Tech LT1604 | 333Khz | +/-2.55 | ? | ? | ? | 90dBmin |  |  | -100dBmax | +/-15ppm/C |  |  |  |  | 36 SSOP | 0.2W
Analog Devices 9260
AKM semi
146BIT ADC | Freq | Ana Input | Input R | Tconvert | Ts&h | S/N 100khz | S/N 500khz | S/N 900khz | THD 100Khz | ref stab | Integral nonlin | Diff nonIin | Stability | Cost | Size | Power
Linear Tech LT1604 | 2.2Mhz | +/-2.5/5/10 | 750/1.5k | 300ns | 200ns | 80dBmin | [8 empty] | 28 SSOP | 175mW+/-5V
Analog Devices 6644
support ADC's | Freq | Ana Input | Input R | Tconvert | Ts&h | S/N 100khz | S/N 500khz | #inputs | THD 100Khz | ref stab | Integral nonlin | Diff nonIin | Stability | Cost | Size | Power
JFET: MAXIM 125-14bit; Vdc: 333KHz; nAdc: +/-5V; mAdc: 3uS | 8channel | [6 empty] | 36SSOP
POWER SUPPLIES | VOLTS | AMPS | WATTS | Vrip RMS | Vrip pp | Load Reg | Line Reg | OVP/OCP | LIN/PWM | Cost | LxWxHxlbs | Display | RS232/on-off | rackmount | rear conn. | parallel
Acopian A5MT900 | 5 | 9 |  | 0.25mV |  | 0.005% | 0.005% | yes/yes | LIN | 250 |  | No
Acopian A15MT550 | 15 | 5.5 |  | 0.25mV |  | 0.005% | 0.005% | yes/yes | LIN | 245 |  | no
Amrel PD8-7A | 0-8V | 7 |  | 0.2mV | 2mV | 0.001% | 0.001% | yes/yes | LIN | 995 | 1/4 rack , 7inch | yes | yes | yes | yes? | ?
Amrel PD20-10AX | 0-20V | 10 |  | 0.2mV | 2mV | 0.001% | 0.001% | yes/yes | LIN | 1895 | 1/2 rack , 7inch | yes | yes | yes | yes? | ?
Kepco ABC 10-10DM | 0-10V | 10 | 100 | 2mV | 3mV | 0.01% | 0.01% | yes/yes | 100KHz | 995 | 2 per rack | yes | GPIB | yes | yes | no need
Kepco ABC 15-7DM | 0-15V | 7 | 105 | 6mV | 6mV | 0.01% | 0.01% | yes/yes | 100KHz | 995 | 2 per rack | yes | GPIB | yes | yes | no need
Kepco MAT 6-32 | 0-6V | 32 | 360 | 0.3mV | 3mV | 0.002% | 0.001% | yes/yes | LIN | 1895 | 3per rack | yes | TMA4882-27 | yes | mil | no need
Kepco MAT 15-20 | 0-15V | 20 | 360 | 0.3mV | 3mV | 0.002% | 0.001% | yes/yes | LIN | 1895 | 3per rack | yes | TMA4882-27 | yes | mil | no need
Kepco MBT 6-32MG | 0-6V | 32 | 360 | 0.3mV | 3mV | 0.002% | 0.001% | yes/yes | LIN | 2499 | 2per RA-37 | yes | SBC | ? | yes | no need
Kepco MBT 6-32M | 0-6V | 32 | 360 | 0.3mV | 3mV | 0.002% | 0.001% | yes/yes | LIN | 1889 | 2per RA-38 | yes | no | ? | yes | no need
Kepco MBT 15-20M | 0-15V | 20 | 360 | 0.3mV | 3mV | 0.002% | 0.001% | yes/yes | LIN | 1889 | 2per RA-39 | yes | no | ? | yes | no need
Kepco MST 6-20M | 0-6V | 20 |  | 2.5mV | 50mV | 0.002% | 0.001% | yes/yes | ? | 1695 | 8per rack | yes | yes | yes | yes | no need
Xantrex XT7-6 | 0-7V | 6 | 43 | <1mV |  | 0.01% | 0.01% | yes/yes | LIN | 665 | 4 per rack | Yes |  |  |  | yes
Xantrex XT15-4 | 0-15V | 4 | 60 | <1mV |  | 0.01% | 0.01% | yes/yes | LIN | 665 | 4 per rack | Yes |  |  |  | yes
Xantrex HPD15-20 | 0-15V | 20 | 300 | <5mV | <100mV | 0.01% | 0.01% | yes/yes | PWM |  | 4 per rack | yes |  |  |  | yes
HP 6641A | 0-8V | 20 | 200 | <300uV | 3mV | 0.01% | 0.01% | yes/yes | LIN | 2500 | 1 (16.75X3.5X17.3) | yes | HPIB | yes | yes | no need
HP 6642A | 0-20V | 10 | 200 | <300uV | 3mV | 0.01% | 0.01% | yes/yes | LIN | 2500 | 1 (16.75X3.5X17.3) | yes | HPIB | yes | yes | no need
HP E3632A | 0-15V | 7 | 105/120 | <350uV | 2mV | 0.01% | 0.01% | yes/yes | LIN | 995 | ? | yes | rs232
HP E3614A | 0-8V | 6 | 60 | <200uV | 1mV | 0.01% | 0.01% | yes/yes | LIN | 525 |  | yes |  |  | yes | yes
HP E3615A | 0-20V | 3 | 60 | <200uV | 1mV | 0.01% | 0.01% | yes/yes | LIN | 526 |  | yes |  |  | yes | yes
Sorensen SRL 10-25 | 0-10V | 25 |  | 350uV | 10mV | 0.01% | 0.01% | yes/yes | LIN | 2645 | 1 per rack | yes | 200 ohm | yes | yes | no need
Sorensen SRL20-12 | 0-20V | 12 |  | 200uV | 20mV | 0.01% | 0.01% | yes/yes | LIN | 2530 | 1 per rack | yes | 200 ohm | yes | yes | no need
JFET: Vicor |  | PWM |  |  | no
JFET ( Isource=100uA) | Idss mA | Vds max | Vgs=.1mA | Vnoise 10hz | Vn 100Khz | Vgs(off) | Isup +/- | siliconix | package | Cost | Vgs=.1mA@15Vds
2N4338 NJFET | .2min/.6max | 50V | -0.2v/vds15V | 22nV/hz-.5 | 5nV/hz-.5 | -.3/-1max | mA | NPA | TO-18 |  | -0.2 to -0.5V
C
Comparators | Tr | Vos | Ib | CM input- | CM input+ | Vdiff abs | Vsupplies | note | DGND pin? | OC out | Package
| ns | mV | uA | V | V | V | V
AD790K | 35 | 0.3 | 3 | -18 | 18 | 15 | 3.5/18 | fast | yes | yes | single
LF311 | 200 | 3 | 0.1 | V+-30min | V-+30max | 30 | 5/30+- | oscillates? | yes | yes
LM311 | 200 | 4 | 5e-05 | V+-30min | V-+30max | 30 | 5/30+- | oscillates | yes | yes | dual2311
LP311 | 2000 | 7.5 | 0.1 | V- | V-+30max | 30 | 5/30+- | oscillates? | yes | yes
LM319 | 80 | 4 | 0.5 | v- | V+ | 36 |  |  | NO | yes
LM360 | 14 | 5 | 20 | v- | V+ | 5 | 4.5/6.5+ - |  | yes | yes | +5 &-5V operation
LT1011 | 150 | 0.5 | 0.025 | V--0.6 | V++0.6 | 36 |  |  | NO | yes
LT1040 | 20000 | 1 | 0.015 | V-+0.3 | V+-0.3 | ? | 2.8/16 |  | yes | yes
EL2018C | 20 | 3 | 0.3 | V- | V+ | 36 | 4.5/18 + - |  | yes | yes | elantec/single
CA3290a | 1000 | 10 | 4e-05 | v--5 | v++5 | 36 | 4.5/18 +only |  | yes | no
JFET: MAXIM MAX4374 |  |  |  | 14SO
Fiber replacement | Gbps | wavelength | interface | Class1? | distance | Vsupply | Wpower |  | size | cost | website | tel | notes
ONM10PHY-SR | 10Gbps | 1310nm | 16bit LVDS | ? | 2km |  |  |  | 2.5x3.5x0.5 | 5000 | networkelements.com | 503-644-7666 | all in one part
TI TLK2201 G-Ethernet | 1.25Gbps | Cu |  |  |  | 2.5V | 200mW |  |  | 4.5 | ti.com/sc
JFET: TI TLK1501-3101; Vdc: 1.6-3.125Gbps |  |  |  | ti.com/sc

## POWERSYS
POWER DISSIPATION | [8 empty] | NEW SYSTEM ESTIMATE:POWER DISSIPATION
| PER BOARD (AMPS) |  |  | BD/SYS | PER SYSTEM (AMPS) |  |  |  |  | PER BOARD (AMPS) |  |  | BD/SYS | PER SYSTEM (AMPS)
| [+5V] | [+15V] | [-15V] |  | [+5V] | [+15V] | [-15V] |  |  | [+5V] | [+15V] | [-15V] |  | [+5V] | [+15V] | [-15V]
ADC1614 | 0.275 | 0.2 | 0.2 | 8 | 2.2 | 1.6 | 1.6 |  | ADC8 | 0.696 | 0.568 | 0.488 | 5 | 3.48 | 2.84 | 2.44
CKBS | 0.209 | 0.483 | 0.483 | 3 | 0.627 | 1.449 | 1.449 |  | CKBS | 0.209 | 0.483 | 0.483 | 6 | 1.254 | 2.898 | 2.898
OPTO24 | 0.234 | 0 | 0 | 3 | 0.702 | 0 | 0 |  | FCVME | 1.8936 | 0 | 0 | 3 | 5.6809 | 0 | 0
PREAMP4 | 0.106 | 0.2 | 0.2 | 4 | 0.424 | 0.8 | 0.8 |  | PREAMP8 | 0.12 | 0.4 | 0.4 | 5 | 0.6 | 2 | 2
RDB1 | 0.05 | 0 | 0 | 2 | 0.1 | 0 | 0 |  | BKPLN7 | 0.4 | 0 | 0 | 7 | 2.8 | 0 | 0
BKPLN1 | 0.3125 | 0 | 0 | 3 | 0.9375 | 0 | 0 | [5 empty] | AMPS | 13.8149 | 7.738 | 7.338
|  |  |  | AMPS | 4.9905 | 3.849 | 3.849 | [5 empty] | WATTS | 69.0745 | 116.07 | 110.07 | 295.2145 | TOTAL_CRYO
|  |  |  | WATTS | 24.9525 | 57.735 | 57.735 | 140.4225
HSDSP | PC POWER | [6 empty] | TOTAL P | [5 empty] | WATTS | WATTS | WATTS
[9 empty] | LINEAR SUPPLY |  | eff. | 50% |  | 138.1491 | 232.14 | 220.14 | 590.4291 | TOTAL_LINEAR
| PER BOARD (AMPS) |  |  | BD/SYS | PER SYSTEM (AMPS) |  |  |  | SWITCHING SUPPLY |  | eff. | 80% |  | 86.3432 | 145.0875 | 137.5875 | 369.0182 | TOTAL_SWTICHER
| [+5V] | [+15V] | [-15V] |  | [+5V] | [+15V] | [-15V]
CKBS | 0.209 | 0.483 | 0.483 | 2 | 0.418 | 0.966 | 0.966
OPTO24 | 0.234 | 0 | 0 | 0 | 0 | 0 | 0
PREAMP4 | 0.106 | 0.2 | 0.2 | 1 | 0.106 | 0.2 | 0.2 |  |  | VREGdrop= | 2 | CABLE_FT= | 100
RDB1 | 0.05 | 0 | 0 | 0 | 0 | 0 | 0 |  | CABLE |  |  |  |  | WATTS | WATTS | WATTS | WATTS
BKPLN1 | 0.3125 | 0 | 0 | 2 | 0.625 | 0 | 0 |  | GAUGE | ohms/ft | 5Vsupply | 15Vsupply | 15Vsupply | +5Vcable_W | +15Vcable_W | -15Vcable_W | CBL_TOTAL
|  |  |  | AMPS | 1.149 | 1.166 | 1.166 |  | AWG14 | 0.00252 | #VALUE! | #VALUE! | 18.8492 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
|  |  |  | WATTS | 5.745 | 17.49 | 17.49 | 40.725 | AWG16 | 0.00402 | #VALUE! | #VALUE! | 19.9499 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
HSDSP | PC POWER | [6 empty] | TOTAL P | AWG18 | 0.00638 | #VALUE! | #VALUE! | 21.6816 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[9 empty] | AWG20 | 0.0101 | #VALUE! | #VALUE! | 24.4114 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[10 empty] | LINEAR | SWITCH
FIBER | 0.85 | 0 | 0 | 1 | 0.85 | 0 | 0 |  | W/CABLE | NO CABLE | NO CABLE
[9 empty] | DOME | DOME | DOME
[9 empty] | #VALUE! | #VALUE! | #VALUE!
#VALUE!
#VALUE!
#VALUE!
[10 empty] | PER BOARD (AMPS) |  |  | BD/SYS | PER SYSTEM (AMPS)
[10 empty] | [+5V] | [+12V] | [-12V] |  | [+5V] | [+12V] | [-12V]
[9 empty] | SPARC5 | 2.5 | 0 | 0 | 1 | 2.5 | 0 | 0
[9 empty] | IXZ4444 | 2.5 | 0 | 0 | 2 | 5 | 0 | 0
[9 empty] | FCVME | 1.8936 | 0 | 0 | 2 | 3.7873 | 0 | 0
[9 empty] | SCSI | 1 | 0 | 0 | 1 | 1 | 0 | 0
[9 empty] | VMEBKPLN | 1.6 | 0 | 0 | 1 | 1.6 | 0 | 0
[13 empty] | AMPS | 13.8873 | 0 | 0
[13 empty] | WATTS | 69.4364 | 0 | 0 | 69.4364
TOTAL P

## PRTSEL2
COMPONENT SELECTION TABLES
D TO A OFFSET | Bits | Gain Drift | Vout | B zero | Iout | Tsettle | Isup +/- | INdigital | Pdiss | Cost | Package
| # | ppm/C | V | %FSR | mA | us | mA | ttl/cmos | mW |  | pins/size/suffix
ANA. DEV. AD667 | 12 | +/-30 | B 2.5/5/10 | .05/.1 | 5 +/- | 3/4us | 12/25 mA | TTL latched | 1W | 11.95 | 28/600mil/J-KN,A-BD
ANA. DEV. AD767 | 12 | +/-30 | B 2.5/5/10 | .05/.1 | 5 +/- | 3/4us | 13/23 mA | TTL latched | 600 | 9.85 | 28/300mil/J-KN,A-BD
BURR BRN DAC811 | 12 | +/-30 | B 5/10+10 | 10ppmFSR | 5 +/- | 3/4us | 15/35 mA | TTL latched | 800 | 20 | 28/600mil/J-KP,A-BH
BURR BRN DAC1201 | 12 | +/-30 | B 5/10+10 | 6ppmFSR | 5 +/- | 7us | 25/35 mA | TTL latched | 1W | 11.2 | 28/600mil/KPV
DATEL DAC-HK12B | 12 | +/-20 | B 2.5/5/10 | 5ppmFSR | 5 +/- | 3/4us | 10/25 mA | TTL latched | 700 |  | 24/600mil/G-MC-2,MM
DATEL DAC-HZ12B | 12 | +/-20 | B 2.5/5/10 | 0.05 | 1 | 3/4us | 10/16 mA | TTL nl | 390 |  | 24/600mil/G-MC,MM
PMI DAC-210 | 10+SIGN | +/-40 | B 10 | 0.05 | 5 | 1.5/-us | 9/12 mA | TTL nl | 300 |  | 18/300mil/E-F-GX
POWER BUFFER | Iout | Slew Rate | Bandwidth | Voffset | Rout | Noise out | Isup +/- |  |  | Cost | Package
| mA | V/us | Mhz | mV | Ohms | nV/[Hz]-1/2 | mA |  |  |  | pins/size/suffix
LINEAR TECH LT1010 | 150 +/- | 75 | 20 | 20/110 | 10 | 100to25 | 40max | BUFFER |  | 3.25 | 8/std/CN8
NATIONAL LM6361 | 25 | 300 | 50 | 22 |  | 15 | 6.9 | OPAMP
CLK LVL SWITCH | Vrange | Ron | Ton | Toff | Topen | Ids max | Isup +/- | IN digital | ID lk out | Cost | Package
Dual SPDT | V | Ohms | nsec | ns | ns | cont mA | mA | levels | nA |  | pins/size/suffix
ANA. DEV. AD7512DI | 25+/- | 75/100max | > | > | 300typ | 50/150max | 0.8 | TTL | 1500max | 5.8 | 14/std/KN
ANA. DEV. AD7592DI | 25+/- | 60/120max | > | > | 450max | 50/150max | 1.5 | TTL latched | 400max | 4.95 | 14/std/KN
HARRIS HI303 | 15+/- | 35/50max | 210/300max | 160/250max | 60 BBM | ? | 1 | CMOS | 1/100max |  | 14/?/-5
MAXIM DG303A | 15+/- | 35/50max | 210/300max | 160/250max | 60 BBM | 30/100max | 0.1 | TTL | 0.25 | ? | 14/?/-5
MAXIM MXDG403 | 15+/- | 35 | 150 | 100 | 60 BBM | ? | 0.001 | CMOS | 1/100max | 2.78per1k | 14/?/-5
SILICONIX DG303AA
HARRIS HI201HS | 15+/- | 30/50max | 30/50max | 150 |  | 30/80max | 4.5 | TTL | 10 |  | 16/std/
D TO A CLK LVL | Bits | Gain Drift | Vout | B zero | Iout | Tsettle | Isup +/- | INdigital | Pdiss | Cost | Package
| # | ppm/C | V | %FSR | mA | us | mA | ttl/cmos | mW |  | pins/size/suffix
ANA. DEV. AD8426 | 8 quad | +/-30 | B 10 | .05/.1 | 10 | 3/4us | 12/25 mA | TTL latched | ? | ? | 28/600mil/J-KN,A-BD
PMI DAC8229 | 8 dual | +/-40 | B 10 | 0.05 | 5 | 1.5/-us | 9/12 mA | TTL nl | 300 |  | 18/300mil/E-F-GX
LOW NOISE OPAMP | Vos | Drift Vos | Ib | Vnoise | Inoise | Tsettle | Gain stable | GBWP | SLEW | Vin diff | Iout | Notes
| uV | uV/C | nA | nv/rtHz(1k) | pA/rtHz | nsec(.1%) | Min | MHz | V/usmin | V
Lin Tech LT1028AC | 40 | 0.8 | 90 | .85/1.1 | 1.0/1.6 |  | 1* | 75 | 11/15 typ | 1.8 +/- 25ma
Lin Tech LT1128AC | 40 | 0.8 |  |  |  |  | 1 | 20 | 4.5/6 typ | 1.8 +/- 25ma
OP27E | 25 | 0.6 | 40 | 3.8 | .4/.6 |  | 1 | 5/8 typ | 1.7/2.8 typ | .7 +/-
Lin Tech LT1055A | 150 | 4 | 0.05 | 14/20 | .0018/.004 | 1200 | 1 | 5 | 10/13 yp | 40 +/-
HARRIS HA7-5127 | 100 | 0.6 | 80 | 3.2/4.5 | .4/.6 | 1500 | 1 | 8.5 | 7/10 typ | .7 +/- 25ma
HARRIS HA7-5127A | 25 | 0.2 | 40 | 3.0/3.8 | .4/.6 | 1500 | 1 | 8.5 | 7/10 typ | .7 +/- 25ma
HARRIS HA7-5137 | 100 | 0.4 | 80 | 3.2/4.5 | .4/.6 | 1000 | 3 | 63 | 14/20 typ | .7 +/- 25ma
HARRIS HA7-5137A | 25 | 0.2 | 40 | 3.0/3.8 | .4/.6 | 1000 | 3 | 63 | 14/20 typ | .7 +/- 25ma
HARRIS HA7-5147 | 100 | 0.4 | 80 | 3.2/4.5 | .4/.6 | 400 | 10 | 140 | 28/35 typ | .7 +/- 25ma
HARRIS HA7-5147A | 25 | 0.2 | 40 | 3.0/3.8 | .4/.6 | 400 | 10 | 120 | 28/35 typ | .7 +/- 25ma
Ana Dev AD846AN or Q | 200 | .8/5.5 | 450 | 2 | 20 | 80 | 1 | 450 | 450typ | 1 +/-
Ana Dev AD829 | 200 | 0.3 | 3.3 | 1.7 | 1.5 | 65 | 1 | 600 | 230 | ? | ?
Ana Dev AD841 | 1mV | 35 | 3.5uA | 15 | 1.5 | 65 | 1 | 600 | 230 | ? | ?
Ana Dev AD797BN | 40 | 0.6 | 250 | 0.9 | 2 | 800(.001%) | 1 | 110 | 20 | .7 +/- 25ma | 50mA (200K)
Burr Brown OPA627 Difet | 40 | 0.4 | 1 | 4.5 | 0.0015 | 550(.01%) | 1 | 16 | 40 | ? | 45mA (1K)
Burr Brown OPA642 | +/-1.5mV | 4 | 18uA | 2.5 | 2.4 | 11.5 | 1 | 450 | 380 | ? | 40mA | +/-5V rails
opa655 jfet
Ana Dev AD745
OPA627
16BIT ADC | Freq | Ana Input | Input R | Tconvert | Ts&h | S/N 100khz | S/N 500khz | S/N 900khz | THD 100Khz | ref stab | Integral nonlin | Diff nonIin | Stability | Cost | Size | Power
Analogic ADC4322 | 2Mhz | +/-2.5/5/10 | 750/1.5k | 300ns | 200ns | 86dBmin | 76dBmin | 75dBmin | -86dBmax | 15ppm/C | +/- .003% | +/- .75LSB | +/-15ppm/C | 1080 | 2.4"x1.5" | 2.3W
Datel ADS-932MC | 2Mhz | +/-2.75 | 500 | ? | 200ns | 81dBmin | 81dBmin | 80dBmin | -88dBmax | +/-30ppm/C | +/- 1LSB | +1LSB | +/-15ppm/C | 845 | 2.4"x1.5" | 1.85W

## Sheet1
SPEX science grade array selection
| SBRC | Fowler | SBRC | Fowler
| SCA411729 | SCA411729 | SCA47672 | SCA47672
Readout die | 206-001-11-E01 | 206-001-11-E01 | 152/5/2/B1 | 152/5/2/B1
Detector die | NI960927-C02 | NI960927-C02 | NI960927-238-C2 | NI960927-238
Doping (cm-3) |  |  | 330000000000000 | 330000000000000
Temp | 35 | 35 | 35 | 35
Idark
Dark Current,mean,e/sec,32s-10s# | 1.87 |  | 151.3
Dark Current,s.e/sec | 8.91 |  | 65.8
Dark Current,mean,e/sec,256s-32s |  | 3.5 |  | 0.42
Dark Current,median.e/sec |  | 1.65 |  | 0.24
Detector Characteristics
Zero Bias point, V | -3.9 |  | -3.95
VdetCom Optimized Value,V | -3.4
Vdet sweep Linear V |  | 0.6
Vdet sweep pedestal V |  | 0.1
"Full Well", 400mV,nonlinear |  | 168000 |  | 145000
"Full Well", 600mV,nonlinear |  | 256000 |  | 225000
"Full Well", 800mV,nonlinear |  | 348000 |  | 308000
Capacitance,F | 61 | 72.6 |  | 64
Noise
Noise,Rochester,V | 239 |  | 424
Noise,Rochester,e | 121 |  | 177.9
Noise,Fowler,e(60Khz,400mV) |  | 35 |  | 44
Residual Image
1st read,%* |  | 1.3% |  | 1.9%
1st read,e* |  | 1690 |  | 1900
3rd read,%* |  | 0.1% |  | 0.2%
3rd read,e* |  | 130 |  | 200
Biases
VDDCL | -3.8 | -3.54 | -3.8 | -3.54
VDDUC | -4 | -3.5 | -4 | -4.108
VGGCL | 0 | -2.768 | -1.5 | -2.064
VdetCom | -3.4 | -3.097 | -3.5 | -4.101
VDDOUT | -1.5 | -0.97 | -1.5 | 0.969
Iidle/slew | -2.5 | -1.79 | -2.5 | -1.79
VNCOL | -6 | -3.979 | -6 | -3.977
VNROW | -6 | -6.21 | -6 | -6.21
Anomalies | q4 scratch | q4 scratch |  | q3-q4 crack
Gain
gain.V/V, reset stuck on | 0.755 | 0.75 | 0.744 | 0.77
Power | 8.744 |  | 6.39
# Itime=4sec
* Test parameters different 130Ke vs 100Ke

## ALLADDIN 2 PINOUT
ALLADIN 2 pinout (SB0152 ROIC PCHANNEL OUTPUTS) | [5 empty] | 700-150 | 700-151
socket Yamaichi IC51-1244-410-1 | [5 empty] | Array Fanout | Config bd |  | 61 Pin Hermetic Connector Bd.
pin num | signal name | high level | low level | function |  |  |  |  | Pin # | J1 | J2 | signal name | high level | low level | function |  | test
1 | Q1_VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk |  |  |  | c |  | 3 | SUBSTRATES
2 | Q1_VDDCL | -1.2 | VDDUC | column line overload protection | clk | [6 empty] | VSUB | 0 gnd |  | readout substrate | bias | 0
3 | Q1_VDDUC | -4.0 to -3.5 |  | unit cell drain | bias | [6 empty] | VDETCOM | -2.8 to-3.4 |  | detector common supply/substrate | bias | 0
4 | Q1_PHISS | VpROW&COL | VnROW | starts slow (row) shift register | clk | [6 empty] | UNIT CELL
5 | Q1_PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk | [6 empty] | VDDUC | -4.0 to -3.5 |  | unit cell drain ~8mA | bias | -4
6 | Q1_PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk | [6 empty] | VSS | 0 |  | unit cell source | bias | 0
7 | Q1_PHIROE | VpROW&COL | VnROW | odd/even clock | clk | [6 empty] | VRSTG | -3.5 | -5.6 | global reset | aclk | -3.75
8 | Q1_PHIDES | VpROW&COL | VnROW | row deselect | clk | [6 empty] | VRSTR | -3.5 | -5.6 | row reset | aclk | -6
9 | Q1_PHIFS | VpROW&COL | VnCOL | starts fast (col) shift register | clk | [6 empty] | VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position |  | 0
10 | Q1_PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk | [6 empty] | VROWON | -6.0 to -5.6 |  | unit cell row enable on position | aclk | -6
11 | Q1_PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk | [6 empty] | OUTPUTS
12 | Q1_VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias | [6 empty] | OUT# |  |  | 100 to 300uA I load
13 | Q1_VnCOL | -6.0 to -3.5 |  | negative supply for fast shift reg | bias | [6 empty] | VDDOUT | -1.0 to -1.3 |  | output drain ~16mA |  | -1.3
14 | Q1_Vp | 0.0 to -0.5 |  | positive supply for shift reg | bias | [6 empty] | CURRENT SOURCES
15 | VSS | 0 |  | unit cell source | bias | [6 empty] | IREF* | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource | 250K
16 | VSUB | 0 gnd |  | readout substrate | bias | [6 empty] | SHIFT REGISTERS
17 | VSS | 0 |  | unit cell source | bias | [6 empty] | Vp | 0.0 to -0.5 |  | positive supply for shift reg | bias | 0
18 | Q3_Vp | 0.0 to -0.5 |  | positive supply for shift reg | bias | [6 empty] | VnCOL | -6.0 to -3.5 |  | negative supply for fast shift reg | bias | -6
19 | Q3_VnCOL | -6.0 to -3.5 |  | negative supply for fast shift reg | bias | [6 empty] | VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias | -6
20 | Q3_VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias | [6 empty] | DIAGNOSTICS
21 | Q3_PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk | [6 empty] | TEND |  |  | 10Kohm to gnd shift reg diag |  | -5
22 | Q3_PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk | [6 empty] | CLOCKS
23 | Q3_PHIFS | VpROW&COL | VnCOL | starts fast (col) shift register | clk | [6 empty] | PHISS | VpROW&COL | VnROW | starts slow (row) shift register | clk | 0/-6
24 | Q3_PHIDES | VpROW&COL | VnROW | row deselect | clk | [6 empty] | PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk | 0/-6
25 | Q3_PHIROE | VpROW&COL | VnROW | odd/even clock | clk | [6 empty] | PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk | 0/-6
26 | Q3_PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk | [6 empty] | PHIROE | VpROW&COL | VnROW | odd/even clock | clk | 0/-6
27 | Q3_PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk | [6 empty] | PHIFS | VpROW&COL | VnCOL | starts fast (col) shift register | clk | 0/-6
28 | Q3_PHISS | VpROW&COL | VnROW | starts slow (row) shift register | clk | [6 empty] | PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk | 0/-6
29 | Q3_VDDUC | -4.0 to -3.5 |  | unit cell drain | bias | [6 empty] | PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk | 0/-6
30 | Q3_VDDCL | -1.2 | VDDUC | column line overload protection | clk | [6 empty] | PHIDES | VpROW&COL | VnROW | row deselect | clk
31 | Q3_VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk | [6 empty] | ANALOG CLOCKS
32 | Q3_TEND |  |  | diagnostic | [7 empty] | VRSTG | -3.5 | -5.6 | global reset | clk | -3.75
33 | Q3_VRSTG | -3.5 | -5.6 | global reset | clk | [6 empty] | VRSTR | -3.5 | -5.6 | row reset | clk | -6
34 | Q3_VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position | [7 empty] | VROWON | -6.0 to -5.6 |  | unit cell row enable | aclk | -6
35 | Q3_IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource | [6 empty] | VGGCL | VDETCOM | -5 | column bus gate voltage | aclk | 0
36 | Q3_VROWON | -6.0 to -5.6 |  | unit cell row enable on position | [7 empty] | VDDCL | -1.2 | VDDUC | column bus drain voltage | aclk | 0
37 | Q3_VRSTR | -3.5 | -5.6 | row reset | clk
38 | Q3_OUT1 |  |  | p channel output 1 | output
39 | Q3_OUT2 |  |  | p channel output 2 | output
40 | Q3_OUT3 |  |  | p channel output 3 | output
41 | Q3_OUT4 |  |  | p channel output 4 | output
42 | Q3_OUT5 |  |  | p channel output 5 | output
43 | Q3_OUT6 |  |  | p channel output 6 | output
44 | Q3_OUT7 |  |  | p channel output 7 | output
45 | Q3_OUT8 |  |  | p channel output 8 | output
46 | Q3_VDDOUT | -1.0 to -1.3 |  | output drain | bias
47 | Q3_VDETCOM | -2.8 to-3.4 |  | detector common supply | bias
48 | Q4_VDDOUT | -1.0 to -1.3 |  | output drain | bias
49 | Q4_OUT8 |  |  | p channel output 8 | output
50 | Q4_OUT7 |  |  | p channel output 7 | output
51 | Q4_OUT6 |  |  | p channel output 6 | output
52 | Q4_OUT5 |  |  | p channel output 5 | output
53 | Q4_OUT4 |  |  | p channel output 4 | output
54 | Q4_OUT3 |  |  | p channel output 3 | output
55 | Q4_OUT2 |  |  | p channel output 2 | output
56 | Q4_OUT1 |  |  | p channel output 1 | output
57 | Q4_VRSTR | -3.5 | -5.6 | row reset | clk
58 | Q4_VROWON | -6.0 to -5.6 |  | unit cell row enable on position
59 | Q4_IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource
60 | Q4_VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position
61 | Q4_VRSTG | -3.5 | -5.6 | global reset | clk
62 | Q4_TEND |  |  | diagnostic
63 | Q4_VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk
64 | Q4_VDDCL | -1.2 | VDDUC | column line overload protection | clk
65 | Q4_VDDUC | -4.0 to -3.5 |  | unit cell drain | bias
66 | Q4_PHISS | VpROW&COL | VnROW | starts slow (row) shift register | clk
67 | Q4_PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk
68 | Q4_PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk
69 | Q4_PHIROE | VpROW&COL | VnROW |  | clk
70 | Q4_PHIDES | VpROW&COL | VnROW |  | clk
71 | Q4_PHIFS | VpROW&COL | VnCOL | starts fast (col) shift register | clk
72 | Q4_PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk
73 | Q4_PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk
74 | Q4_VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias
75 | Q4_VnCOL | -6.0 to -3.5 |  | negative supply for fast shift reg | bias
76 | Q4_VP | 0.0 to -0.5 |  | positive supply for shift reg | bias
77 | VSS | 0 |  | unit cell source | bias
78 | VSUB | 0 gnd |  | readout substrate | bias
79 | VSS | 0 |  | unit cell source | bias
80 | Q2_Vp | 0.0 to -0.5 |  | positive supply for shift reg | bias
81 | Q2_VnCOL | -6.0 to -3.5 |  | negative supply for fast shift reg | bias
82 | Q2_VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias
83 | Q2_PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk
84 | Q2_PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk
85 | Q2_PHIFS | VpROW&COL | VnCOL | starts fast (col) shift register | clk
86 | Q2_PHIDES | VpROW&COL | VnROW | row deselect | clk
87 | Q2_PHIROE | VpROW&COL | VnROW | odd/even clock | clk
88 | Q2_PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk
89 | Q2_PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk
90 | Q2_PHISS | VpROW&COL | VnROW | starts slow (row) shift register | clk
91 | Q2_VDDUC | -4.0 to -3.5 |  | unit cell drain | bias
92 | Q2_VDDCL | -1.2 | VDDUC | column line overload protection | clk
93 | Q2_VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk
94 | Q2_TEND |  |  | diagnostic |  |  |  | ????
95 | Q2_VRSTG | -3.5 | -5.6 | global reset | clk
96 | Q2_VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position
97 | Q2_IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource
98 | Q2_VROWON | -6.0 to -5.6 |  | unit cell row enable on position
99 | Q2_VRSTR | -3.5 | -5.6 | unit cell reset on position = row reset | clk
100 | Q2_OUT1 |  |  | p channel output 1 | output
101 | Q2_OUT2 |  |  | p channel output 2 | output
102 | Q2_OUT3 |  |  | p channel output 3 | output
103 | Q2_OUT4 |  |  | p channel output 4 | output
104 | Q2_OUT5 |  |  | p channel output 5 | output
105 | Q2_OUT6 |  |  | p channel output 6 | output
106 | Q2_OUT7 |  |  | p channel output 7 | output
107 | Q2_OUT8 |  |  | p channel output 8 | output
108 | Q2_VDDOUT | -1.0 to -1.3 |  | output drain | bias
109 | VDETCOM | -2.8 to-3.4 |  | detector common supply | bias |  |  |  | Z | 1
110 | Q1_VDDOUT | -1.0 to -1.3 |  | output drain | bias |  |  |  | J |  | 28
111 | Q1_OUT8 |  |  | p channel output 8 | output |  |  |  | G |  | 26
112 | Q1_OUT7 |  |  | p channel output 7 | output |  |  |  | h |  | 24
113 | Q1_OUT6 |  |  | p channel output 6 | output |  |  |  | z |  | 22
114 | Q1_OUT5 |  |  | p channel output 5 | output |  |  |  | KK |  | 20
115 | Q1_OUT4 |  |  | p channel output 4 | output |  |  |  | E |  | 18
116 | Q1_OUT3 |  |  | p channel output 3 | output |  |  |  | f |  | 16
117 | Q1_OUT2 |  |  | p channel output 2 | output |  |  |  | C |  | 14
118 | Q1_OUT1 |  |  | p channel output 1 | output |  |  |  | x |  | 12
119 | Q1_VRSTR | -3.5 | -5.6 | row reset | clk |  |  |  | JJ |  | 9
120 | Q1_VROWON | -6.0 to -5.6 |  | unit cell row enable on position
121 | Q1_IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource
122 | Q1_VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position
123 | Q1_VRSTG | -3.5 | -5.6 | global reset | clk
124 | Q1_TEND |  |  | diagnostic |  |  |  | ??? | j |  | 30
1
31
62
63
32
93
94
124
quad 1?
quad 3
quad 4
quad 2
first pixel
first pixel
first pixel
first pixel
last pixels
slow mux
fast mux
fast mux
slow mux
fast mux
fast mux

## wire count
SPECTROGRAPH ARRAY
two half array readouts= quads 1&2, 3&4
array pin | function | high level | low level | function |  | quad | signal | fanout bd | count
1 | VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk | 1 | Q12_VGGCL | 1,93 | 1
2 | VDDCL | -1.2 | VDDUC | column line overload protection | clk | 1 | Q12_VDDCL | 2,92 | 2
3 | VDDUC | -4.0 to -3.5 |  | unit cell drain | bias | 1 | Q12_VDDUC | 3,91 | 3
4 | PHISSYNC | VpROW&COL | VnROW | starts slow (row) shift register | clk | 1 | Q12_PSSYNC | 4,90 | 4
5 | PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk | 1 | Q12_PS1 | 5,89 | 5
6 | PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk | 1 | Q12_PS2 | 6,88 | 6
7 | PHIS3 | VpROW&COL | VnROW | phase 3 clock | clk | 1 | Q12_PS3 | 7,87 | 7
8 | PHIRST | VpROW&COL | VnROW | resets detector bias | clk | 1 | Q12_PRST | 8,86 | 8
9 | PHIFSYNC | VpROW&COL | VnCOL | starts fast (col) shift register | clk | 1 | Q12_PFSYNC | 9,85 | 9
10 | PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk | 1 | Q12_PF1 | 10,84 | 10
11 | PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk | 1 | Q12_PF2 | 11,83 | 11
12 | VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias | 1 | Q12_VNROW | 12,82 | 12
13 | VnCOL,MPEN | -6.0 to -3.5 |  | negative supply for fast shift reg | bias | 1 | Q12_VNCOL | 13,81 | 13
14 | VpROW&COL | 0.0 to -0.5 |  | positive supply for shift reg | bias | 1 | Q12_VPROW | 14,80 | 14
15 | VSSUC | 0 |  | unit cell source | bias | 1 | Q12_VSSUC | 15,79 | 15
16 | VSUB | 0 gnd |  | readout substrate | bias | 1,2,3,4 | Q12_VSUB | 16 | 16
79 | VSSUC | 0 |  | unit cell source | bias | 2
80 | VpROW&COL | 0.0 to -0.5 |  | positive supply for shift reg | bias | 2
81 | VnCOL,MPEN | -6.0 to -3.5 |  | negative supply for fast shift reg | bias | 2
82 | VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias | 2
83 | PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk | 2
84 | PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk | 2
85 | PHIFSYNC | VpROW&COL | VnCOL | starts fast (col) shift register | clk | 2
86 | PHIRST | VpROW&COL | VnROW | resets detector bias | clk | 2
87 | PHIS3 | VpROW&COL | VnROW | phase 3 clock | clk | 2
88 | PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk | 2
89 | PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk | 2
90 | PHISSYNC | VpROW&COL | VnROW | starts slow (row) shift register | clk | 2
91 | VDDUC | -4.0 to -3.5 |  | unit cell drain | bias | 2
92 | VDDCL | -1.2 | VDDUC | column line overload protection | clk | 2
93 | VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk | 2
94 | TEND |  |  | diagnostic |  | 2 | Q2_TEND | 94 | 17
95 | VRSTOFF | -3.5 | -5.6 | unit cell reset off position = global reset | clk | 2 | Q12_VRSTOFF | 95123 | 18
96 | VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position |  | 2 | Q12_VROWOFF | 96122 | 19
97 | IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource | 2 | Q2_IREF | 97 | 20
98 | VROWON | -6.0 to -5.6 |  | unit cell row enable on position |  | 2 | Q12_VROWON | 98120 | 21
99 | VRSTON | -3.5 | -5.6 | unit cell reset on position = row reset | clk | 2 | Q12_RSTON | 99119 | 22
100 | POUT1 |  |  | p channel output 1 | output | 2 | Q2_OUT1 | 100 | 23
101 | POUT2 |  |  | p channel output 2 | output | 2 | Q2_OUT2 | 101 | 24
102 | POUT3 |  |  | p channel output 3 | output | 2 | Q2_OUT3 | 102 | 25
103 | POUT4 |  |  | p channel output 4 | output | 2 | Q2_OUT4 | 103 | 26
104 | POUT5 |  |  | p channel output 5 | output | 2 | Q2_OUT5 | 104 | 27
105 | POUT6 |  |  | p channel output 6 | output | 2 | Q2_OUT6 | 105 | 28
106 | POUT7 |  |  | p channel output 7 | output | 2 | Q2_OUT7 | 106 | 29
107 | POUT8 |  |  | p channel output 8 | output | 2 | Q2_OUT8 | 107 | 30
108 | VDDOUT | -1.0 to -1.3 |  | output drain | bias | 2 | Q12_VDDOUT | 108110 | 31
109 | VDETCOM | -2.8 to-3.4 |  | detector common supply | bias | 1,2,3,4 | Q12_VDETCOM | 109 | 32
110 | VDDOUT | -1.0 to -1.3 |  | output drain | bias | 1
111 | POUT8 |  |  | p channel output 8 | output | 1 | Q1_OUT8
112 | POUT7 |  |  | p channel output 7 | output | 1 | Q1_OUT7
113 | POUT6 |  |  | p channel output 6 | output | 1 | Q1_OUT6
114 | POUT5 |  |  | p channel output 5 | output | 1 | Q1_OUT5
115 | POUT4 |  |  | p channel output 4 | output | 1 | Q1_OUT4
116 | POUT3 |  |  | p channel output 3 | output | 1 | Q1_OUT3
117 | POUT2 |  |  | p channel output 2 | output | 1 | Q1_OUT2
118 | POUT1 |  |  | p channel output 1 | output | 1 | Q1_OUT1
119 | VRSTON | -3.5 | -5.6 | unit cell reset on position = row reset | clk | 1
120 | VROWON | -6.0 to -5.6 |  | unit cell row enable on position |  | 1
121 | IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource | 1 | Q1_IREF | 121 | 33
122 | VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position |  | 1
123 | VRSTOFF | -3.5 | -5.6 | unit cell reset off position = global reset | clk | 1
124 | TEND |  |  | diagnostic |  | 1 | Q1_TEND | 124 | 34
QUADS 3&4
array pin | function | high level | low level | function |  | quad | signal | fanout bd | count
17 | VSSUC | 0 |  | unit cell source | bias | 3 | Q34_VSSUC | 17,77 | 35
18 | VpROW&COL | 0.0 to -0.5 |  | positive supply for shift reg | bias | 3 | Q34_VPROW | 18,76 | 36
19 | VnCOL,MPEN | -6.0 to -3.5 |  | negative supply for fast shift reg | bias | 3 | Q34_VNCOL | 19,75 | 37
20 | VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias | 3 | Q34_VNROW | 20,74 | 38
21 | PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk | 3 | Q34_PF2 | 21,73 | 39
22 | PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk | 3 | Q34_PF1 | 22,72 | 40
23 | PHIFSYNC | VpROW&COL | VnCOL | starts fast (col) shift register | clk | 3 | Q34_PFSYNC | 23,71 | 41
24 | PHIRST | VpROW&COL | VnROW | resets detector bias | clk | 3 | Q34_PRST | 24,70 | 42
25 | PHIS3 | VpROW&COL | VnROW | phase 3 clock | clk | 3 | Q34_PS3 | 25,69 | 43
26 | PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk | 3 | Q34_PS2 | 26,68 | 44
27 | PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk | 3 | Q34_PS1 | 27,67 | 45
28 | PHISSYNC | VpROW&COL | VnROW | starts slow (row) shift register | clk | 3 | Q34_PSSYNC | 28,66 | 46
29 | VDDUC | -4.0 to -3.5 |  | unit cell drain | bias | 3 | Q34_VDDUC | 29,65 | 47
30 | VDDCL | -1.2 | VDDUC | column line overload protection | clk | 3 | Q34_VDDCL | 30,64 | 48
31 | VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk | 3 | Q34_VGGCL | 31,63 | 49
32 | TEND |  |  | diagnostic |  | 3 | Q3_TEND | 32 | 50
33 | VRSTOFF | -3.5 | -5.6 | unit cell reset off position = global reset | clk | 3 | Q34_VRSTOFF | 33,61 | 51
34 | VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position |  | 3 | Q34_VROWOFF | 34,60 | 52
35 | IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource | 3 | Q3_IREF | 35 | 53
36 | VROWON | -6.0 to -5.6 |  | unit cell row enable on position |  | 3 | Q34_VROWON | 36,58 | 54
37 | VRSTON | -3.5 | -5.6 | unit cell reset on position = row reset | clk | 3 | Q34_RSTON | 37,57 | 55
38 | POUT1 |  |  | p channel output 1 | output | 3 | Q3_OUT1 | 38 | 56
39 | POUT2 |  |  | p channel output 2 | output | 3 | Q3_OUT2 | 39 | 57
40 | POUT3 |  |  | p channel output 3 | output | 3 | Q3_OUT3 | 40 | 58
41 | POUT4 |  |  | p channel output 4 | output | 3 | Q3_OUT4 | 41 | 59
42 | POUT5 |  |  | p channel output 5 | output | 3 | Q3_OUT5 | 42 | 60
43 | POUT6 |  |  | p channel output 6 | output | 3 | Q3_OUT6 | 43 | 61
44 | POUT7 |  |  | p channel output 7 | output | 3 | Q3_OUT7 | 44 | 62
45 | POUT8 |  |  | p channel output 8 | output | 3 | Q3_OUT8 | 45 | 63
46 | VDDOUT | -1.0 to -1.3 |  | output drain | bias | 3 | Q34_VDDOUT | 46,48 | 64
47 | VDETCOM | -2.8 to-3.4 |  | detector common supply | bias | 1,2,3,4 | Q34_VDETCOM
48 | VDDOUT | -1.0 to -1.3 |  | output drain | bias | 4
49 | POUT8 |  |  | p channel output 8 | output | 4 | Q1_OUT8 | 49 | 65
50 | POUT7 |  |  | p channel output 7 | output | 4 | Q1_OUT7 | 50 | 66
51 | POUT6 |  |  | p channel output 6 | output | 4 | Q1_OUT6 | 51 | 67
52 | POUT5 |  |  | p channel output 5 | output | 4 | Q1_OUT5 | 52 | 68
53 | POUT4 |  |  | p channel output 4 | output | 4 | Q1_OUT4 | 53 | 69
54 | POUT3 |  |  | p channel output 3 | output | 4 | Q1_OUT3 | 54 | 70
55 | POUT2 |  |  | p channel output 2 | output | 4 | Q1_OUT2 | 55 | 71
56 | POUT1 |  |  | p channel output 1 | output | 4 | Q1_OUT1 | 56 | 72
57 | VRSTON | -3.5 | -5.6 | unit cell reset on position = row reset | clk | 4
58 | VROWON | -6.0 to -5.6 |  | unit cell row enable on position |  | 4
59 | IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource | 4 | Q4_IREF | 59 | 73
60 | VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position |  | 4
61 | VRSTOFF | -3.5 | -5.6 | unit cell reset off position = global reset | clk | 4
62 | TEND |  |  | diagnostic |  | 4 | Q4_TEND | 62 | 74
63 | VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk | 4
64 | VDDCL | -1.2 | VDDUC | column line overload protection | clk | 4
65 | VDDUC | -4.0 to -3.5 |  | unit cell drain | bias | 4
66 | PHISSYNC | VpROW&COL | VnROW | starts slow (row) shift register | clk | 4
67 | PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk | 4
68 | PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk | 4
69 | PHIS3 | VpROW&COL | VnROW | phase 3 clock | clk | 4
70 | PHIRST | VpROW&COL | VnROW | resets detector bias | clk | 4
71 | PHIFSYNC | VpROW&COL | VnCOL | starts fast (col) shift register | clk | 4
72 | PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk | 4
73 | PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk | 4
74 | VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias | 4
75 | VnCOL,MPEN | -6.0 to -3.5 |  | negative supply for fast shift reg | bias | 4
76 | VpROW&COL | 0.0 to -0.5 |  | positive supply for shift reg | bias | 4
77 | VSSUC | 0 |  | unit cell source | bias | 4
78 | VSUB | 0 gnd |  | readout substrate | bias | 1,2,3,4 | Q34_VSUB | 78 | 75
IMAGING ARRAY
ONE QUAD #TBD
array pin | function |  |  |  |  | quad | signal | fanout bd | count
1 | VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk | 1 | IA_VGGCL | 1,93,31,63 | 76
2 | VDDCL | -1.2 | VDDUC | column line overload protection | clk | 1 | IA_VDDCL | 2,92,30,64 | 77
3 | VDDUC | -4.0 to -3.5 |  | unit cell drain | bias | 1 | IA_VDDUC | 3,91,29,65 | 78
4 | PHISSYNC | VpROW&COL | VnROW | starts slow (row) shift register | clk | 1 | IA_PSSYNC | 4,90,28,66 | 79
5 | PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk | 1 | IA_PS1 | 5,89,27,67 | 80
6 | PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk | 1 | IA_PS2 | 6,88,26,68 | 81
7 | PHIS3 | VpROW&COL | VnROW | phase 3 clock | clk | 1 | IA_PS3 | 7,87,25,69 | 82
8 | PHIRST | VpROW&COL | VnROW | resets detector bias | clk | 1 | IA_PRST | 8,86,24,70 | 83
9 | PHIFSYNC | VpROW&COL | VnCOL | starts fast (col) shift register | clk | 1 | IA_PFSYNC | 9,85,23,71 | 84
10 | PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk | 1 | IA_PF1 | 10,84,22,72 | 85
11 | PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk | 1 | IA_PF2 | 11,83,21,73 | 86
12 | VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias | 1 | IA_VNROW | 12,82,20,74 | 87
13 | VnCOL,MPEN | -6.0 to -3.5 |  | negative supply for fast shift reg | bias | 1 | IA_VNCOL | 13,81,19,75 | 88
14 | VpROW&COL | 0.0 to -0.5 |  | positive supply for shift reg | bias | 1 | IA_VPROW | 14,80,18,76 | 89
15 | VSSUC | 0 |  | unit cell source | bias | 1 | IA_VSSUC | 15,79,17,77 | 90
16 | VSUB | 0 gnd |  | readout substrate | bias | 1,2,3,4 | IA_VSUB | 16 | 91
17 | VSSUC | 0 |  | unit cell source | bias | 3
18 | VpROW&COL | 0.0 to -0.5 |  | positive supply for shift reg | bias | 3
19 | VnCOL,MPEN | -6.0 to -3.5 |  | negative supply for fast shift reg | bias | 3
20 | VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias | 3
21 | PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk | 3
22 | PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk | 3
23 | PHIFSYNC | VpROW&COL | VnCOL | starts fast (col) shift register | clk | 3
24 | PHIRST | VpROW&COL | VnROW | resets detector bias | clk | 3
25 | PHIS3 | VpROW&COL | VnROW | phase 3 clock | clk | 3
26 | PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk | 3
27 | PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk | 3
28 | PHISSYNC | VpROW&COL | VnROW | starts slow (row) shift register | clk | 3
29 | VDDUC | -4.0 to -3.5 |  | unit cell drain | bias | 3
30 | VDDCL | -1.2 | VDDUC | column line overload protection | clk | 3
31 | VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk | 3
32 | TEND |  |  | diagnostic |  | 3 | Q3_TEND | 32 | 92
33 | VRSTOFF | -3.5 | -5.6 | unit cell reset off position = global reset | clk | 3
34 | VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position |  | 3
35 | IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource | 3 | Q3_IREF | 35 | 93
36 | VROWON | -6.0 to -5.6 |  | unit cell row enable on position |  | 3
37 | VRSTON | -3.5 | -5.6 | unit cell reset on position = row reset | clk | 3
38 | POUT1 |  |  | p channel output 1 | output | 3 | Q3_OUT1 | 38 | 94
39 | POUT2 |  |  | p channel output 2 | output | 3 | Q3_OUT2 | 39 | 95
40 | POUT3 |  |  | p channel output 3 | output | 3 | Q3_OUT3 | 40 | 96
41 | POUT4 |  |  | p channel output 4 | output | 3 | Q3_OUT4 | 41 | 97
42 | POUT5 |  |  | p channel output 5 | output | 3 | Q3_OUT5 | 42 | 98
43 | POUT6 |  |  | p channel output 6 | output | 3 | Q3_OUT6 | 43 | 99
44 | POUT7 |  |  | p channel output 7 | output | 3 | Q3_OUT7 | 44 | 100
45 | POUT8 |  |  | p channel output 8 | output | 3 | Q3_OUT8 | 45 | 101
46 | VDDOUT | -1.0 to -1.3 |  | output drain | bias | 3
47 | VDETCOM | -2.8 to-3.4 |  | detector common supply | bias | 1,2,3,4 | Q34_VDETCOM
48 | VDDOUT | -1.0 to -1.3 |  | output drain | bias | 4
49 | POUT8 |  |  | p channel output 8 | output | 4 | Q1_OUT8 | 49 | 102
50 | POUT7 |  |  | p channel output 7 | output | 4 | Q1_OUT7 | 50 | 103
51 | POUT6 |  |  | p channel output 6 | output | 4 | Q1_OUT6 | 51 | 104
52 | POUT5 |  |  | p channel output 5 | output | 4 | Q1_OUT5 | 52 | 105
53 | POUT4 |  |  | p channel output 4 | output | 4 | Q1_OUT4 | 53 | 106
54 | POUT3 |  |  | p channel output 3 | output | 4 | Q1_OUT3 | 54 | 107
55 | POUT2 |  |  | p channel output 2 | output | 4 | Q1_OUT2 | 55 | 108
56 | POUT1 |  |  | p channel output 1 | output | 4 | Q1_OUT1 | 56 | 109
57 | VRSTON | -3.5 | -5.6 | unit cell reset on position = row reset | clk | 4
58 | VROWON | -6.0 to -5.6 |  | unit cell row enable on position |  | 4
59 | IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource | 4 | Q4_IREF | 59 | 110
60 | VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position |  | 4
61 | VRSTOFF | -3.5 | -5.6 | unit cell reset off position = global reset | clk | 4
62 | TEND |  |  | diagnostic |  | 4 | Q4_TEND | 62 | 111
63 | VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk | 4
64 | VDDCL | -1.2 | VDDUC | column line overload protection | clk | 4
65 | VDDUC | -4.0 to -3.5 |  | unit cell drain | bias | 4
66 | PHISSYNC | VpROW&COL | VnROW | starts slow (row) shift register | clk | 4
67 | PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk | 4
68 | PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk | 4
69 | PHIS3 | VpROW&COL | VnROW | phase 3 clock | clk | 4
70 | PHIRST | VpROW&COL | VnROW | resets detector bias | clk | 4
71 | PHIFSYNC | VpROW&COL | VnCOL | starts fast (col) shift register | clk | 4
72 | PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk | 4
73 | PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk | 4
74 | VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias | 4
75 | VnCOL,MPEN | -6.0 to -3.5 |  | negative supply for fast shift reg | bias | 4
76 | VpROW&COL | 0.0 to -0.5 |  | positive supply for shift reg | bias | 4
77 | VSSUC | 0 |  | unit cell source | bias | 4
78 | VSUB | 0 gnd |  | readout substrate | bias | 1,2,3,4 | Q34_VSUB | 78 | 112
79 | VSSUC | 0 |  | unit cell source | bias | 2
80 | VpROW&COL | 0.0 to -0.5 |  | positive supply for shift reg | bias | 2
81 | VnCOL,MPEN | -6.0 to -3.5 |  | negative supply for fast shift reg | bias | 2
82 | VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias | 2
83 | PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk | 2
84 | PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk | 2
85 | PHIFSYNC | VpROW&COL | VnCOL | starts fast (col) shift register | clk | 2
86 | PHIRST | VpROW&COL | VnROW | resets detector bias | clk | 2
87 | PHIS3 | VpROW&COL | VnROW | phase 3 clock | clk | 2
88 | PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk | 2
89 | PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk | 2
90 | PHISSYNC | VpROW&COL | VnROW | starts slow (row) shift register | clk | 2
91 | VDDUC | -4.0 to -3.5 |  | unit cell drain | bias | 2
92 | VDDCL | -1.2 | VDDUC | column line overload protection | clk | 2
93 | VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk | 2
94 | TEND |  |  | diagnostic |  | 2 | Q2_TEND | 94 | 113
95 | VRSTOFF | -3.5 | -5.6 | unit cell reset off position = global reset | clk | 2 | IA_VRSTOFF | 95,123,33,61 | 114
96 | VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position |  | 2 | IA_VROWOFF | 96,122,34,60 | 115
97 | IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource | 2 | Q2_IREF | 97 | 116
98 | VROWON | -6.0 to -5.6 |  | unit cell row enable on position |  | 2 | IA_VROWON | 98,120,36,58 | 117
99 | VRSTON | -3.5 | -5.6 | unit cell reset on position = row reset | clk | 2 | IA_RSTON | 99,119,37,57 | 118
100 | POUT1 |  |  | p channel output 1 | output | 2 | Q2_OUT1 | 100 | 119
101 | POUT2 |  |  | p channel output 2 | output | 2 | Q2_OUT2 | 101 | 120
102 | POUT3 |  |  | p channel output 3 | output | 2 | Q2_OUT3 | 102 | 121
103 | POUT4 |  |  | p channel output 4 | output | 2 | Q2_OUT4 | 103 | 122
104 | POUT5 |  |  | p channel output 5 | output | 2 | Q2_OUT5 | 104 | 123
105 | POUT6 |  |  | p channel output 6 | output | 2 | Q2_OUT6 | 105 | 124
106 | POUT7 |  |  | p channel output 7 | output | 2 | Q2_OUT7 | 106 | 125
107 | POUT8 |  |  | p channel output 8 | output | 2 | Q2_OUT8 | 107 | 126
108 | VDDOUT | -1.0 to -1.3 |  | output drain | bias | 2 | IA_VDDOUT | 108,110,46,48 | 127
109 | VDETCOM | -2.8 to-3.4 |  | detector common supply | bias | 1,2,3,4 | IA_VDETCOM | 109 | 128
110 | VDDOUT | -1.0 to -1.3 |  | output drain | bias | 1
111 | POUT8 |  |  | p channel output 8 | output | 1 | Q1_OUT8
112 | POUT7 |  |  | p channel output 7 | output | 1 | Q1_OUT7
113 | POUT6 |  |  | p channel output 6 | output | 1 | Q1_OUT6
114 | POUT5 |  |  | p channel output 5 | output | 1 | Q1_OUT5
115 | POUT4 |  |  | p channel output 4 | output | 1 | Q1_OUT4
116 | POUT3 |  |  | p channel output 3 | output | 1 | Q1_OUT3
117 | POUT2 |  |  | p channel output 2 | output | 1 | Q1_OUT2
118 | POUT1 |  |  | p channel output 1 | output | 1 | Q1_OUT1
119 | VRSTON | -3.5 | -5.6 | unit cell reset on position = row reset | clk | 1
120 | VROWON | -6.0 to -5.6 |  | unit cell row enable on position |  | 1
121 | IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource | 1 | Q1_IREF | 121 | 129
122 | VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position |  | 1
123 | VRSTOFF | -3.5 | -5.6 | unit cell reset off position = global reset | clk | 1
124 | TEND |  |  | diagnostic |  | 1 | Q1_TEND | 124 | 130
HALL EFFECT SENSORS
F.W.BELL FH-301-040 HALL GENERATOR
DICHROIC TURRET
pin | function | spec | color | notes |  |  | signal
1 | I+ | 15-30 mA | red | [5 empty] | 131
2 | I- |  | blk | [5 empty] | 132
3 | OUT- |  | yel | Rin=20 to 80ohms |  |  |  |  | 133
4 | OUT+ |  | blu | Rout=50 to 200ohms |  |  |  |  | 134
ORDER SORTER #1
1 | I+ | 15-30 mA | red | [5 empty] | 135
2 | I- |  | blk | [5 empty] | 136
3 | OUT- |  | yel | Rin=20 to 80ohms |  |  |  |  | 137
4 | OUT+ |  | blu | Rout=50 to 200ohms |  |  |  |  | 138
K MIRROR ROTATOR
1 | I+ | 15-30 mA | red | [5 empty] | 139
2 | I- |  | blk | [5 empty] | 140
3 | OUT- |  | yel | Rin=20 to 80ohms |  |  |  |  | 141
4 | OUT+ |  | blu | Rout=50 to 200ohms |  |  |  |  | 142
SLIT WHEEL
1 | I+ | 15-30 mA | red | [5 empty] | 143
2 | I- |  | blk | [5 empty] | 144
3 | OUT- |  | yel | Rin=20 to 80ohms |  |  |  |  | 145
4 | OUT+ |  | blu | Rout=50 to 200ohms |  |  |  |  | 146
ORDER SORTER #2
1 | I+ | 15-30 mA | red | [5 empty] | 147
2 | I- |  | blk | [5 empty] | 148
3 | OUT- |  | yel | Rin=20 to 80ohms |  |  |  |  | 149
4 | OUT+ |  | blu | Rout=50 to 200ohms |  |  |  |  | 150
GRATING TURRET
1 | I+ | 15-30 mA | red | [5 empty] | 151
2 | I- |  | blk | [5 empty] | 152
3 | OUT- |  | yel | Rin=20 to 80ohms |  |  |  |  | 153
4 | OUT+ |  | blu | Rout=50 to 200ohms |  |  |  |  | 154
1 | I+ | 15-30 mA | red | [5 empty] | 155
2 | I- |  | blk | [5 empty] | 156
3 | OUT- |  | yel | Rin=20 to 80ohms |  |  |  |  | 157
4 | OUT+ |  | blu | Rout=50 to 200ohms |  |  |  |  | 158
TEMPERATURE SENSORS
LAKESHORE SILICON DIODES
SPECTROGRAPH ARRAY
pin | function | spec | color | notes |  |  | signal
1 | V+ | [7 empty] | 159
2 | I+ | [7 empty] | 160
3 | V- | [7 empty] | 161
4 | I- | [7 empty] | 162
IMAGING QUAD
1 | V+ | [7 empty] | 163
2 | I+ | [7 empty] | 164
3 | V- | [7 empty] | 165
4 | I- | [7 empty] | 166
COLD STRUCTURE
1 | V+ | [7 empty] | 167
2 | I+ | [7 empty] | 168
3 | V- | [7 empty] | 169
4 | I- | [7 empty] | 170
TEMP4
1 | V+ | [7 empty] | 171
2 | I+ | [7 empty] | 172
3 | V- | [7 empty] | 173
4 | I- | [7 empty] | 174
TEMP5
1 | V+ | [7 empty] | 175
2 | I+ | [7 empty] | 176
3 | V- | [7 empty] | 177
4 | I- | [7 empty] | 178
TEMP6
1 | V+ | [7 empty] | 179
2 | I+ | [7 empty] | 180
3 | V- | [7 empty] | 181
4 | I- | [7 empty] | 182
TEMP7
1 | V+ | [7 empty] | 183
2 | I+ | [7 empty] | 184
3 | V- | [7 empty] | 185
4 | I- | [7 empty] | 186
TEMP8
1 | V+ | [7 empty] | 187
2 | I+ | [7 empty] | 188
3 | V- | [7 empty] | 189
4 | I- | [7 empty] | 190
SAFETY LIMIT SWITCHES
MICROSWITCH
pin
1 | [8 empty] | 191
2 | [8 empty] | 192
1 | [8 empty] | 193
2 | [8 empty] | 194

## ARRAY PINOUT
ALLADIN 1 pinout (SB084 ROIC PCHANNEL OUTPUTS)
socket Yamaichi IC51-1244-410-1
pin num | signal name | high level | low level | function |  | signal name | high level | low level | function |  | test
1 | VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk | SUBSTRATES
2 | VDDCL | -1.2 | VDDUC | column line overload protection | clk | VSUB | 0 gnd |  | readout substrate | bias | 0
3 | VDDUC | -4.0 to -3.5 |  | unit cell drain | bias | VDETCOM | -2.8 to-3.4 |  | detector common supply/substrate | bias | 0
4 | PHISSYNC | VpROW&COL | VnROW | starts slow (row) shift register | clk | UNIT CELL
5 | PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk | VDDUC | -4.0 to -3.5 |  | unit cell drain ~8mA | bias | -4
6 | PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk | VSSUC | 0 |  | unit cell source | bias | 0
7 | PHIS3 | VpROW&COL | VnROW | phase 3 clock | clk | VRSTOFF | -3.5 | -5.6 | unit cell reset off position = global reset | clk | -3.75
8 | PHIRST | VpROW&COL | VnROW | resets detector bias | clk | VRSTON | -3.5 | -5.6 | unit cell reset on position = row reset | clk | -6
9 | PHIFSYNC | VpROW&COL | VnCOL | starts fast (col) shift register | clk | VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position |  | 0
10 | PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk | VROWON | -6.0 to -5.6 |  | unit cell row enable on position |  | -6
11 | PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk | OUTPUTS
12 | VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias | POUT# |  |  | 100 to 300uA I load
13 | VnCOL,MPEN | -6.0 to -3.5 |  | negative supply for fast shift reg | bias | VDDOUT | -1.0 to -1.3 |  | output drain ~16mA |  | -1.3
14 | VpROW&COL | 0.0 to -0.5 |  | positive supply for shift reg | bias | CURRENT SOURCES
15 | VSSUC | 0 |  | unit cell source | bias | IREF* | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource | 250K
16 | VSUB | 0 gnd |  | readout substrate | bias | SHIFT REGISTERS
17 | VSSUC | 0 |  | unit cell source | bias | VpROW&COL | 0.0 to -0.5 |  | positive supply for shift reg | bias | 0
18 | VpROW&COL | 0.0 to -0.5 |  | positive supply for shift reg | bias | VnCOL,MPEN | -6.0 to -3.5 |  | negative supply for fast shift reg | bias | -6
19 | VnCOL,MPEN | -6.0 to -3.5 |  | negative supply for fast shift reg | bias | VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias | -6
20 | VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias | CLOCKS
21 | PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk | PHISSYNC | VpROW&COL | VnROW | starts slow (row) shift register | clk | 0/-6
22 | PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk | PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk | 0/-6
23 | PHIFSYNC | VpROW&COL | VnCOL | starts fast (col) shift register | clk | PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk | 0/-6
24 | PHIRST | VpROW&COL | VnROW | resets detector bias | clk | PHIS3 | VpROW&COL | VnROW | phase 3 clock | clk | 0/-6
25 | PHIS3 | VpROW&COL | VnROW | phase 3 clock | clk | PHIRST | VpROW&COL | VnROW | resets detector bias | clk | 0/-6
26 | PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk | PHIFSYNC | VpROW&COL | VnCOL | starts fast (col) shift register | clk | 0/-6
27 | PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk | PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk | 0/-6
28 | PHISSYNC | VpROW&COL | VnROW | starts slow (row) shift register | clk | PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk | 0/-6
29 | VDDUC | -4.0 to -3.5 |  | unit cell drain | bias | PHIRST | VpROW&COL | VnROW | resets detector bias | clk | 0/-6
30 | VDDCL | -1.2 | VDDUC | column line overload protection | clk | VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk | 0
31 | VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk | VDDCL | -1.2 | VDDUC | column line overload protection | clk | 0
32 | TEND |  |  | diagnostic |  | VRSTOFF | -3.5 | -5.6 | unit cell reset off position = global reset | clk | -3.75
33 | VRSTOFF | -3.5 | -5.6 | unit cell reset off position = global reset | clk | VRSTON | -3.5 | -5.6 | unit cell reset on position = row reset | clk | -6
34 | VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position
35 | IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource | TEND |  |  | 10Kohm to gnd shift reg diag |  | -5
36 | VROWON | -6.0 to -5.6 |  | unit cell row enable on position
37 | VRSTON | -3.5 | -5.6 | unit cell reset on position = row reset | clk
38 | POUT1 |  |  | p channel output 1 | output
39 | POUT2 |  |  | p channel output 2 | output
40 | POUT3 |  |  | p channel output 3 | output
41 | POUT4 |  |  | p channel output 4 | output
42 | POUT5 |  |  | p channel output 5 | output
43 | POUT6 |  |  | p channel output 6 | output
44 | POUT7 |  |  | p channel output 7 | output
45 | POUT8 |  |  | p channel output 8 | output
46 | VDDOUT | -1.0 to -1.3 |  | output drain | bias
47 | VDETCOM | -2.8 to-3.4 |  | detector common supply | bias
48 | VDDOUT | -1.0 to -1.3 |  | output drain | bias
49 | POUT8 |  |  | p channel output 8 | output
50 | POUT7 |  |  | p channel output 7 | output
51 | POUT6 |  |  | p channel output 6 | output
52 | POUT5 |  |  | p channel output 5 | output
53 | POUT4 |  |  | p channel output 4 | output
54 | POUT3 |  |  | p channel output 3 | output
55 | POUT2 |  |  | p channel output 2 | output
56 | POUT1 |  |  | p channel output 1 | output
57 | VRSTON | -3.5 | -5.6 | unit cell reset on position = row reset | clk
58 | VROWON | -6.0 to -5.6 |  | unit cell row enable on position
59 | IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource
60 | VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position
61 | VRSTOFF | -3.5 | -5.6 | unit cell reset off position = global reset | clk
62 | TEND |  |  | diagnostic
63 | VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk
64 | VDDCL | -1.2 | VDDUC | column line overload protection | clk
65 | VDDUC | -4.0 to -3.5 |  | unit cell drain | bias
66 | PHISSYNC | VpROW&COL | VnROW | starts slow (row) shift register | clk
67 | PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk
68 | PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk
69 | PHIS3 | VpROW&COL | VnROW | phase 3 clock | clk
70 | PHIRST | VpROW&COL | VnROW | resets detector bias | clk
71 | PHIFSYNC | VpROW&COL | VnCOL | starts fast (col) shift register | clk
72 | PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk
73 | PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk
74 | VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias
75 | VnCOL,MPEN | -6.0 to -3.5 |  | negative supply for fast shift reg | bias
76 | VpROW&COL | 0.0 to -0.5 |  | positive supply for shift reg | bias
77 | VSSUC | 0 |  | unit cell source | bias
78 | VSUB | 0 gnd |  | readout substrate | bias
79 | VSSUC | 0 |  | unit cell source | bias
80 | VpROW&COL | 0.0 to -0.5 |  | positive supply for shift reg | bias
81 | VnCOL,MPEN | -6.0 to -3.5 |  | negative supply for fast shift reg | bias
82 | VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias
83 | PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk
84 | PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk
85 | PHIFSYNC | VpROW&COL | VnCOL | starts fast (col) shift register | clk
86 | PHIRST | VpROW&COL | VnROW | resets detector bias | clk
87 | PHIS3 | VpROW&COL | VnROW | phase 3 clock | clk
88 | PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk
89 | PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk
90 | PHISSYNC | VpROW&COL | VnROW | starts slow (row) shift register | clk
91 | VDDUC | -4.0 to -3.5 |  | unit cell drain | bias
92 | VDDCL | -1.2 | VDDUC | column line overload protection | clk
93 | VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk
94 | TEND |  |  | diagnostic
95 | VRSTOFF | -3.5 | -5.6 | unit cell reset off position = global reset | clk
96 | VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position
97 | IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource
98 | VROWON | -6.0 to -5.6 |  | unit cell row enable on position
99 | VRSTON | -3.5 | -5.6 | unit cell reset on position = row reset | clk
100 | POUT1 |  |  | p channel output 1 | output
101 | POUT2 |  |  | p channel output 2 | output
102 | POUT3 |  |  | p channel output 3 | output
103 | POUT4 |  |  | p channel output 4 | output
104 | POUT5 |  |  | p channel output 5 | output
105 | POUT6 |  |  | p channel output 6 | output
106 | POUT7 |  |  | p channel output 7 | output
107 | POUT8 |  |  | p channel output 8 | output
108 | VDDOUT | -1.0 to -1.3 |  | output drain | bias
109 | VDETCOM | -2.8 to-3.4 |  | detector common supply | bias
110 | VDDOUT | -1.0 to -1.3 |  | output drain | bias
111 | POUT8 |  |  | p channel output 8 | output
112 | POUT7 |  |  | p channel output 7 | output
113 | POUT6 |  |  | p channel output 6 | output
114 | POUT5 |  |  | p channel output 5 | output
115 | POUT4 |  |  | p channel output 4 | output
116 | POUT3 |  |  | p channel output 3 | output
117 | POUT2 |  |  | p channel output 2 | output
118 | POUT1 |  |  | p channel output 1 | output
119 | VRSTON | -3.5 | -5.6 | unit cell reset on position = row reset | clk
120 | VROWON | -6.0 to -5.6 |  | unit cell row enable on position
121 | IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource
122 | VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position
123 | VRSTOFF | -3.5 | -5.6 | unit cell reset off position = global reset | clk
124 | TEND |  |  | diagnostic
1
31
62
63
32
93
94
124
quad 1?
quad 3
quad 4
quad 2
first pixel
first pixel
first pixel
first pixel
last pixels
slow mux
fast mux
fast mux
slow mux
fast mux
fast mux

## FAST CLOCK (2)
FRAME START TIMING
high level resets register
phiFS
|  | rising edge/high level clocks 1st stage of register | [5 empty] | rising edge/high level clocks 2nd stage of register
phiF1
rising edge/high level enables outputs
phiF2 | [6 empty] | non overlapping edges
phiCOL1 (1-8/quad) |  | columns 1-8 addressed to col bus but not output bus enabled
1.5usec
phiCEN1 (1-8/quad) |  |  | columns 1-8 output bus enabled
phiCOL2 (9-16/quad) |  |  | columns 9-16 addressed to column bus but not output bus enabled
1.5usec
phiCEN2 (9-16/quad) | [7 empty] | columns 9-16 output bus enabled
phiCOL3
phiCEN3
|  |  |  | 1.1us |  |  | 500ns |  | 200n |  |  |  | 1.5us
Signal Chain |  |  | output & preamp settle |  | ADC convert |  |  | ship |  | DSP process and store first 32 pixels
[8 empty] | output & preamp settle |  |  | ADC convert |  | ship | DSP process and store second 32 pixels

## assignments
PRELIMINARY PIN ASSIGNMENTS
2026-07-28 19:36
31
30
29
28
27
26
25
24
23
22
21
20
19
18
17
16
15 | phiF2_B |  |  | non overlapping edges
|  |  | rising edge/high level clocks 1st stage of register |  | rising edge/high level clocks 2nd stage of register
14 | phiF1_B
high level resets register
13 | phiFS_B
12 | phiDES_B
11 | phiROE_B
10 | phiS2_B
9 | phiS1_B
8 | VDDCL_B
7 | phiSS_B
6 | FRAME_B
5 | CONVERT_B
4 | READ_DATA_B
3 | VRSTG_B
2 | VRSTR_B
1 | VROWON_B
0 | VGGCL_B
BOTTOM HALF
TOP HALF

## spexguiderloc (2)
Array Mount/location Worksheet |  | SBRC=27um pixels X 1024 = |  |  | 27648 |  | substrate=
Spex guider array mount w/sca46509 |  |  | quad= | 13.824
2000-02-04
0
0
| -21.486 |  |  |  |  | -19.366
| -11.199 |  |  |  |  | 10.227
[6 empty] | -33.349 |  | -33.451
[6 empty] | 10.297 |  | 24.281
-
-
-19.366mm
9.079mm
-34.1mm
10.517mm
-33.044mm
9.155mm
-34.235mm
24.281mm

## spexgmuxloc
Array Mount/location Worksheet |  | SBRC=27um pixels X 1024 = |  |  | 27648 | substrate=
Spex guider array mount w/mux |  |  | quad= | 13.824
2000-02-04
11.177
21.318
0
0
-21.486
-11.199
-6.740mm
21.256mm left
-6.740mm,
21.439rt,
-6.740mm,
21.342mm cntr
-7.804mm
6.181mm
-6.761mm
7.543mm
-21.712mm
6.231mm
-36.544mm
7.634mm
-35.512mm
6.271mm
-36.509mm
21.438mm
-6.730mm,
35.168mm
-7.774mm,
36.525mm
-21.577mm,
36.575mm
-35.406mm,
36.610mm
-36.460mm,
35.255mm

## spexmuxloc
Array Mount/location Worksheet |  | SBRC=27um pixels X 1024 = |  |  | 27648 | substrate=
Spex spec array mount w/mux |  |  | quad= | 13.824
2000-02-04
11.411
21.4
0
0
-21.195
-11.052
-6.519mm
21.148mm left
-6.521mm,
21.347rt,
-6.521mm,
21.243mm cntr
-7.588mm
6.081mm
-6.535mm
7.433mm
-21.394mm
6.146mm
-36.274mm
7.572mm
-35.221mm
6.195mm
-36.241mm
21.374mm
-6.473mm,
35.062mm
-7.498mm,
36.434mm
-21.321mm,
36.486mm
-35.174mm,
36.520mm
-36.187mm,
35.184mm

## sca412202loc
Array Mount/location Worksheet |  | SBRC=27um pixels X 1024 = |  |  | 27648 |  | substrate=
Spex spec array mount w/sca412202 array |  |  | quad= | 13.824
2000-02-03
0
0
[6 empty] | -7.694 |  | -7.551 | -7.573
[6 empty] | 7.694 |  | 21.647 | 35.843
[6 empty] | -21.693 |  |  | -21.751
| -21.232 |  |  |  |  | 7.563 |  |  | 35.879
-11.209
[6 empty] | -35.772 |  | -35.794 | -35.787
[6 empty] | 7.597 |  | 21.679 | 35.922
-21.257mm
1.482mm
-1.5mm
21.491mm
-6.817mm
21.554mm left
-6.817mm,
21.744 rt,
-6.817mm,
21.646mm cntr
-7.878mm
6.455mm
-6.833mm
7.836mm
-21.695mm
6.468mm
-36.548mm
7.849mm
-35.55mm
6.494mm
-36.538mm
21.679mm
-6.802mm,
35.462mm
-7.87mm,
36.833mm
-21.693mm,
36.836mm
-35.481mm,
36.849mm
-36.558mm,
35.476mm

## mount_loc
Array Mount/location Worksheet |  | SBRC=27um pixels X 1024 = |  |  | 27648 |  | substrate=
Spex spec array mount w/206 mux |  |  | quad= | 13.824
1999-12-14
23.826
14.134
0.053
[6 empty] | 0 |  | 0.011 | 27.803
[6 empty] | 0 |  | 13.895
[6 empty] | -13.896 |  |  | -13.852
| -14.576 |  |  |  |  | 0.033 |  |  | 27.831
23.68
[6 empty] | -27.812 |  | -27.794 | -27.75
[6 empty] | 0.043 |  | 13.959 | 28.858
-52.348
14.878

## signal list (array)
ALLADIN 2 pinout (SB0152 ROIC PCHANNEL OUTPUTS) | [5 empty] | 700-150 |  |  | 700-151 |  |  |  |  | xxx |  |  |  |  | 700-149 |  |  |  |  | 600-010 |  |  |  | 700-148 |  |  |  |  | Backplane |  |  | 700-148 | [7 empty] | 2047
socket Yamaichi IC51-1244-410-1 | [5 empty] | Array Fanout |  |  | Config bd | J1,2,3,4= 2mm 60 pin header |  | P1,2=96 pin |  | Thermal feedthrough |  | COLD | WARM |  | 61 pin Bd | J1,2=34 pin headers |  |  |  | 61 pin CABLES | conn=MS3116-E24-61S | conn=MS3116-E24-61S |  | 60 pin Bd |  |  | J1,2=34 pin headers |  |  |  |  | ClockBias Bd | [7 empty] | 0.0048828
pin num | signal name | high level | low level | function |  | signal | conn | pin | signal | conn | pin | conn | pin | signal |  | pin | pin | Cu? | signal | conn | pin | pin | bd# | signal | cable# | pin | cond/shld | signal | pin | bd# | conn | pin | signal | slot# | pin | DAC#_high | DAC#_low | BD# | slot | bd | DAC# | signal_high | volts | [13 empty] | sw1 | sw2 | sw3 | sw4 | sw5 | sw6 | sw7 | sw8 | sw9 | sw10 | jhxx | jhxx | slot | bd | DAC# | signal_low | volts | [13 empty] | sw1 | sw2 | sw3 | sw4 | sw5 | sw6 | sw7 | sw8 | sw9 | sw10 | jhxx | jhxx
1 | Q1_VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk | q1_VGGCL | J1 | 1 | VGGCL_T | J1 | 1 | P1 | A2,B2 | VGGCL_T | JH1_T | 4 | 3 |  | VGGCL_T | J1 | 3 | a | 1 | VGGCL_T | T-Q1&2 | a | cond | VGGCL_T | a | 1 | J1 | 3 | VGGCL_T | slot_7 | C17 |  |  |  | 14 | 2T | 1 | VGGCL_T_H | -3.75 | -768 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ON | OFF | ON | OFF | ON | ON | ON | ON | ON | ON | ON | ON | 14 | 2T | 2 | VGGCL_T_L | -4.5 | -922 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | ON | OFF | ON | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON
2 | Q1_VDDCL | -1.2 | VDDUC | column line overload protection | clk | q1_VDDCL | J1 | 3 | VDDCL_T | J1 | 3 | P1 | A27,B27 | VDDCL_T | JH2_T | 20 | 19 |  | VDDCL_T | J2 | 19 | LL | 1 | VDDCL_T | T-Q1&2 | LL | shield | VDDCL_T | LL | 1 | J2 | 19 | VDDCL_T | slot_1 | C17 |  |  |  | 13 | 1T | 1 | VDDCL_T_H | -4.5 | -922 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | ON | OFF | ON | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | 13 | 1T | 2 | VDDCL_T_L | -1.3 | -267 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | ON | OFF | OFF | ON | OFF | OFF | OFF | OFF | ON | OFF | ON | OFF
3 | Q1_VDDUC | -4.0 to -3.5 |  | unit cell drain | bias | q1_VDDUC | J1 | 5 | VDDUC_T | J1 | 5 | P1 | C27 | VDDUC_T | JH2_T | 19 | 20 |  | VDDUC_T | J2 | 20 | KK | 1 | VDDUC_T | T-Q1&2 | KK | cond | VDDUC_T | KK | 1 | J2 | 20 | VDDUC_T | slot_7 | C28 |  |  |  | 14 | 2T | 12 | VDDUC_T | -3.75 | -768 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ON | OFF | ON | OFF | ON | ON | ON | ON | ON | ON | ON | ON
4 | Q1_PHISS | VpROW&COL | VnROW | starts slow (row) shift register | clk | q1_PSS | J1 | 7 | PSS_T | J1 | 7 | P1 | A28,B28 | PSS_T | JH2_T | 22 | 21 |  | PSS_T | J2 | 21 | AA | 1 | PSS_T | T-Q1&2 | AA | shield | PSS_T | AA | 1 | J2 | 21 | PSS_T | slot_7 | C31 |  |  |  | 14 | 2T | 15 | PSS_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 14 | 2T | 16 | PSS_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
5 | Q1_PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk | q1_PS1 | J1 | 9 | PS1_T | J1 | 9 | P1 | C28 | PS1_T | JH2_T | 22 | 21 |  | PS1_T | J2 | 22 | z | 1 | PS1_T | T-Q1&2 | z | cond | PS1_T | z | 1 | J2 | 22 | PS1_T | slot_1 | C19 |  |  |  | 13 | 1T | 3 | PS1_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 13 | 1T | 4 | PS1_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
6 | Q1_PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk | q1_PS2 | J1 | 11 | PS2_T | J1 | 11 | P1 | A29,B29 | PS2_T | JH2_T | 24 | 23 |  | PS2_T | J2 | 23 | i | 1 | PS2_T | T-Q1&2 | i | shield | PS2_T | i | 1 | J2 | 23 | PS2_T | slot_1 | C21 |  |  |  | 13 | 1T | 5 | PS2_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 13 | 1T | 6 | PS2_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
7 | Q1_PHIROE | VpROW&COL | VnROW | odd/even clock | clk | q1_PROE | J1 | 13 | PROE_T | J1 | 13 | P1 | C29 | PROE_T | JH2_T | 23 | 24 |  | PROE_T | J2 | 24 | h | 1 | PROE_T | T-Q1&2 | h | cond | PROE_T | h | 1 | J2 | 24 | PROE_T | slot_1 | C23 |  |  |  | 13 | 1T | 7 | PROE_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 13 | 1T | 8 | PROE_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
8 | Q1_PHIDES | VpROW&COL | VnROW | row deselect | clk | q1_PDES | J1 | 15 | PDES_T | J1 | 15 | P1 | A30,B30 | PDES_T | JH2_T | 26 | 25 |  | PDES_T | J2 | 25 | H | 1 | PDES_T | T-Q1&2 | H | shield | PDES_T | H | 1 | J2 | 25 | PDES_T | slot_1 | C25 |  |  |  | 13 | 1T | 9 | PDES_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 13 | 1T | 10 | PDES_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
9 | Q1_PHIFS | VpROW&COL | VnCOL | starts fast (col) shift register | clk | q1_PFS | J1 | 17 | PFS_T | J1 | 17 | P1 | C30 | PFS_T | JH2_T | 25 | 26 |  | PFS_T | J2 | 26 | G | 1 | PFS_T | T-Q1&2 | G | cond | PFS_T | G | 1 | J2 | 26 | PFS_T | slot_1 | C27 |  |  |  | 13 | 1T | 11 | PFS_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 13 | 1T | 12 | PFS_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
10 | Q1_PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk | q1_PF1 | J1 | 19 | PF1_T | J1 | 19 | P1 | A31,B31 | PF1_T | JH2_T | 28 | 27 |  | PF1_T | J2 | 27 | K | 1 | PF1_T | T-Q1&2 | K | shield | PF1_T | K | 1 | J2 | 27 | PF1_T | slot_1 | C29 |  |  |  | 13 | 1T | 13 | PF1_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 13 | 1T | 14 | PF1_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
11 | Q1_PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk | q1_PF2 | J1 | 21 | PF2_T | J1 | 21 | P1 | C31 | PF2_T | JH2_T | 27 | 28 |  | PF2_T | J2 | 28 | J | 1 | PF2_T | T-Q1&2 | J | cond | PF2_T | J | 1 | J2 | 28 | PF2_T | slot_1 | C31 |  |  |  | 13 | 1T | 15 | PF2_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 13 | 1T | 16 | PF2_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
12 | Q1_VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias | q1_VNROW | J1 | 23 | VNROW_T | J1 | 23 | P1 | A32,B32 | VNROW_T | JH2_T | 30 | 29 |  | VNROW_T | J2 | 29 | k | 1 | VNROW_T | T-Q1&2 | k | shield | VNROW_T | k | 1 | J2 | 29 | VNROW_T | slot_7 | C30 |  |  |  | 14 | 2T | 14 | VNROW_T | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | 14 | 2T | 4 | VROWON_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
13 | Q1_VnCOL | -6.0 to -3.5 |  | negative supply for fast shift reg | bias | q1_VNCOL | J1 | 25 | VNCOL_T | J1 | 25 | P1 | C32 | VNCOL_T | JH2_T | 29 | 30 |  | VNCOL_T | J2 | 30 | j | 1 | VNCOL_T | T-Q1&2 | j | cond | VNCOL_T | j | 1 | J2 | 30 | VNCOL_T | slot_7 | C29 |  |  |  | 14 | 2T | 13 | VNCOL_T | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
14 | Q1_Vp | 0.0 to -0.5 |  | positive supply for shift reg | bias | q1_VP | J1 | 27 | VP_T | J1 | 27 | P1 | A1,B1 | VP_T | JH1_T | 2 | 1 | Cu | VP_T | J1 | 1 | Z | 1 | VP_T | T-Q1&2 | Z | cond | VP_T | Z | 1 | J1 | 1 | VP_T | AGND1 | A12S?
15 | VSS | 0 |  | unit cell source | bias | q1_VSSCM | J1 | 29 | VSSCM_T | J1 | 29 | P1 | C2 | VSSCM_T | JH1_T | 3 | 4 |  | VSSCM_T | J2 | 4 | u | 1 | VSSCM_T | T-Q1&2 | u | shield | VSSCM_T | u | 1 | J2 | 4 | VSSCM_T | AGND7 | A22s?
16 | VSUB | 0 gnd |  | readout substrate | bias | AGND | J1 | 31 | AGND | J1 | 31 | P1 | C1 | AGND_T | JH1_T | 1 | 2 | Cu | AGND_T | J1 | 2 | Y | 1 | AGND_T | T-Q1&2 | Y | shield | AGND_T | Y | 1 | J1 | 2 | AGND_T | AGND2 | A2s?
17 | VSS | 0 |  | unit cell source | bias | q3_VSSCM | J1 | 33 | VSSCM_B | J1 | 33 | P2 | C2 | VSSCM_B | JH1_B | 3 | 4 |  | VSSCM_B | J1 | 4 | u | 2 | VSSCM_B | B-Q3&4 | u | shield | VSSCM_B | u | 2 | J1 | 4 | VSSCM_B | AGND6 | A12s?
18 | Q3_Vp | 0.0 to -0.5 |  | positive supply for shift reg | bias | q3_VP | J1 | 35 | VP_B | J1 | 35 | P2 | A1,B1 | VP_B | JH1_B | 1 | 2 | Cu | VP_B | J1 | 2 | Y | 2 | VP_B | B-Q3&4 | Y | Y | VP_B | Y | 2 | J1 | 2 | VP_B | AGND4 | A2s?
19 | Q3_VnCOL | -6.0 to -3.5 |  | negative supply for fast shift reg | bias | q3_VNCOL | J1 | 37 | VNCOL_B | J1 | 37 | P2 | C32 | VNCOL_B | JH2_B | 29 | 30 |  | VNCOL_B | J2 | 30 | j | 2 | VNCOL_B | B-Q3&4 | j | cond | VNCOL_B | j | 2 | J2 | 30 | VNCOL_B | slot_14 | C29 |  |  |  | 6 | 2B | 13 | VNCOL_B | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
20 | Q3_VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias | q3_VNROW | J1 | 39 | VNROW_B | J1 | 39 | P2 | A32,B32 | VNROW_B | JH2_B | 30 | 29 |  | VNROW_B | J2 | 29 | k | 2 | VNROW_B | B-Q3&4 | k | shield | VNROW_B | k | 2 | J2 | 29 | VNROW_B | slot_14 | C30 |  |  |  | 6 | 2B | 14 | VNROW_B | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
21 | Q3_PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk | q3_PF2 | J1 | 41 | PF2_B | J1 | 41 | P2 | C31 | PF2_B | JH2_B | 27 | 28 |  | PF2_B | J2 | 28 | J | 2 | PF2_B | B-Q3&4 | J | cond | PF2_B | J | 2 | J2 | 28 | PF2_B | slot_8 | C31 |  |  |  | 5 | 1B | 15 | PF2_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 5 | 1B | 16 | PF2_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
22 | Q3_PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk | q3_PF1 | J1 | 43 | PF1_B | J1 | 43 | P2 | A31,B31 | PF1_B | JH2_B | 28 | 27 |  | PF1_B | J2 | 27 | K | 2 | PF1_B | B-Q3&4 | K | shield | PF1_B | K | 2 | J2 | 27 | PF1_B | slot_8 | C29 |  |  |  | 5 | 1B | 13 | PF1_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 5 | 1B | 14 | PF1_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
23 | Q3_PHIFS | VpROW&COL | VnCOL | starts fast (col) shift register | clk | q3_PFS | J1 | 45 | PFS_B | J1 | 45 | P2 | C30 | PFS_B | JH2_B | 25 | 26 |  | PFS_B | J2 | 26 | G | 2 | PFS_B | B-Q3&4 | G | cond | PFS_B | G | 2 | J2 | 26 | PFS_B | slot_8 | C27 |  |  |  | 5 | 1B | 11 | PFS_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 5 | 1B | 12 | PFS_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
24 | Q3_PHIDES | VpROW&COL | VnROW | row deselect | clk | q3_PDES | J1 | 47 | PDES_B | J1 | 47 | P2 | A30,B30 | PDES_B | JH2_B | 26 | 25 |  | PDES_B | J2 | 25 | H | 2 | PDES_B | B-Q3&4 | H | shield | PDES_B | H | 2 | J2 | 25 | PDES_B | slot_8 | C25 |  |  |  | 5 | 1B | 9 | PDES_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 5 | 1B | 10 | PDES_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
25 | Q3_PHIROE | VpROW&COL | VnROW | phase 3 clock | clk | q3_PROE | J1 | 49 | PROE_B | J1 | 49 | P2 | C29 | PROE_B | JH2_B | 23 | 24 |  | PROE_B | J2 | 24 | h | 2 | PROE_B | B-Q3&4 | h | cond | PROE_B | h | 2 | J2 | 24 | PROE_B | slot_8 | C23 |  |  |  | 5 | 1B | 7 | PROE_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 5 | 1B | 8 | PROE_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
26 | Q3_PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk | q3_PS2 | J1 | 51 | PS2_B | J1 | 51 | P2 | A29,B29 | PS2_B | JH2_B | 24 | 23 |  | PS2_B | J2 | 23 | i | 2 | PS2_B | B-Q3&4 | i | shield | PS2_B | i | 2 | J2 | 23 | PS2_B | slot_8 | C21 |  |  |  | 5 | 1B | 5 | PS2_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 5 | 1B | 6 | PS2_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
27 | Q3_PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk | q3_PS1 | J1 | 53 | PS1_B | J1 | 53 | P2 | C28 | PS1_B | JH2_B | 21 | 22 |  | PS1_B | J2 | 22 | z | 2 | PS1_B | B-Q3&4 | z | cond | PS1_B | z | 2 | J2 | 22 | PS1_B | slot_8 | C19 |  |  |  | 5 | 1B | 3 | PS1_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 5 | 1B | 4 | PS1_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
28 | Q3_PHISS | VpROW&COL | VnROW | starts slow (row) shift register | clk | q3_PSS | J1 | 55 | PSS_B | J1 | 55 | P2 | A28,B28 | PSS_B | JH2_B | 22 | 21 |  | PSS_B | J2 | 21 | AA | 2 | PSS_B | B-Q3&4 | AA | shield | PSS_B | AA | 2 | J2 | 21 | PSS_B | slot_14 | C31 |  |  |  | 6 | 2B | 15 | PSS_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 6 | 2B | 16 | PSS_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
29 | Q3_VDDUC | -4.0 to -3.5 |  | unit cell drain | bias | q3_VDDUC | J1 | 57 | VDDUC_B | J1 | 57 | P2 | C27 | VDDUC_B | JH2_B | 19 | 20 |  | VDDUC_B | J2 | 20 | KK | 2 | VDDUC_B | B-Q3&4 | KK | cond | VDDUC_B | KK | 2 | J2 | 20 | VDDUC_B | slot_14 | C28
30 | Q3_VDDCL | -1.2 | VDDUC | column line overload protection | clk | q3_VDDCL | J1 | 59 | VDDCL_B | J1 | 59 | P2 | A27,B27 | VDDCL_B | JH2_B | 20 | 19 |  | VDDCL_B | J2 | 19 | LL | 2 | VDDCL_B | B-Q3&4 | LL | shield | VDDCL_B | LL | 2 | J2 | 19 | VDDCL_B | slot_8 | C17 |  |  |  | 5 | 1B | 1 | VDDCL_B_H | -4.5 | -922 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | ON | OFF | ON | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | 5 | 1B | 2 | VDDCL_B_L | -1.3 | -267 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | ON | OFF | OFF | ON | OFF | OFF | OFF | OFF | ON | OFF | ON | OFF
31 | Q3_VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk | q3_VGGCL | J1 | 60 | VGGCL_B | J1 | 60 | P2 | A2,B2 | VGGCL_B | JH1_B | 4 | 3 |  | VGGCL_B | J1 | 3 | a | 2 | VGGCL_B | B-Q3&4 | a | cond | VGGCL_B | a | 2 | J1 | 3 | VGGCL_B | slot_14 | C17 |  |  |  | 6 | 2B | 1 | VGGCL_B_H | -3.75 | -768 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ON | OFF | ON | OFF | ON | ON | ON | ON | ON | ON | ON | ON | 6 | 2B | 2 | VGGCL_B_L | -4.5 | -922 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | ON | OFF | ON | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON
32 | Q3_TEND |  |  | diagnostic |  | q3_TEND | J2 | 1 | Q3_TEND | J2 | 1 | P2 | C15 | Q3_TEND | JH1_B | 29 | 30 |  | Q3_TEND | J1 | 30 | nc | 2 |  |  |  |  | Q3_TEND | nc | 2 | J1 | 30 | Q3_TEND
33 | Q3_VRSTG | -3.5 | -5.6 | global reset | clk | q3_VRSTG | J2 | 3 | VRSTG_B | J2 | 3 | P2 | C3 | VRSTG_B | JH1_B | 5 | 6 |  | VRSTG_B | J1 | 6 | FF | 2 | VRSTG_B | B-Q3&4 | FF | cond | VRSTG_B | FF | 2 | J1 | 6 | VRSTG_B | slot_14 | C23 |  |  |  | 6 | 2B | 7 | VRSTG_B_H | -2 | -410 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | 6 | 2B | 8 | VRSTG_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
34 | Q3_VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position |  | q3_VROWOFF | J2 | 5 | VROWOFF_B | J2 | 5 | P2 | A4,B4 | VROWOFF_B | JH1_B | 8 | 7 |  | VROWOFF_B | J1 | 7 | t | 2 | VROWOFF_B | B-Q3&4 | t | shield | VROWOFF_B | t | 2 | J1 | 7 | VROWOFF_B | AGND5 | A22s?
35 | Q3_IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource | q3_IREF | J2 | 7 | VIREF_B | J2 | 7 | P2 | C4 | VIREF_B | JH1_B | 7 | 8 |  | VIREF_B | J1 | 8 | s | 2 | VIREF_B | B-Q3&4 | s | cond | VIREF_B | s | 2 | J1 | 8 | VIREF_B | slot_14 | C25 |  |  |  | 6 | 2B | 9 | VIREF_B | -1.8 | -369 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | ON | OFF | OFF | ON | OFF | ON | ON | ON | OFF | OFF | OFF | OFF
36 | Q3_VROWON | -6.0 to -5.6 |  | unit cell row enable on position |  | q3_VROWON | J2 | 9 | VROWON_B | J2 | 9 | P2 | A3,B3 | VROWON_B | JH1_B | 6 | 5 |  | VROWON_B | J1 | 5 | GG | 2 | VROWON_B | B-Q3&4 | GG | shield | VROWON_B | GG | 2 | J1 | 5 | VROWON_B | slot_14 | C19 |  |  |  | 6 | 2B | 3 | VROWON_B_H | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | 6 | 2B | 4 | VROWON_B_L | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
37 | Q3_VRSTR | -3.5 | -5.6 | row reset | clk | q3_VRSTR | J2 | 11 | VRSTR_B | J2 | 11 | P2 | A5,B5 | VRSTR_B | JH1_B | 10 | 9 |  | VRSTR_B | J1 | 9 | MM | 2 | VRSTR_B | B-Q3&4 | MM | cond | VRSTR_B | MM | 2 | J1 | 9 | VRSTR_B | slot_14 | C21 |  |  |  | 6 | 2B | 5 | VRSTR_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 6 | 2B | 6 | VRSTR_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
38 | Q3_OUT1 |  |  | p channel output 1 | output | q3_OUT1 | J2 | 13 | Q3OUT1+ | J2 | 13 | P2 | C6 | Q3OUT1+ | JH1_B | 11 | 12 |  | Q3OUT1+ | J1 | 12 | DD | 2 | Q3OUT1+ | B-Q3&4 | DD | cond | Q3OUT1+ | DD | 2 | J1 | 12 | Q3OUT1+ | slot_13 | C3
39 | Q3_OUT2 |  |  | p channel output 2 | output | q3_OUT2 | J2 | 15 | Q3OUT2+ | J2 | 15 | P2 | C7 | Q3OUT2+ | JH1_B | 13 | 14 |  | Q3OUT2+ | J1 | 14 | W | 2 | Q3OUT2+ | B-Q3&4 | W | cond | Q3OUT2+ | W | 2 | J1 | 14 | Q3OUT2+ | slot_13 | C5
40 | Q3_OUT3 |  |  | p channel output 3 | output | q3_OUT3 | J2 | 17 | Q3OUT3+ | J2 | 17 | P2 | C8 | Q3OUT3+ | JH1_B | 15 | 16 |  | Q3OUT3+ | J1 | 16 | q | 2 | Q3OUT3+ | B-Q3&4 | q | cond | Q3OUT3+ | q | 2 | J1 | 16 | Q3OUT3+ | slot_13 | C7
41 | Q3_OUT4 |  |  | p channel output 4 | output | q3_OUT4 | J2 | 19 | Q3OUT4+ | J2 | 19 | P2 | C9 | Q3OUT4+ | JH1_B | 17 | 18 |  | Q3OUT4+ | J1 | 18 | U | 2 | Q3OUT4+ | B-Q3&4 | U | cond | Q3OUT4+ | U | 2 | J1 | 18 | Q3OUT4+ | slot_13 | C9
42 | Q3_OUT5 |  |  | p channel output 5 | output | q3_OUT5 | J2 | 21 | Q3OUT5+ | J2 | 21 | P2 | C10 | Q3OUT5+ | JH1_B | 19 | 20 |  | Q3OUT5+ | J1 | 20 | S | 2 | Q3OUT5+ | B-Q3&4 | S | cond | Q3OUT5+ | S | 2 | J1 | 20 | Q3OUT5+ | slot_13 | C11
43 | Q3_OUT6 |  |  | p channel output 6 | output | q3_OUT6 | J2 | 23 | Q3OUT6+ | J2 | 23 | P2 | C11 | Q3OUT6+ | JH1_B | 21 | 22 |  | Q3OUT6+ | J1 | 22 | n | 2 | Q3OUT6+ | B-Q3&4 | n | cond | Q3OUT6+ | n | 2 | J1 | 22 | Q3OUT6+ | slot_13 | C13
44 | Q3_OUT7 |  |  | p channel output 7 | output | q3_OUT7 | J2 | 25 | Q3OUT7+ | J2 | 25 | P2 | C12 | Q3OUT7+ | JH1_B | 23 | 24 |  | Q3OUT7+ | J1 | 24 | BB | 2 | Q3OUT7+ | B-Q3&4 | BB | cond | Q3OUT7+ | BB | 2 | J1 | 24 | Q3OUT7+ | slot_13 | C15
45 | Q3_OUT8 |  |  | p channel output 8 | output | q3_OUT8 | J2 | 27 | Q3OUT8+ | J2 | 27 | P2 | C13 | Q3OUT8+ | JH1_B | 25 | 26 |  | Q3OUT8+ | J1 | 26 | P | 2 | Q3OUT8+ | B-Q3&4 | P | cond | Q3OUT8+ | P | 2 | J1 | 26 | Q3OUT8+ | slot_13 | C17
46 | Q3_VDDOUT | -1.0 to -1.3 |  | output drain | bias | q3_VDDOUT | J2 | 29 | VDDOUT_B | J2 | 29 | P2 | C14 | VDDOUT_B | JH1&2 | 27_1 | 28_2 | _Cu | VDDOUT_B | J1&2 | 28_2 | N_A | 2 | VDDOUT_B | B-Q3&4 | N_A | c_s | VDDOUT_B | N_A | 2 | J1&2 | 28_2 | VDDOUT_B | slot_14 | C26 |  |  |  | 6 | 2B | 10 | VDDOUT_B | -1.3 | -267 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | ON | OFF | OFF | ON | OFF | OFF | OFF | OFF | ON | OFF | ON | OFF
47 | Q3_VDETCOM | -2.8 to-3.4 |  | detector common supply | bias | q3_VDETCOM | J2 | 31 | VDETCOM | J2 | 31 | P2 | A18,B18 | VDETCOM_B | JH2_B | 2 | 1 | Cu | VDETCOM_B | J2 | 1 | B | 2 | VDETCOM_B | B-Q3&4 | B | cond | VDETCOM_B | B | 2 | J2 | 1 | VDETCOM_B | slot_14 | C27 |  |  |  | 6 | 2B | 11 | VDETCOM_B | -3.75 | -768 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ON | OFF | ON | OFF | ON | ON | ON | ON | ON | ON | ON | ON
48 | Q4_VDDOUT | -1.0 to -1.3 |  | output drain | bias | q3_VDDOUT | J2 | 33 | VDDOUT_B | J2 | 33 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | B-Q3&4 | * | * | * | * | * | * | * | *
49 | Q4_OUT8 |  |  | p channel output 8 | output | q4_OUT8 | J2 | 35 | Q4OUT8+ | J2 | 35 | P2 | C19 | Q4OUT8+ | JH2_B | 3 | 4 |  | Q4OUT8+ | J2 | 4 | b | 2 | Q4OUT8+ | B-Q3&4 | b | cond | Q4OUT8+ | b | 2 | J2 | 4 | Q4OUT8+ | slot_10 | C17
50 | Q4_OUT7 |  |  | p channel output 7 | output | q4_OUT7 | J2 | 37 | Q4OUT7+ | J2 | 37 | P2 | C20 | Q4OUT7+ | JH2_B | 5 | 6 |  | Q4OUT7+ | J2 | 6 | v |  | Q4OUT7+ | B-Q3&4 | v | cond | Q4OUT7+ | v |  | J2 | 6 | Q4OUT7+ | slot_10 | C15
51 | Q4_OUT6 |  |  | p channel output 6 | output | q4_OUT6 | J2 | 39 | Q4OUT6+ | J2 | 39 | P2 | C21 | Q4OUT6+ | JH2_B | 7 | 8 |  | Q4OUT6+ | J2 | 8 | d |  | Q4OUT6+ | B-Q3&4 | d | cond | Q4OUT6+ | d |  | J2 | 8 | Q4OUT6+ | slot_10 | C13
52 | Q4_OUT5 |  |  | p channel output 5 | output | q4_OUT5 | J2 | 41 | Q4OUT5+ | J2 | 41 | P2 | C22 | Q4OUT5+ | JH2_B | 9 | 10 |  | Q4OUT5+ | J2 | 10 | HH |  | Q4OUT5+ | B-Q3&4 | HH | cond | Q4OUT5+ | HH |  | J2 | 10 | Q4OUT5+ | slot_10 | C11
53 | Q4_OUT4 |  |  | p channel output 4 | output | q4_OUT4 | J2 | 43 | Q4OUT4+ | J2 | 43 | P2 | C23 | Q4OUT4+ | JH2_B | 11 | 12 |  | Q4OUT4+ | J2 | 12 | x |  | Q4OUT4+ | B-Q3&4 | x | cond | Q4OUT4+ | x |  | J2 | 12 | Q4OUT4+ | slot_10 | C9
54 | Q4_OUT3 |  |  | p channel output 3 | output | q4_OUT3 | J2 | 45 | Q4OUT3+ | J2 | 45 | P2 | C24 | Q4OUT3+ | JH2_B | 13 | 14 |  | Q4OUT3+ | J2 | 14 | C |  | Q4OUT3+ | B-Q3&4 | C | cond | Q4OUT3+ | C |  | J2 | 13 | Q4OUT3+ | slot_10 | C7
55 | Q4_OUT2 |  |  | p channel output 2 | output | q4_OUT2 | J2 | 47 | Q4OUT2+ | J2 | 47 | P2 | C25 | Q4OUT2+ | JH2_B | 15 | 16 |  | Q4OUT2+ | J2 | 16 | f |  | Q4OUT2+ | B-Q3&4 | f | cond | Q4OUT2+ | f |  | J2 | 16 | Q4OUT2+ | slot_10 | C5
56 | Q4_OUT1 |  |  | p channel output 1 | output | q4_OUT1 | J2 | 49 | Q4OUT1+ | J2 | 49 | P2 | C26 | Q4OUT1+ | JH2_B | 17 | 18 |  | Q4OUT1+ | J2 | 18 | E |  | Q4OUT1+ | B-Q3&4 | E | cond | Q4OUT1+ | E |  | J2 | 18 | Q4OUT1+ | slot_10 | C3
57 | Q4_VRSTR | -3.5 | -5.6 | row reset | clk | q4_VRSTR | J2 | 51 | VRSTR_B | J2 | 51 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
58 | Q4_VROWON | -6.0 to -5.6 |  | unit cell row enable on position |  | q4_VROWON | J2 | 53 | VROWON_B | J2 | 53 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
59 | Q4_IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource | q4_IREF | J2 | 55 | VIREF_B | J2 | 55 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
60 | Q4_VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position |  | q4_VROWOFF | J2 | 57 | VROWOFF_B | J2 | 57 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
61 | Q4_VRSTG | -3.5 | -5.6 | global reset | clk | q4_VRSTG | J2 | 59 | VRSTG_B | J2 | 59 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
62 | Q4_TEND |  |  | diagnostic |  | q4_TEND | J2 | 60 | Q4_TEND | J2 | 60 | P2 | C17 | Q4_TEND | JH2_B | 33 | 34 |  | Q4_TEND | J2 | 34 | NC | [5 empty] | Q4_TEND | NC |  | J2 | 34 | Q4_TEND
63 | Q4_VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk | q4_VGGCL | J3 | 1 | VGGCL_B | J3 | 1 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
64 | Q4_VDDCL | -1.2 | VDDUC | column line overload protection | clk | q4_VDDCL | J3 | 3 | VDDCL_B | J3 | 3 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
65 | Q4_VDDUC | -4.0 to -3.5 |  | unit cell drain | bias | Q4_VDDUC | J3 | 5 | VDDUC_B | J3 | 5 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
66 | Q4_PHISS | VpROW&COL | VnROW | starts slow (row) shift register | clk | q4_PSS | J3 | 7 | PSS_B | J3 | 7 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
67 | Q4_PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk | q4_PS1 | J3 | 9 | PS1_B | J3 | 9 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
68 | Q4_PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk | q4_PS2 | J3 | 11 | PS2_B | J3 | 11 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
69 | Q4_PHIROE | VpROW&COL | VnROW | phase 3 clock | clk | q4_PROE | J3 | 13 | PROE_B | J3 | 13 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
70 | Q4_PHIDES | VpROW&COL | VnROW | resets detector bias | clk | q4_PDES | J3 | 15 | PDES_B | J3 | 15 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
71 | Q4_PHIFS | VpROW&COL | VnCOL | starts fast (col) shift register | clk | q4_PFS | J3 | 17 | PFS_B | J3 | 17 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
72 | Q4_PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk | q4_PF1 | J3 | 19 | PF1_B | J3 | 19 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
73 | Q4_PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk | q4_PF2 | J3 | 21 | PF2_B | J3 | 21 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
74 | Q4_VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias | q4_VNROW | J3 | 23 | VNROW_B | J3 | 23 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
75 | Q4_VnCOL | -6.0 to -3.5 |  | negative supply for fast shift reg | bias | q4_VNCOL | J3 | 25 | VNCOL_B | J3 | 25 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
76 | Q4_VP | 0.0 to -0.5 |  | positive supply for shift reg | bias | q4_VP | J3 | 27 | VP_B | J3 | 27 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
77 | VSS | 0 |  | unit cell source | bias | q4_VSSCM | J3 | 29 | VSSCM_B | J3 | 29 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
78 | VSUB | 0 gnd |  | readout substrate | bias | AGND | J3 | 31 | AGND | J3 | 31 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
79 | VSS | 0 |  | unit cell source | bias | q2_VSSCM | J3 | 33 | VSSCM_T | J3 | 33 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
80 | Q2_Vp | 0.0 to -0.5 |  | positive supply for shift reg | bias | q2_VP | J3 | 35 | VP_T | J3 | 35 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
81 | Q2_VnCOL | -6.0 to -3.5 |  | negative supply for fast shift reg | bias | q2_VNCOL | J3 | 37 | VNCOL_T | J3 | 37 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
82 | Q2_VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias | q2_VNROW | J3 | 39 | VNROW_T | J3 | 39 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
83 | Q2_PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk | q2_PF2 | J3 | 41 | PF2_T | J3 | 41 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
84 | Q2_PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk | q2_PF1 | J3 | 43 | PF1_T | J3 | 43 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
85 | Q2_PHIFS | VpROW&COL | VnCOL | starts fast (col) shift register | clk | q2_PFS | J3 | 45 | PFS_T | J3 | 45 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
86 | Q2_PHIDES | VpROW&COL | VnROW | row deselect | clk | q2_PDES | J3 | 47 | PDES_T | J3 | 47 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
87 | Q2_PHIROE | VpROW&COL | VnROW |  | clk | q2_PROE | J3 | 49 | PROE_T | J3 | 49 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
88 | Q2_PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk | q2_PHIS2 | J3 | 51 | PS2_T | J3 | 51 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
89 | Q2_PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk | q2_PS1 | J3 | 53 | PS1_T | J3 | 53 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
90 | Q2_PHISS | VpROW&COL | VnROW | starts slow (row) shift register | clk | q2_PSS | J3 | 55 | PSS_T | J3 | 55 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
91 | Q2_VDDUC | -4.0 to -3.5 |  | unit cell drain | bias | q2_VDDUC | J3 | 57 | VDDUC_T | J3 | 57 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
92 | Q2_VDDCL | -1.2 | VDDUC | column line overload protection | clk | q2_VDDCL | J3 | 59 | VDDCL_T | J3 | 59 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
93 | Q2_VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk | q2_VGGCL | J3 | 60 | VDDCL_T | J3 | 60 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
94 | Q2_TEND |  |  | diagnostic |  | q2_TEND | J4 | 1 | Q2TEND | J4 | 1 | P1 | C17 | Q2TEND | JH1_T | 33 | 34 |  | Q2TEND | [8 empty] | Q2TEND |  |  |  |  | Q2TEND
95 | Q2_VRSTG | -3.5 | -5.6 | global reset | clk | q2_VRSTG | J4 | 3 | VRSTG_T | J4 | 3 | P1 | C3 | VRSTG_T | JH1_T | 5 | 6 |  | VRSTG_T | J1 | 6 | FF | 1 | VRSTG_T | T-Q1&2 | FF | cond | VRSTG_T | FF | 1 | J1 | 6 | VRSTG_T | slot_7 | C23 |  |  |  | 14 | 2T | 7 | VRSTG_T_H | -2 | -410 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | 14 | 2T | 8 | VRSTG_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
96 | Q2_VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position |  | q2_VROWOFF | J4 | 5 | VROWOFF_T | J4 | 5 | P1 | A4,B4 | VROWOFF_T | JH1_T | 5 | 5 |  | VROWOFF_T | J1 | 5 | t | 1 | VROWOFF_T | T-Q1&2 | t | shield | VROWOFF_T | t | 1 | J1 | 5 | VROWOFF_T | AGND3
97 | Q2_IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource | q2_IREF | J4 | 7 | VIREF_T | J4 | 7 | P1 | C4 | VIREF_T | JH1_T | 7 | 8 |  | VIREF_T | J1 | 8 | s | 1 | VIREF_T | T-Q1&2 | s | cond | VIREF_T | s | 1 | J1 | 8 | VIREF_T | slot_7 | C25 |  |  |  | 14 | 2T | 9 | VIREF_T | -1.8 | -369 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | ON | OFF | OFF | ON | OFF | ON | ON | ON | OFF | OFF | OFF | OFF
98 | Q2_VROWON | -6.0 to -5.6 |  | unit cell row enable on position |  | q2_VROWON | J4 | 9 | VROWON_T | J4 | 9 | P1 | A3,B3 | VROWON_T | JH1_T | 8 | 7 |  | VROWON_T | J1 | 7 | GG | 1 | VROWON_T | T-Q1&2 | GG | shield | VROWON_T | GG | 1 | J1 | 7 | VROWON_T | slot_7 | C19
99 | Q2_VRSTR | -3.5 | -5.6 | unit cell reset on position = row reset | clk | q2_VRSTR | J4 | 11 | VRSTR_T | J4 | 11 | P1 | A5,B5 | VRSTR_T | JH1_T | 10 | 9 |  | VRSTR_T | J1 | 9 | MM | 1 | VRSTR_T | T-Q1&2 | MM | shield | VRSTR_T | MM | 1 | J1 | 9 | VRSTR_T | slot_7 | C21 |  |  |  | 14 | 2T | 5 | VRSTR_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 14 | 2T | 6 | VRSTR_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
100 | Q2_OUT1 |  |  | p channel output 1 | output | q2_OUT1 | J4 | 13 | Q2OUT1+ | J4 | 13 | P1 | C26 | Q2OUT1+ | JH2_T | 17 | 18 |  | Q2OUT1+ | J2 | 18 | E | 1 | Q2OUT1+ | T-Q1&2 | E | cond | Q2OUT1+ | E | 1 | J2 | 18 | Q2OUT1+ | slot_3 | C3
101 | Q2_OUT2 |  |  | p channel output 2 | output | q2_OUT2 | J4 | 15 | Q2OUT2+ | J4 | 15 | P1 | C25 | Q2OUT2+ | JH2_T | 15 | 16 |  | Q2OUT2+ | J2 | 16 | f | 1 | Q2OUT2+ | T-Q1&2 | f | cond | Q2OUT2+ | f | 1 | J2 | 16 | Q2OUT2+ | slot_3 | C5
102 | Q2_OUT3 |  |  | p channel output 3 | output | q2_OUT3 | J4 | 17 | Q2OUT3+ | J4 | 17 | P1 | C24 | Q2OUT3+ | JH2_T | 13 | 14 |  | Q2OUT3+ | J2 | 14 | C | 1 | Q2OUT3+ | T-Q1&2 | C | cond | Q2OUT3+ | C | 1 | J2 | 14 | Q2OUT3+ | slot_3 | C7
103 | Q2_OUT4 |  |  | p channel output 4 | output | q2_OUT4 | J4 | 19 | Q2OUT4+ | J4 | 19 | P1 | C23 | Q2OUT4+ | JH2_T | 11 | 12 |  | Q2OUT4+ | J2 | 12 | x | 1 | Q2OUT4+ | T-Q1&2 | x | cond | Q2OUT4+ | x | 1 | J2 | 12 | Q2OUT4+ | slot_3 | C9
104 | Q2_OUT5 |  |  | p channel output 5 | output | q2_OUT5 | J4 | 21 | Q2OUT5+ | J4 | 21 | P1 | C22 | Q2OUT5+ | JH2_T | 9 | 10 |  | Q2OUT5+ | J2 | 10 | HH | 1 | Q2OUT5+ | T-Q1&2 | HH | cond | Q2OUT5+ | HH | 1 | J2 | 10 | Q2OUT5+ | slot_3 | C11
105 | Q2_OUT6 |  |  | p channel output 6 | output | q2_OUT6 | J4 | 23 | Q2OUT6+ | J4 | 23 | P1 | C21 | Q2OUT6+ | JH2_T | 7 | 8 |  | Q2OUT6+ | J2 | 8 | d | 1 | Q2OUT6+ | T-Q1&2 | d | cond | Q2OUT6+ | d | 1 | J2 | 8 | Q2OUT6+ | slot_3 | C13
106 | Q2_OUT7 |  |  | p channel output 7 | output | q2_OUT7 | J4 | 25 | Q2OUT7+ | J4 | 25 | P1 | C20 | Q2OUT7+ | JH2_T | 5 | 6 |  | Q2OUT7+ | J2 | 6 | v | 1 | Q2OUT7+ | T-Q1&2 | v | cond | Q2OUT7+ | v | 1 | J2 | 6 | Q2OUT7+ | slot_3 | C15
107 | Q2_OUT8 |  |  | p channel output 8 | output | q2_OUT8 | J4 | 27 | Q2OUT8+ | J4 | 27 | P1 | C19 | Q2OUT8+ | JH2_T | 3 | 4 |  | Q2OUT8+ | J2 | 4 | b | 1 | Q2OUT8+ | T-Q1&2 | b | cond | Q2OUT8+ | b | 1 | J2 | 4 | Q2OUT8+ | slot_3 | C17
108 | Q2_VDDOUT | -1.0 to -1.3 |  | output drain | bias | q2_VDDOUT | J4 | 29 | VDDOUT_T | J4 | 29 | P1 | C14,?C18 | VDDOUT_T | JH1&2 | 27_1 | 28_2 | _Cu | VDDOUT_T | J1&2 | 28_2 | N_A | 1 | VDDOUT_T | T-Q1&2 | N_A | c_s | VDDOUT_T | N_A | 1 | J1&2 | 28_2 | VDDOUT_T | slot_7 | C26 |  |  |  | 14 | 2T | 10 | VDDOUT_T | -1.3 | -267 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | ON | OFF | OFF | ON | OFF | OFF | OFF | OFF | ON | OFF | ON | OFF
109 | VDETCOM | -2.8 to-3.4 |  | detector common supply | bias | VDETCOM | J4 | 31 | VDETCOM | J4 | 31 | P1 | A18,B18 | VDETCOM | JH2_T | 2 | 1 | Cu | VDETCOM | J2 | 1 | B | 1 | VDETCOM | T-Q1&2 | B | cond | VDETCOM | B | 1 | J2 | 1 | VDETCOM | slot_7 | C27 |  |  |  | 14 | 2T | 11 | VDETCOM_T | -3.75 | -768 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ON | OFF | ON | OFF | ON | ON | ON | ON | ON | ON | ON | ON
110 | Q1_VDDOUT | -1.0 to -1.3 |  | output drain | bias | q1_VDDOUT | J4 | 33 | VDDOUT_T | J4 | 33 | P1 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
111 | Q1_OUT8 |  |  | p channel output 8 | output | q1_OUT8 | J4 | 35 | Q1OUT8+ | J4 | 35 | P1 | C13 | Q1OUT8+ | JH1_T | 25 | 26 |  | Q1OUT8+ | J1 | 26 | P | 1 | Q1OUT8+ | T-Q1&2 | P | cond | Q1OUT8+ | P | 1 | J1 | 26 | Q1OUT8+ | slot_6 | C17
112 | Q1_OUT7 |  |  | p channel output 7 | output | q1_OUT7 | J4 | 37 | Q1OUT7+ | J4 | 37 | P1 | C12 | Q1OUT7+ | JH1_T | 23 | 24 |  | Q1OUT7+ | J1 | 24 | BB | 1 | Q1OUT7+ | T-Q1&2 | BB | cond | Q1OUT7+ | BB | 1 | J1 | 24 | Q1OUT7+ | slot_6 | C15
113 | Q1_OUT6 |  |  | p channel output 6 | output | q1_OUT6 | J4 | 39 | Q1OUT6+ | J4 | 39 | P1 | C11 | Q1OUT6+ | JH1_T | 21 | 22 |  | Q1OUT6+ | J1 | 22 | n | 1 | Q1OUT6+ | T-Q1&2 | n | cond | Q1OUT6+ | n | 1 | J1 | 22 | Q1OUT6+ | slot_6 | C13
114 | Q1_OUT5 |  |  | p channel output 5 | output | q1_OUT5 | J4 | 41 | Q1OUT5+ | J4 | 41 | P1 | C10 | Q1OUT5+ | JH1_T | 19 | 20 |  | Q1OUT5+ | J1 | 20 | S | 1 | Q1OUT5+ | T-Q1&2 | S | cond | Q1OUT5+ | S | 1 | J1 | 20 | Q1OUT5+ | slot_6 | C11
115 | Q1_OUT4 |  |  | p channel output 4 | output | q1_OUT4 | J4 | 43 | Q1OUT4+ | J4 | 43 | P1 | C9 | Q1OUT4+ | JH1_T | 17 | 18 |  | Q1OUT4+ | J1 | 18 | U | 1 | Q1OUT4+ | T-Q1&2 | U | cond | Q1OUT4+ | U | 1 | J1 | 18 | Q1OUT4+ | slot_6 | C9
116 | Q1_OUT3 |  |  | p channel output 3 | output | q1_OUT3 | J4 | 45 | Q1OUT3+ | J4 | 45 | P1 | C8 | Q1OUT3+ | JH1_T | 15 | 16 |  | Q1OUT3+ | J1 | 16 | q | 1 | Q1OUT3+ | T-Q1&2 | q | cond | Q1OUT3+ | q | 1 | J1 | 16 | Q1OUT3+ | slot_6 | C7
117 | Q1_OUT2 |  |  | p channel output 2 | output | q1_OUT2 | J4 | 47 | Q1OUT2+ | J4 | 47 | P1 | C7 | Q1OUT2+ | JH1_T | 13 | 14 |  | Q1OUT2+ | J1 | 14 | W | 1 | Q1OUT2+ | T-Q1&2 | W | cond | Q1OUT2+ | W | 1 | J1 | 14 | Q1OUT2+ | slot_6 | C5
118 | Q1_OUT1 |  |  | p channel output 1 | output | q1_OUT1 | J4 | 49 | Q1OUT1+ | J4 | 49 | P1 | C6 | Q1OUT1+ | JH1_T | 11 | 12 |  | Q1OUT1+ | J1 | 12 | DD | 1 | Q1OUT1+ | T-Q1&2 | DD | cond | Q1OUT1+ | DD | 1 | J1 | 12 | Q1OUT1+ | slot_6 | C3
119 | Q1_VRSTR | -3.5 | -5.6 | row reset | clk | q1_VRSTR | J4 | 51 | VRSTR_T | J4 | 51 | P1 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
120 | Q1_VROWON | -6.0 to -5.6 |  | unit cell row enable on position |  | q1_VROWON | J4 | 53 | VROWON_T | J4 | 53 | P1 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
121 | Q1_IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource | q1_IREF | J4 | 55 | VIREF_T | J4 | 55 | P1 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
122 | Q1_VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position |  | q1_VROWOFF | J4 | 57 | VROWOFF_T | J4 | 57 | P1 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
123 | Q1_VRSTG | -3.5 | -5.6 | global reset | clk | q1_VRSTG | J4 | 59 | VRSTG_T | J4 | 59 | P1 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
124 | Q1_TEND |  |  | diagnostic |  | q1_TEND | J4 | 60 | Q1TEND | J4 | 60 | P1 | C15 | Q1TEND | JH1_T | 30 | 29 |  | Q1TEND | J1 | 29 | PP | 1 | Q1TEND | T-Q1&2 | PP | PP | Q1TEND | PP | 1 | J1 | 29 | Q1TEND
[42 empty] | high level | low level | function |  | test
[42 empty] | 0 gnd |  | readout substrate | bias | 0
[42 empty] | -2.8 to-3.4 |  | detector common supply/substrate | bias | 0
[42 empty] | -4.0 to -3.5 |  | unit cell drain ~8mA | bias | -4
[42 empty] | 0 |  | unit cell source | bias | 0
[42 empty] | -3.5 | -5.6 | global reset | aclk | -3.75
[42 empty] | -3.5 | -5.6 | row reset | aclk | -6
[42 empty] | 0.0 to -0.7 |  | unit cell row enable off position |  | 0
[42 empty] | -6.0 to -5.6 |  | unit cell row enable on position | aclk | -6
100 to 300uA I load
[42 empty] | -1.0 to -1.3 |  | output drain ~16mA |  | -1.3
[42 empty] | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource | 250K
[42 empty] | 0.0 to -0.5 |  | positive supply for shift reg | bias | 0
[42 empty] | -6.0 to -3.5 |  | negative supply for fast shift reg | bias | -6
[42 empty] | -6.0 to -5.6 |  | negative supply for row shift reg | bias | -6
[44 empty] | 10Kohm to gnd shift reg diag |  | -5
[42 empty] | VpROW&COL | VnROW | starts slow (row) shift register | clk | 0/-6
[42 empty] | VpROW&COL | VnROW | phase 1 clock | clk | 0/-6
[42 empty] | VpROW&COL | VnROW | phase 2 clock | clk | 0/-6
[42 empty] | VpROW&COL | VnROW | odd/even clock | clk | 0/-6
[42 empty] | VpROW&COL | VnCOL | starts fast (col) shift register | clk | 0/-6
[42 empty] | VpROW&COL | VnCOL | phase1 clock | clk | 0/-6
[42 empty] | VpROW&COL | VnCOL | phase 2 clock | clk | 0/-6
[42 empty] | VpROW&COL | VnROW | row deselect | clk
[42 empty] | -3.5 | -5.6 | global reset | clk | -3.75
[42 empty] | -3.5 | -5.6 | row reset | clk | -6
[42 empty] | -6.0 to -5.6 |  | unit cell row enable | aclk | -6
[42 empty] | VDETCOM | -5 | column bus gate voltage | aclk | 0
[42 empty] | -1.2 | VDDUC | column bus drain voltage | aclk | 0
1
31
32
124
quad 1?
quad 3
first pixel
first pixel
last pixels
slow mux
fast mux
fast mux

## signal list P1P2(array) (2)
ALLADIN 2 pinout (SB0152 ROIC PCHANNEL OUTPUTS) | [5 empty] | 700-150 |  |  | 700-151 |  |  |  |  | xxx |  |  |  |  | 700-149 |  |  |  |  | 600-010 |  |  |  | 700-148 |  |  |  |  | Backplane |  |  | 700-148 | [7 empty] | 2047
socket Yamaichi IC51-1244-410-1 | [5 empty] | Array Fanout |  |  | Config bd | J1,2,3,4= 2mm 60 pin header |  | P1,2=96 pin |  | Thermal feedthrough |  | COLD | WARM |  | 61 pin Bd | J1,2=34 pin headers |  |  |  | 61 pin CABLES | conn=MS3116-E24-61S | conn=MS3116-E24-61S |  | 60 pin Bd |  |  | J1,2=34 pin headers |  |  |  |  | ClockBias Bd | [7 empty] | 0.0048828
pin num | signal name | high level | low level | function |  | signal | conn | pin | signal | conn | pin | conn | pin | signal |  | pin | pin | Cu? | signal | conn | pin | pin | bd# | signal | cable# | pin | cond/shld | signal | pin | bd# | conn | pin | signal | slot# | pin | DAC#_high | DAC#_low | BD# | slot | bd | DAC# | signal_high | volts | [13 empty] | sw1 | sw2 | sw3 | sw4 | sw5 | sw6 | sw7 | sw8 | sw9 | sw10 | jhxx | jhxx | slot | bd | DAC# | signal_low | volts | [13 empty] | sw1 | sw2 | sw3 | sw4 | sw5 | sw6 | sw7 | sw8 | sw9 | sw10 | jhxx | jhxx
110 | Q1_VDDOUT | -1.0 to -1.3 |  | output drain | bias | q1_VDDOUT | J4 | 33 | VDDOUT_T | J4 | 33 | P1 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
119 | Q1_VRSTR | -3.5 | -5.6 | row reset | clk | q1_VRSTR | J4 | 51 | VRSTR_T | J4 | 51 | P1 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
120 | Q1_VROWON | -6.0 to -5.6 |  | unit cell row enable on position |  | q1_VROWON | J4 | 53 | VROWON_T | J4 | 53 | P1 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
121 | Q1_IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource | q1_IREF | J4 | 55 | VIREF_T | J4 | 55 | P1 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
122 | Q1_VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position |  | q1_VROWOFF | J4 | 57 | VROWOFF_T | J4 | 57 | P1 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
123 | Q1_VRSTG | -3.5 | -5.6 | global reset | clk | q1_VRSTG | J4 | 59 | VRSTG_T | J4 | 59 | P1 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
14 | Q1_Vp | 0.0 to -0.5 |  | positive supply for shift reg | bias | q1_VP | J1 | 27 | VP_T | J1 | 27 | P1 | A1,B1 | VP_T | JH1 | 2 | 1 | Cu | VP_T | J1 | 1 | Z | 1 | VP_T | T-Q1&2 | Z | cond | VP_T | Z | 1 | J1 | 1 | VP_T | AGND1
109 | VDETCOM | -2.8 to-3.4 |  | detector common supply | bias | VDETCOM | J4 | 31 | VDETCOM | J4 | 31 | P1 | A18,B18 | VDETCOM | JH2 | 2 | 1 | Cu | VDETCOM | J2 | 1 | A | 1 | VDETCOM | T-Q1&2 | A | cond | VDETCOM | A | 1 | J2 | 1 | VDETCOM | slot 14 | C27 |  |  |  | 14 | 2T | 11 | VDETCOM_T | -3 | -615 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | ON | OFF | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF
1 | Q1_VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk | q1_VGGCL | J1 | 1 | VGGCL_T | J1 | 1 | P1 | A2,B2 | VGGCL_T | JH1 | 4 | 3 |  | VGGCL_T | J1 | 3 | a | 1 | VGGCL_T | T-Q1&2 | a | cond | VGGCL_T | a | 1 | J1 | 3 | VGGCL_T | slot 14 | C17 |  |  |  | 14 | 2T | 1 | VGGCL_T_H | -2.7 | [25 empty] | 14 | 2T | 2 | VGGCL_T_L | -4.9 | -1004 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | ON | OFF | ON | ON | ON | ON | ON | OFF | ON | OFF | ON | ON
2 | Q1_VDDCL | -1.2 | VDDUC | column line overload protection | clk | q1_VDDCL | J1 | 3 | VDDCL_T | J1 | 3 | P1 | A27,B27 | VDDCL_T | JH2 | 20 | 19 |  | VDDCL_T | J2 | 19 | LL | 1 | VDDCL_T | T-Q1&2 | LL | shield | VDDCL_T | LL | 1 | J2 | 19 | VDDCL_T | slot 13 | C17 |  |  |  | 13 | 1T | 1 | VDDCL_T_H | -4.5 | -922 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | ON | OFF | ON | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | 13 | 1T | 2 | VDDCL_T_L | -1.5 | -308 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | ON | OFF | OFF | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON
4 | Q1_PHISS | VpROW&COL | VnROW | starts slow (row) shift register | clk | q1_PSS | J1 | 7 | PSS_T | J1 | 7 | P1 | A28,B28 | PSS_T | JH2 | 22 | 21 |  | PSS_T | J2 | 21 | AA | 1 | PSS_T | T-Q1&2 | AA | shield | PSS_T | AA | 1 | J2 | 21 | PSS_T | slot 14 | C31 |  |  |  | 14 | 2T | 15 | PSS_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 14 | 2T | 16 | PSS_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
6 | Q1_PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk | q1_PS2 | J1 | 11 | PS2_T | J1 | 11 | P1 | A29,B29 | PS2_T | JH2 | 24 | 23 |  | PS2_T | J2 | 23 | I | 1 | PS2_T | T-Q1&2 | I | shield | PS2_T | I | 1 | J2 | 23 | PS2_T | slot 13 | C21 |  |  |  | 13 | 1T | 5 | PS2_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 13 | 1T | 6 | PS2_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
98 | Q2_VROWON | -6.0 to -5.6 |  | unit cell row enable on position |  | q2_VROWON | J4 | 9 | VROWON_T | J4 | 9 | P1 | A3,B3 | VROWON_T | JH1 | 8 | 7 |  | VROWON_T | J1 | 7 | t | 1 | VROWON_T | T-Q1&2 | t | shield | VROWON_T | t | 1 | J1 | 7 | VROWON_T | slot 14 | C19
8 | Q1_PHIDES | VpROW&COL | VnROW | row deselect | clk | q1_PDES | J1 | 15 | PDES_T | J1 | 15 | P1 | A30,B30 | PDES_T | JH2 | 26 | 25 |  | PDES_T | J2 | 25 | H | 1 | PDES_T | T-Q1&2 | H | shield | PDES_T | H | 1 | J2 | 25 | PDES_T | slot 13 | C25 |  |  |  | 13 | 1T | 9 | PDES_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 13 | 1T | 10 | PDES_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
10 | Q1_PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk | q1_PF1 | J1 | 19 | PF1_T | J1 | 19 | P1 | A31,B31 | PF1_T | JH2 | 28 | 27 |  | PF1_T | J2 | 27 | K | 1 | PF1_T | T-Q1&2 | K | shield | PF1_T | K | 1 | J2 | 27 | PF1_T | slot 13 | C29 |  |  |  | 13 | 1T | 13 | PF1_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 13 | 1T | 14 | PF1_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
12 | Q1_VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias | q1_VNROW | J1 | 23 | VNROW_T | J1 | 23 | P1 | A32,B32 | VNROW_T | JH2 | 30 | 29 |  | VNROW_T | J2 | 29 | k | 1 | VNROW_T | T-Q1&2 | k | shield | VNROW_T | k | 1 | J2 | 29 | VNROW_T | slot 14 | C30 |  |  |  | 14 | 2T | 14 | VNROW_T | -5.6 | -1147 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | ON | ON | OFF | OFF | OFF | ON | ON | ON | ON | OFF | ON | OFF | 14 | 2T | 4 | VROWON_T_L | -5.6 | -1147 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | ON | ON | OFF | OFF | OFF | ON | ON | ON | ON | OFF | ON | OFF
96 | Q2_VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position |  | q2_VROWOFF | J4 | 5 | VROWOFF_T | J4 | 5 | P1 | A4,B4 | VROWOFF_T | JH1 | 5 | 5 |  | VROWOFF_T | J1 | 5 | GG | 1 | VROWOFF_T | T-Q1&2 | GG | shield | VROWOFF_T | GG | 1 | J1 | 5 | VROWOFF_T | AGND3
99 | Q2_VRSTR | -3.5 | -5.6 | unit cell reset on position = row reset | clk | q2_VRSTR | J4 | 11 | VRSTR_T | J4 | 11 | P1 | A5,B5 | VRSTR_T | JH1 | 10 | 9 |  | VRSTR_T | J1 | 9 | JJ | 1 | VRSTR_T | T-Q1&2 | JJ | shield | VRSTR_T | JJ | 1 | J1 | 9 | VRSTR_T | slot 14 | C21 |  |  |  | 14 | 2T | 5 | VRSTR_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 14 | 2T | 6 | VRSTR_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
16 | VSUB | 0 gnd |  | readout substrate | bias | AGND | J1 | 31 | AGND | J1 | 31 | P1 | C1 | AGND_T | JH1 | 1 | 2 | Cu | AGND_T | J1 | 2 | Y | 1 | AGND_T | T-Q1&2 | Y | shield | AGND_T | Y | 1 | J1 | 2 | AGND_T | AGND2
114 | Q1_OUT5 |  |  | p channel output 5 | output | q1_OUT5 | J4 | 41 | Q1OUT5+ | J4 | 41 | P1 | C10 | Q1OUT5+ | JH1 | 19 | 20 |  | Q1OUT5+ | J1 | 20 | S | 1 | Q1OUT5+ | T-Q1&2 | S | cond | Q1OUT5+ | S | 1 | J1 | 20 | Q1OUT5+ | slot 11 | C11
113 | Q1_OUT6 |  |  | p channel output 6 | output | q1_OUT6 | J4 | 39 | Q1OUT6+ | J4 | 39 | P1 | C11 | Q1OUT6+ | JH1 | 21 | 22 |  | Q1OUT6+ | J1 | 22 | n | 1 | Q1OUT6+ | T-Q1&2 | n | cond | Q1OUT6+ | n | 1 | J1 | 22 | Q1OUT6+ | slot 11 | C13
112 | Q1_OUT7 |  |  | p channel output 7 | output | q1_OUT7 | J4 | 37 | Q1OUT7+ | J4 | 37 | P1 | C12 | Q1OUT7+ | JH1 | 23 | 24 |  | Q1OUT7+ | J1 | 24 | BB | 1 | Q1OUT7+ | T-Q1&2 | BB | cond | Q1OUT7+ | BB | 1 | J1 | 24 | Q1OUT7+ | slot 11 | C15
111 | Q1_OUT8 |  |  | p channel output 8 | output | q1_OUT8 | J4 | 35 | Q1OUT8+ | J4 | 35 | P1 | C13 | Q1OUT8+ | JH1 | 25 | 26 |  | Q1OUT8+ | J1 | 26 | P | 1 | Q1OUT8+ | T-Q1&2 | P | cond | Q1OUT8+ | P | 1 | J1 | 26 | Q1OUT8+ | slot 11 | C17
108 | Q2_VDDOUT | -1.0 to -1.3 |  | output drain | bias | q2_VDDOUT | J4 | 29 | VDDOUT_T | J4 | 29 | P1 | C14,?C18 | VDDOUT_T | JH1&2 | 27_1 | 28_2 | _Cu | VDDOUT_T | J1&2 | 28_2 | N_B | 1 | VDDOUT_T | T-Q1&2 | N_B | c_s | VDDOUT_T | N_B | 1 | J1&2 | 28_2 | VDDOUT_T | slot 14 | C26 |  |  |  | 14 | 2T | 10 | VDDOUT_T | -1.3 | -267 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | ON | OFF | OFF | ON | OFF | OFF | OFF | OFF | ON | OFF | ON | OFF
124 | Q1_TEND |  |  | diagnostic |  | q1_TEND | J4 | 60 | Q1TEND | J4 | 60 | P1 | C15 | Q1TEND | JH1 | 30 | 29 |  | Q1TEND | J1 | 29 | PP | 1 | Q1TEND | T-Q1&2 | PP | PP | Q1TEND | PP | 1 | J1 | 29 | Q1TEND
94 | Q2_TEND |  |  | diagnostic |  | q2_TEND | J4 | 1 | Q2TEND | J4 | 1 | P1 | C17 | Q2TEND | JH1 | 33 | 34 |  | Q2TEND | [8 empty] | Q2TEND |  |  |  |  | Q2TEND
107 | Q2_OUT8 |  |  | p channel output 8 | output | q2_OUT8 | J4 | 27 | Q2OUT8+ | J4 | 27 | P1 | C19 | Q2OUT8+ | JH2 | 3 | 4 |  | Q2OUT8+ | J2 | 4 | b | 1 | Q2OUT8+ | T-Q1&2 | b | cond | Q2OUT8+ | b | 1 | J2 | 4 | Q2OUT8+ | slot 9 | C17
15 | VSS | 0 |  | unit cell source | bias | q1_VSSCM | J1 | 29 | VSSCM_T | J1 | 29 | P1 | C2 | VSSCM_T | JH1 | 3 | 4 |  | VSSCM_T | J2 | 4 | u | 1 | VSSCM_T | T-Q1&2 | u | shield | VSSCM_T | u | 1 | J2 | 4 | VSSCM_T | AGND7
106 | Q2_OUT7 |  |  | p channel output 7 | output | q2_OUT7 | J4 | 25 | Q2OUT7+ | J4 | 25 | P1 | C20 | Q2OUT7+ | JH2 | 5 | 6 |  | Q2OUT7+ | J2 | 6 | v | 1 | Q2OUT7+ | T-Q1&2 | v | cond | Q2OUT7+ | v | 1 | J2 | 6 | Q2OUT7+ | slot 9 | C15
105 | Q2_OUT6 |  |  | p channel output 6 | output | q2_OUT6 | J4 | 23 | Q2OUT6+ | J4 | 23 | P1 | C21 | Q2OUT6+ | JH2 | 7 | 8 |  | Q2OUT6+ | J2 | 8 | d | 1 | Q2OUT6+ | T-Q1&2 | d | cond | Q2OUT6+ | d | 1 | J2 | 8 | Q2OUT6+ | slot 9 | C13
104 | Q2_OUT5 |  |  | p channel output 5 | output | q2_OUT5 | J4 | 21 | Q2OUT5+ | J4 | 21 | P1 | C22 | Q2OUT5+ | JH2 | 9 | 10 |  | Q2OUT5+ | J2 | 10 | HH | 1 | Q2OUT5+ | T-Q1&2 | HH | cond | Q2OUT5+ | HH | 1 | J2 | 10 | Q2OUT5+ | slot 9 | C11
103 | Q2_OUT4 |  |  | p channel output 4 | output | q2_OUT4 | J4 | 19 | Q2OUT4+ | J4 | 19 | P1 | C23 | Q2OUT4+ | JH2 | 11 | 12 |  | Q2OUT4+ | J2 | 12 | x | 1 | Q2OUT4+ | T-Q1&2 | x | cond | Q2OUT4+ | x | 1 | J2 | 12 | Q2OUT4+ | slot 9 | C9
102 | Q2_OUT3 |  |  | p channel output 3 | output | q2_OUT3 | J4 | 17 | Q2OUT3+ | J4 | 17 | P1 | C24 | Q2OUT3+ | JH2 | 13 | 14 |  | Q2OUT3+ | J2 | 14 | C | 1 | Q2OUT3+ | T-Q1&2 | C | cond | Q2OUT3+ | C | 1 | J2 | 14 | Q2OUT3+ | slot 9 | C7
101 | Q2_OUT2 |  |  | p channel output 2 | output | q2_OUT2 | J4 | 15 | Q2OUT2+ | J4 | 15 | P1 | C25 | Q2OUT2+ | JH2 | 15 | 16 |  | Q2OUT2+ | J2 | 16 | f | 1 | Q2OUT2+ | T-Q1&2 | f | cond | Q2OUT2+ | f | 1 | J2 | 16 | Q2OUT2+ | slot 9 | C5
100 | Q2_OUT1 |  |  | p channel output 1 | output | q2_OUT1 | J4 | 13 | Q2OUT1+ | J4 | 13 | P1 | C26 | Q2OUT1+ | JH2 | 17 | 18 |  | Q2OUT1+ | J2 | 18 | E | 1 | Q2OUT1+ | T-Q1&2 | E | cond | Q2OUT1+ | E | 1 | J2 | 18 | Q2OUT1+ | slot 9 | C3
3 | Q1_VDDUC | -4.0 to -3.5 |  | unit cell drain | bias | q1_VDDUC | J1 | 5 | VDDUC_T | J1 | 5 | P1 | C27 | VDDUC_T | JH2 | 19 | 20 |  | VDDUC_T | J2 | 20 | KK | 1 | VDDUC_T | T-Q1&2 | KK | cond | VDDUC_T | KK | 1 | J2 | 20 | VDDUC_T | slot 14 | C28 |  |  |  | 14 | 2T | 12 | VDDUC_T | -3.7 | -758 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | ON | OFF | ON | OFF | ON | ON | ON | ON | OFF | ON | OFF | ON
5 | Q1_PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk | q1_PS1 | J1 | 9 | PS1_T | J1 | 9 | P1 | C28 | PS1_T | JH2 | 22 | 21 |  | PS1_T | J2 | 22 | z | 1 | PS1_T | T-Q1&2 | z | cond | PS1_T | z | 1 | J2 | 22 | PS1_T | slot 13 | C19 |  |  |  | 13 | 1T | 3 | PS1_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 13 | 1T | 4 | PS1_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
7 | Q1_PHIROE | VpROW&COL | VnROW | odd/even clock | clk | q1_PROE | J1 | 13 | PROE_T | J1 | 13 | P1 | C29 | PROE_T | JH2 | 23 | 24 |  | PROE_T | J2 | 24 | h | 1 | PROE_T | T-Q1&2 | h | cond | PROE_T | h | 1 | J2 | 24 | PROE_T | slot 13 | C23 |  |  |  | 13 | 1T | 7 | PROE_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 13 | 1T | 8 | PROE_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
95 | Q2_VRSTG | -3.5 | -5.6 | global reset | clk | q2_VRSTG | J4 | 3 | VRSTG_T | J4 | 3 | P1 | C3 | VRSTG_T | JH1 | 5 | 6 |  | VRSTG_T | J1 | 6 | FF | 1 | VRSTG_T | T-Q1&2 | FF | cond | VRSTG_T | FF | 1 | J1 | 6 | VRSTG_T | slot 14 | C23 |  |  |  | 14 | 2T | 7 | VRSTG_T_H | -3.5 | -717 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | OFF | ON | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | 14 | 2T | 8 | VRSTG_T_L | -5.6 | -1147 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | ON | ON | OFF | OFF | OFF | ON | ON | ON | ON | OFF | ON | OFF
9 | Q1_PHIFS | VpROW&COL | VnCOL | starts fast (col) shift register | clk | q1_PFS | J1 | 17 | PFS_T | J1 | 17 | P1 | C30 | PFS_T | JH2 | 25 | 26 |  | PFS_T | J2 | 26 | G | 1 | PFS_T | T-Q1&2 | G | cond | PFS_T | G | 1 | J2 | 26 | PFS_T | slot 13 | C27 |  |  |  | 13 | 1T | 11 | PFS_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 13 | 1T | 12 | PFS_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
11 | Q1_PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk | q1_PF2 | J1 | 21 | PF2_T | J1 | 21 | P1 | C31 | PF2_T | JH2 | 27 | 28 |  | PF2_T | J2 | 28 | J | 1 | PF2_T | T-Q1&2 | J | cond | PF2_T | J | 1 | J2 | 28 | PF2_T | slot 13 | C31 |  |  |  | 13 | 1T | 15 | PF2_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 13 | 1T | 16 | PF2_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
13 | Q1_VnCOL | -6.0 to -3.5 |  | negative supply for fast shift reg | bias | q1_VNCOL | J1 | 25 | VNCOL_T | J1 | 25 | P1 | C32 | VNCOL_T | JH2 | 29 | 30 |  | VNCOL_T | J2 | 30 | j | 1 | VNCOL_T | T-Q1&2 | j | cond | VNCOL_T | j | 1 | J2 | 30 | VNCOL_T | slot 14 | C29 |  |  |  | 14 | 2T | 13 | VNCOL_T | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
97 | Q2_IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource | q2_IREF | J4 | 7 | VIREF_T | J4 | 7 | P1 | C4 | VIREF_T | JH1 | 7 | 8 |  | VIREF_T | J1 | 8 | s | 1 | VIREF_T | T-Q1&2 | s | cond | VIREF_T | s | 1 | J1 | 8 | VIREF_T | slot 14 | C25 |  |  |  | 14 | 2T | 9 | VIREF_T | -1.8 | -369 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | ON | OFF | OFF | ON | OFF | ON | ON | ON | OFF | OFF | OFF | OFF
118 | Q1_OUT1 |  |  | p channel output 1 | output | q1_OUT1 | J4 | 49 | Q1OUT1+ | J4 | 49 | P1 | C6 | Q1OUT1+ | JH1 | 11 | 12 |  | Q1OUT1+ | J1 | 12 | DD | 1 | Q1OUT1+ | T-Q1&2 | DD | cond | Q1OUT1+ | DD | 1 | J1 | 12 | Q1OUT1+ | slot 11 | C3
117 | Q1_OUT2 |  |  | p channel output 2 | output | q1_OUT2 | J4 | 47 | Q1OUT2+ | J4 | 47 | P1 | C7 | Q1OUT2+ | JH1 | 13 | 14 |  | Q1OUT2+ | J1 | 14 | W | 1 | Q1OUT2+ | T-Q1&2 | W | cond | Q1OUT2+ | W | 1 | J1 | 14 | Q1OUT2+ | slot 11 | C5
116 | Q1_OUT3 |  |  | p channel output 3 | output | q1_OUT3 | J4 | 45 | Q1OUT3+ | J4 | 45 | P1 | C8 | Q1OUT3+ | JH1 | 15 | 16 |  | Q1OUT3+ | J1 | 16 | q | 1 | Q1OUT3+ | T-Q1&2 | q | cond | Q1OUT3+ | q | 1 | J1 | 16 | Q1OUT3+ | slot 11 | C7
115 | Q1_OUT4 |  |  | p channel output 4 | output | q1_OUT4 | J4 | 43 | Q1OUT4+ | J4 | 43 | P1 | C9 | Q1OUT4+ | JH1 | 17 | 18 |  | Q1OUT4+ | J1 | 18 | U | 1 | Q1OUT4+ | T-Q1&2 | U | cond | Q1OUT4+ | U | 1 | J1 | 18 | Q1OUT4+ | slot 11 | C9
48 | Q4_VDDOUT | -1.0 to -1.3 |  | output drain | bias | q3_VDDOUT | J2 | 33 | VDDOUT_B | J2 | 33 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | B-Q3&4 | * | * | * | * | * | * | * | *
57 | Q4_VRSTR | -3.5 | -5.6 | row reset | clk | q4_VRSTR | J2 | 51 | VRSTR_B | J2 | 51 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
58 | Q4_VROWON | -6.0 to -5.6 |  | unit cell row enable on position |  | q4_VROWON | J2 | 53 | VROWON_B | J2 | 53 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
59 | Q4_IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource | q4_IREF | J2 | 55 | VIREF_B | J2 | 55 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
60 | Q4_VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position |  | q4_VROWOFF | J2 | 57 | VROWOFF_B | J2 | 57 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
61 | Q4_VRSTG | -3.5 | -5.6 | global reset | clk | q4_VRSTG | J2 | 59 | VRSTG_B | J2 | 59 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
63 | Q4_VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk | q4_VGGCL | J3 | 1 | VGGCL_B | J3 | 1 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
64 | Q4_VDDCL | -1.2 | VDDUC | column line overload protection | clk | q4_VDDCL | J3 | 3 | VDDCL_B | J3 | 3 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
65 | Q4_VDDUC | -4.0 to -3.5 |  | unit cell drain | bias | Q4_VDDUC | J3 | 5 | VDDUC_B | J3 | 5 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
66 | Q4_PHISS | VpROW&COL | VnROW | starts slow (row) shift register | clk | q4_PSS | J3 | 7 | PSS_B | J3 | 7 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
67 | Q4_PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk | q4_PS1 | J3 | 9 | PS1_B | J3 | 9 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
68 | Q4_PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk | q4_PS2 | J3 | 11 | PS2_B | J3 | 11 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
69 | Q4_PHIROE | VpROW&COL | VnROW | phase 3 clock | clk | q4_PROE | J3 | 13 | PROE_B | J3 | 13 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
70 | Q4_PHIDES | VpROW&COL | VnROW | resets detector bias | clk | q4_PDES | J3 | 15 | PDES_B | J3 | 15 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
71 | Q4_PHIFS | VpROW&COL | VnCOL | starts fast (col) shift register | clk | q4_PFS | J3 | 17 | PFS_B | J3 | 17 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
72 | Q4_PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk | q4_PF1 | J3 | 19 | PF1_B | J3 | 19 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
73 | Q4_PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk | q4_PF2 | J3 | 21 | PF2_B | J3 | 21 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
74 | Q4_VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias | q4_VNROW | J3 | 23 | VNROW_B | J3 | 23 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
75 | Q4_VnCOL | -6.0 to -3.5 |  | negative supply for fast shift reg | bias | q4_VNCOL | J3 | 25 | VNCOL_B | J3 | 25 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
76 | Q4_VP | 0.0 to -0.5 |  | positive supply for shift reg | bias | q4_VP | J3 | 27 | VP_B | J3 | 27 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
77 | VSS | 0 |  | unit cell source | bias | q4_VSSCM | J3 | 29 | VSSCM_B | J3 | 29 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
78 | VSUB | 0 gnd |  | readout substrate | bias | AGND | J3 | 31 | AGND | J3 | 31 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
79 | VSS | 0 |  | unit cell source | bias | q2_VSSCM | J3 | 33 | VSSCM_T | J3 | 33 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
80 | Q2_Vp | 0.0 to -0.5 |  | positive supply for shift reg | bias | q2_VP | J3 | 35 | VP_T | J3 | 35 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
81 | Q2_VnCOL | -6.0 to -3.5 |  | negative supply for fast shift reg | bias | q2_VNCOL | J3 | 37 | VNCOL_T | J3 | 37 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
82 | Q2_VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias | q2_VNROW | J3 | 39 | VNROW_T | J3 | 39 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
83 | Q2_PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk | q2_PF2 | J3 | 41 | PF2_T | J3 | 41 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
84 | Q2_PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk | q2_PF1 | J3 | 43 | PF1_T | J3 | 43 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
85 | Q2_PHIFS | VpROW&COL | VnCOL | starts fast (col) shift register | clk | q2_PFS | J3 | 45 | PFS_T | J3 | 45 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
86 | Q2_PHIDES | VpROW&COL | VnROW | row deselect | clk | q2_PDES | J3 | 47 | PDES_T | J3 | 47 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
87 | Q2_PHIROE | VpROW&COL | VnROW |  | clk | q2_PROE | J3 | 49 | PROE_T | J3 | 49 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
88 | Q2_PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk | q2_PHIS2 | J3 | 51 | PS2_T | J3 | 51 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
89 | Q2_PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk | q2_PS1 | J3 | 53 | PS1_T | J3 | 53 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
90 | Q2_PHISS | VpROW&COL | VnROW | starts slow (row) shift register | clk | q2_PSS | J3 | 55 | PSS_T | J3 | 55 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
91 | Q2_VDDUC | -4.0 to -3.5 |  | unit cell drain | bias | q2_VDDUC | J3 | 57 | VDDUC_T | J3 | 57 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
92 | Q2_VDDCL | -1.2 | VDDUC | column line overload protection | clk | q2_VDDCL | J3 | 59 | VDDCL_T | J3 | 59 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
93 | Q2_VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk | q2_VGGCL | J3 | 60 | VDDCL_T | J3 | 60 | P2 | * | * | * | * | * |  | * | * | * | * | * | * | * | * | * | * | * | * | * | * | *
18 | Q3_Vp | 0.0 to -0.5 |  | positive supply for shift reg | bias | q3_VP | J1 | 35 | VP_B | J1 | 35 | P2 | A1,B1 | VP_B | JH1 | 1 | 2 | Cu | VP_B | J1 | 2 | Y | 2 | VP_B | B-Q3&4 | Y | Y | VP_B | Y | 2 | J1 | 2 | VP_B | AGND4
47 | Q3_VDETCOM | -2.8 to-3.4 |  | detector common supply | bias | q3_VDETCOM | J2 | 31 | VDETCOM | J2 | 31 | P2 | A18,B18 | VDETCOM_B | JH2 | 2 | 1 | Cu | VDETCOM_B | J2 | 1 | A | 2 | VDETCOM_B | B-Q3&4 | A | cond | VDETCOM_B | A | 2 | J2 | 1 | VDETCOM_B | slot 6 | C27 |  |  |  | 6 | 2B | 11 | VDETCOM_B | -3 | -615 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | ON | OFF | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF
31 | Q3_VGGCL | VDETCOM | -5 | connects column lines to VDDCL | clk | q3_VGGCL | J1 | 60 | VGGCL_B | J1 | 60 | P2 | A2,B2 | VGGCL_B | JH1 | 4 | 3 |  | VGGCL_B | J1 | 3 | a | 2 | VGGCL_B | B-Q3&4 | a | cond | VGGCL_B | a | 2 | J1 | 3 | VGGCL_B | slot 6 | C17 |  |  |  | 6 | 2B | 1 | VGGCL_B_H | -2.7 | -553 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | ON | OFF | ON | OFF | OFF | OFF | ON | OFF | ON | OFF | OFF | OFF | 6 | 2B | 2 | VGGCL_B_L | -4.9 | -1004 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | ON | OFF | ON | ON | ON | ON | ON | OFF | ON | OFF | ON | ON
30 | Q3_VDDCL | -1.2 | VDDUC | column line overload protection | clk | q3_VDDCL | J1 | 59 | VDDCL_B | J1 | 59 | P2 | A27,B27 | VDDCL_B | JH2 | 20 | 19 |  | VDDCL_B | J2 | 19 | LL | 2 | VDDCL_B | B-Q3&4 | LL | shield | VDDCL_B | LL | 2 | J2 | 19 | VDDCL_B | slot 5 | C17 |  |  |  | 5 | 1B | 1 | VDDCL_B_H | -3.4 | -697 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | ON | OFF | ON | OFF | ON | OFF | ON | ON | ON | OFF | OFF | OFF | 5 | 1B | 2 | VDDCL_B_L | -3.4 | -697 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | ON | OFF | ON | OFF | ON | OFF | ON | ON | ON | OFF | OFF | OFF
28 | Q3_PHISS | VpROW&COL | VnROW | starts slow (row) shift register | clk | q3_PSS | J1 | 55 | PSS_B | J1 | 55 | P2 | A28,B28 | PSS_B | JH2 | 22 | 21 |  | PSS_B | J2 | 21 | AA | 2 | PSS_B | B-Q3&4 | AA | shield | PSS_B | AA | 2 | J2 | 21 | PSS_B | slot 6 | C31 |  |  |  | 6 | 2B | 15 | PSS_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 6 | 2B | 16 | PSS_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
26 | Q3_PHIS2 | VpROW&COL | VnROW | phase 2 clock | clk | q3_PS2 | J1 | 51 | PS2_B | J1 | 51 | P2 | A29,B29 | PS2_B | JH2 | 24 | 23 |  | PS2_B | J2 | 23 | i | 2 | PS2_B | B-Q3&4 | i | shield | PS2_B | i | 2 | J2 | 23 | PS2_B | slot 5 | C21 |  |  |  | 5 | 1B | 5 | PS2_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 5 | 1B | 6 | PS2_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
36 | Q3_VROWON | -6.0 to -5.6 |  | unit cell row enable on position |  | q3_VROWON | J2 | 9 | VROWON_B | J2 | 9 | P2 | A3,B3 | VROWON_B | JH1 | 6 | 5 |  | VROWON_B | J1 | 5 | GG | 2 | VROWON_B | B-Q3&4 | GG | shield | VROWON_B | GG | 2 | J1 | 5 | VROWON_B | slot 6 | C19 |  |  |  | 6 | 2B | 3 | VROWON_B_H | -5.6 | -1147 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | ON | ON | OFF | OFF | OFF | ON | ON | ON | ON | OFF | ON | OFF | 6 | 2B | 4 | VROWON_B_L | -5.6 | -1147 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | ON | ON | OFF | OFF | OFF | ON | ON | ON | ON | OFF | ON | OFF
24 | Q3_PHIDES | VpROW&COL | VnROW | row deselect | clk | q3_PDES | J1 | 47 | PDES_B | J1 | 47 | P2 | A30,B30 | PDES_B | JH2 | 26 | 25 |  | PDES_B | J2 | 25 | H | 2 | PDES_B | B-Q3&4 | H | shield | PDES_B | H | 2 | J2 | 25 | PDES_B | slot 5 | C25 |  |  |  | 5 | 1B | 9 | PDES_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 5 | 1B | 10 | PDES_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
22 | Q3_PHIF1 | VpROW&COL | VnCOL | phase1 clock | clk | q3_PF1 | J1 | 43 | PF1_B | J1 | 43 | P2 | A31,B31 | PF1_B | JH2 | 28 | 27 |  | PF1_B | J2 | 27 | K | 2 | PF1_B | B-Q3&4 | K | shield | PF1_B | K | 2 | J2 | 27 | PF1_B | slot 5 | C29 |  |  |  | 5 | 1B | 13 | PF1_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 5 | 1B | 14 | PF1_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
20 | Q3_VnROW | -6.0 to -5.6 |  | negative supply for row shift reg | bias | q3_VNROW | J1 | 39 | VNROW_B | J1 | 39 | P2 | A32,B32 | VNROW_B | JH2 | 30 | 29 |  | VNROW_B | J2 | 29 | k | 2 | VNROW_B | B-Q3&4 | k | shield | VNROW_B | k | 2 | J2 | 29 | VNROW_B | slot 6 | C30 |  |  |  | 6 | 2B | 14 | VNROW_B | -5.6 | -1147 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | ON | ON | OFF | OFF | OFF | ON | ON | ON | ON | OFF | ON | OFF
34 | Q3_VROWOFF | 0.0 to -0.7 |  | unit cell row enable off position |  | q3_VROWOFF | J2 | 5 | VROWOFF_B | J2 | 5 | P2 | A4,B4 | VROWOFF_B | JH1 | 8 | 7 |  | VROWOFF_B | J1 | 7 | t | 2 | VROWOFF_B | B-Q3&4 | t | shield | VROWOFF_B | t | 2 | J1 | 7 | VROWOFF_B | AGND5
37 | Q3_VRSTR | -3.5 | -5.6 | row reset | clk | q3_VRSTR | J2 | 11 | VRSTR_B | J2 | 11 | P2 | A5,B5 | VRSTR_B | JH1 | 10 | 9 |  | VRSTR_B | J1 | 9 | MM | 2 | VRSTR_B | B-Q3&4 | MM | cond | VRSTR_B | MM | 2 | J1 | 9 | VRSTR_B | slot 6 | C21 |  |  |  | 6 | 2B | 5 | VRSTR_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 6 | 2B | 6 | VRSTR_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
42 | Q3_OUT5 |  |  | p channel output 5 | output | q3_OUT5 | J2 | 21 | Q3OUT5+ | J2 | 21 | P2 | C10 | Q3OUT5+ | JH1 | 19 | 20 |  | Q3OUT5+ | J1 | 20 | S | 2 | Q3OUT5+ | B-Q3&4 | S | cond | Q3OUT5+ | S | 2 | J1 | 20 | Q3OUT5+ | slot 3 | C11
43 | Q3_OUT6 |  |  | p channel output 6 | output | q3_OUT6 | J2 | 23 | Q3OUT6+ | J2 | 23 | P2 | C11 | Q3OUT6+ | JH1 | 21 | 22 |  | Q3OUT6+ | J1 | 22 | n | 2 | Q3OUT6+ | B-Q3&4 | n | cond | Q3OUT6+ | n | 2 | J1 | 22 | Q3OUT6+ | slot 3 | C13
44 | Q3_OUT7 |  |  | p channel output 7 | output | q3_OUT7 | J2 | 25 | Q3OUT7+ | J2 | 25 | P2 | C12 | Q3OUT7+ | JH1 | 23 | 24 |  | Q3OUT7+ | J1 | 24 | BB | 2 | Q3OUT7+ | B-Q3&4 | BB | cond | Q3OUT7+ | BB | 2 | J1 | 24 | Q3OUT7+ | slot 3 | C15
45 | Q3_OUT8 |  |  | p channel output 8 | output | q3_OUT8 | J2 | 27 | Q3OUT8+ | J2 | 27 | P2 | C13 | Q3OUT8+ | JH1 | 25 | 26 |  | Q3OUT8+ | J1 | 26 | P | 2 | Q3OUT8+ | B-Q3&4 | P | cond | Q3OUT8+ | P | 2 | J1 | 26 | Q3OUT8+ | slot 3 | C17
46 | Q3_VDDOUT | -1.0 to -1.3 |  | output drain | bias | q3_VDDOUT | J2 | 29 | VDDOUT_B | J2 | 29 | P2 | C14 | VDDOUT_B | JH1&2 | 27_1 | 28_2 | _Cu | VDDOUT_B | J1&2 | 28_2 | N_B | 2 | VDDOUT_B | B-Q3&4 | N_B | c_s | VDDOUT_B | N_B | 2 | J1&2 | 28_2 | VDDOUT_B | slot 6 | C26 |  |  |  | 6 | 2B | 10 | VDDOUT_B | -1.3 | -267 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | ON | OFF | OFF | ON | OFF | OFF | OFF | OFF | ON | OFF | ON | OFF
32 | Q3_TEND |  |  | diagnostic |  | q3_TEND | J2 | 1 | Q3_TEND | J2 | 1 | P2 | C15 | Q3_TEND | JH1 | 29 | 30 |  | Q3_TEND | J1 | 30 | nc | 2 |  |  |  |  | Q3_TEND | nc | 2 | J1 | 30 | Q3_TEND
62 | Q4_TEND |  |  | diagnostic |  | q4_TEND | J2 | 60 | Q4_TEND | J2 | 60 | P2 | C17 | Q4_TEND | JH2 | 33 | 34 |  | Q4_TEND | J2 | 34 | NC | [5 empty] | Q4_TEND | NC |  | J2 | 34 | Q4_TEND
49 | Q4_OUT8 |  |  | p channel output 8 | output | q4_OUT8 | J2 | 35 | Q4OUT8+ | J2 | 35 | P2 | C19 | Q4OUT8+ | JH2 | 3 | 4 |  | Q4OUT8+ | J2 | 4 | b | 2 | Q4OUT8+ | B-Q3&4 | b | cond | Q4OUT8+ | b | 2 | J2 | 4 | Q4OUT8+ | slot 1 | C17
17 | VSS | 0 |  | unit cell source | bias | q3_VSSCM | J1 | 33 | VSSCM_B | J1 | 33 | P2 | C2 | VSSCM_B | JH1 | 3 | 4 |  | VSSCM_B | J1 | 4 | u | 2 | VSSCM_B | B-Q3&4 | u | shield | VSSCM_B | u | 2 | J1 | 4 | VSSCM_B | AGND6
50 | Q4_OUT7 |  |  | p channel output 7 | output | q4_OUT7 | J2 | 37 | Q4OUT7+ | J2 | 37 | P2 | C20 | Q4OUT7+ | JH2 | 5 | 6 |  | Q4OUT7+ | J2 | 6 | v |  | Q4OUT7+ | B-Q3&4 | v | cond | Q4OUT7+ | v |  | J2 | 6 | Q4OUT7+ | slot 1 | C15
51 | Q4_OUT6 |  |  | p channel output 6 | output | q4_OUT6 | J2 | 39 | Q4OUT6+ | J2 | 39 | P2 | C21 | Q4OUT6+ | JH2 | 7 | 8 |  | Q4OUT6+ | J2 | 8 | d |  | Q4OUT6+ | B-Q3&4 | d | cond | Q4OUT6+ | d |  | J2 | 8 | Q4OUT6+ | slot 1 | C13
52 | Q4_OUT5 |  |  | p channel output 5 | output | q4_OUT5 | J2 | 41 | Q4OUT5+ | J2 | 41 | P2 | C22 | Q4OUT5+ | JH2 | 9 | 10 |  | Q4OUT5+ | J2 | 10 | HH |  | Q4OUT5+ | B-Q3&4 | HH | cond | Q4OUT5+ | HH |  | J2 | 10 | Q4OUT5+ | slot 1 | C11
53 | Q4_OUT4 |  |  | p channel output 4 | output | q4_OUT4 | J2 | 43 | Q4OUT4+ | J2 | 43 | P2 | C23 | Q4OUT4+ | JH2 | 11 | 12 |  | Q4OUT4+ | J2 | 12 | x |  | Q4OUT4+ | B-Q3&4 | x | cond | Q4OUT4+ | x |  | J2 | 12 | Q4OUT4+ | slot 1 | C9
54 | Q4_OUT3 |  |  | p channel output 3 | output | q4_OUT3 | J2 | 45 | Q4OUT3+ | J2 | 45 | P2 | C24 | Q4OUT3+ | JH2 | 13 | 14 |  | Q4OUT3+ | J2 | 13 | C |  | Q4OUT3+ | B-Q3&4 | C | cond | Q4OUT3+ | C |  | J2 | 13 | Q4OUT3+ | slot 1 | C7
55 | Q4_OUT2 |  |  | p channel output 2 | output | q4_OUT2 | J2 | 47 | Q4OUT2+ | J2 | 47 | P2 | C25 | Q4OUT2+ | JH2 | 15 | 16 |  | Q4OUT2+ | J2 | 16 | f |  | Q4OUT2+ | B-Q3&4 | f | cond | Q4OUT2+ | f |  | J2 | 16 | Q4OUT2+ | slot 1 | C5
56 | Q4_OUT1 |  |  | p channel output 1 | output | q4_OUT1 | J2 | 49 | Q4OUT1+ | J2 | 49 | P2 | C26 | Q4OUT1+ | JH2 | 17 | 18 |  | Q4OUT1+ | J2 | 18 | E |  | Q4OUT1+ | B-Q3&4 | E | cond | Q4OUT1+ | E |  | J2 | 18 | Q4OUT1+ | slot 1 | C3
29 | Q3_VDDUC | -4.0 to -3.5 |  | unit cell drain | bias | q3_VDDUC | J1 | 57 | VDDUC_B | J1 | 57 | P2 | C27 | VDDUC_B | JH2 | 19 | 20 |  | VDDUC_B | J2 | 20 | KK | 2 | VDDUC_B | B-Q3&4 | KK | cond | VDDUC_B | KK | 2 | J2 | 20 | VDDUC_B | slot 6 | C28
27 | Q3_PHIS1 | VpROW&COL | VnROW | phase 1 clock | clk | q3_PS1 | J1 | 53 | PS1_B | J1 | 53 | P2 | C28 | PS1_B | JH2 | 21 | 22 |  | PS1_B | J2 | 22 | z | 2 | PS1_B | B-Q3&4 | z | cond | PS1_B | z | 2 | J2 | 22 | PS1_B | slot 5 | C19 |  |  |  | 5 | 1B | 3 | PS1_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 5 | 1B | 4 | PS1_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
25 | Q3_PHIROE | VpROW&COL | VnROW | phase 3 clock | clk | q3_PROE | J1 | 49 | PROE_B | J1 | 49 | P2 | C29 | PROE_B | JH2 | 23 | 24 |  | PROE_B | J2 | 24 | h | 2 | PROE_B | B-Q3&4 | h | cond | PROE_B | h | 2 | J2 | 24 | PROE_B | slot 5 | C23 |  |  |  | 5 | 1B | 7 | PROE_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 5 | 1B | 8 | PROE_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
33 | Q3_VRSTG | -3.5 | -5.6 | global reset | clk | q3_VRSTG | J2 | 3 | VRSTG_B | J2 | 3 | P2 | C3 | VRSTG_B | JH1 | 5 | 6 |  | VRSTG_B | J1 | 6 | FF | 2 | VRSTG_B | B-Q3&4 | FF | cond | VRSTG_B | FF | 2 | J1 | 6 | VRSTG_B | slot 6 | C23 |  |  |  | 6 | 2B | 7 | VRSTG_B_H | -3.5 | -717 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | OFF | ON | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | 6 | 2B | 8 | VRSTG_B_L | -5.6 | -1147 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | ON | ON | OFF | OFF | OFF | ON | ON | ON | ON | OFF | ON | OFF
23 | Q3_PHIFS | VpROW&COL | VnCOL | starts fast (col) shift register | clk | q3_PFS | J1 | 45 | PFS_B | J1 | 45 | P2 | C30 | PFS_B | JH2 | 25 | 26 |  | PFS_B | J2 | 26 | G | 2 | PFS_B | B-Q3&4 | G | cond | PFS_B | G | 2 | J2 | 26 | PFS_B | slot 5 | C27 |  |  |  | 5 | 1B | 11 | PFS_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 5 | 1B | 12 | PFS_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
21 | Q3_PHIF2 | VpROW&COL | VnCOL | phase 2 clock | clk | q3_PF2 | J1 | 41 | PF2_B | J1 | 41 | P2 | C31 | PF2_B | JH2 | 27 | 28 |  | PF2_B | J2 | 28 | J | 2 | PF2_B | B-Q3&4 | J | cond | PF2_B | J | 2 | J2 | 28 | PF2_B | slot 5 | C31 |  |  |  | 5 | 1B | 15 | PF2_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | 5 | 1B | 16 | PF2_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
19 | Q3_VnCOL | -6.0 to -3.5 |  | negative supply for fast shift reg | bias | q3_VNCOL | J1 | 37 | VNCOL_B | J1 | 37 | P2 | C32 | VNCOL_B | JH2 | 29 | 30 |  | VNCOL_B | J2 | 30 | j | 2 | VNCOL_B | B-Q3&4 | j | cond | VNCOL_B | j | 2 | J2 | 30 | VNCOL_B | slot 6 | C29 |  |  |  | 6 | 2B | 13 | VNCOL_B | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
35 | Q3_IREF | -1.5 to -2.5 |  | w/125k ohm resistor for current source ea. quad | Isource | q3_IREF | J2 | 7 | VIREF_B | J2 | 7 | P2 | C4 | VIREF_B | JH1 | 7 | 8 |  | VIREF_B | J1 | 8 | s | 2 | VIREF_B | B-Q3&4 | s | cond | VIREF_B | s | 2 | J1 | 8 | VIREF_B | slot 6 | C25 |  |  |  | 6 | 2B | 9 | VIREF_B | -1.8 | -369 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | ON | OFF | OFF | ON | OFF | ON | ON | ON | OFF | OFF | OFF | OFF
38 | Q3_OUT1 |  |  | p channel output 1 | output | q3_OUT1 | J2 | 13 | Q3OUT1+ | J2 | 13 | P2 | C6 | Q3OUT1+ | JH1 | 11 | 12 |  | Q3OUT1+ | J1 | 12 | DD | 2 | Q3OUT1+ | B-Q3&4 | DD | cond | Q3OUT1+ | DD | 2 | J1 | 12 | Q3OUT1+ | slot 3 | C3
39 | Q3_OUT2 |  |  | p channel output 2 | output | q3_OUT2 | J2 | 15 | Q3OUT2+ | J2 | 15 | P2 | C7 | Q3OUT2+ | JH1 | 13 | 14 |  | Q3OUT2+ | J1 | 14 | W | 2 | Q3OUT2+ | B-Q3&4 | W | cond | Q3OUT2+ | W | 2 | J1 | 14 | Q3OUT2+ | slot 3 | C5
40 | Q3_OUT3 |  |  | p channel output 3 | output | q3_OUT3 | J2 | 17 | Q3OUT3+ | J2 | 17 | P2 | C8 | Q3OUT3+ | JH1 | 15 | 16 |  | Q3OUT3+ | J1 | 16 | q | 2 | Q3OUT3+ | B-Q3&4 | q | cond | Q3OUT3+ | q | 2 | J1 | 16 | Q3OUT3+ | slot 3 | C7
41 | Q3_OUT4 |  |  | p channel output 4 | output | q3_OUT4 | J2 | 19 | Q3OUT4+ | J2 | 19 | P2 | C9 | Q3OUT4+ | JH1 | 17 | 18 |  | Q3OUT4+ | J1 | 18 | U | 2 | Q3OUT4+ | B-Q3&4 | U | cond | Q3OUT4+ | U | 2 | J1 | 18 | Q3OUT4+ | slot 3 | C9

## BIASVS
SPEX BIAS VOLTAGES AS OF 10/3/97 USING SBRC USER GUIDE VOLTAGES
|  |  |  | 1 LSB = | 0.0048828 | DAC INPUT BIT CODE | [11 empty] | DIP SWITCH POSITION
| STD BIAS |  |  | VOLTS | # OF LSB'S | 12 | 11 | 10 | 9 | 8 | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | JHX | JHXX
| VDDUC | VDET |  | 9.9951 | 2047 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF
| -3.7 | -3.375 |  | 9 | 1843 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | OFF | OFF | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
|  |  | signal | 2 | 409 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF
slot# | Board # | DAC# | * program | 1 | 204 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | OFF | ON | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON
|  |  |  | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 1B | 1 | VDDCL_B_H | -3.4 | -697 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | ON | OFF | ON | OFF | ON | OFF | ON | ON | ON | OFF | OFF | OFF
5 | 1B | 2 | VDDCL_B_L | -3.4 | -697 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | ON | OFF | ON | OFF | ON | OFF | ON | ON | ON | OFF | OFF | OFF
5 | 1B | 3 | PS1_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 1B | 4 | PS1_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
5 | 1B | 5 | PS2_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 1B | 6 | PS2_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
5 | 1B | 7 | PROE_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 1B | 8 | PROE_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
5 | 1B | 9 | PDES_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 1B | 10 | PDES_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
5 | 1B | 11 | PFS_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 1B | 12 | PFS_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
5 | 1B | 13 | PF1_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 1B | 14 | PF1_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
5 | 1B | 15 | PF2_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 1B | 16 | PF2_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
6 | 2B | 1 | VGGCL_B_H | -2.7 | -553 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | ON | OFF | ON | OFF | OFF | OFF | ON | OFF | ON | OFF | OFF | OFF
6 | 2B | 2 | VGGCL_B_L | -4.9 | -1004 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | ON | OFF | ON | ON | ON | ON | ON | OFF | ON | OFF | ON | ON
6 | 2B | 3 | VROWON_B_H | -5.6 | -1147 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | ON | ON | OFF | OFF | OFF | ON | ON | ON | ON | OFF | ON | OFF
6 | 2B | 4 | VROWON_B_L | -5.6 | -1147 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | ON | ON | OFF | OFF | OFF | ON | ON | ON | ON | OFF | ON | OFF
6 | 2B | 5 | VRSTR_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
6 | 2B | 6 | VRSTR_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
6 | 2B | 7 | VRSTG_B_H | -3.5 | -717 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | OFF | ON | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
6 | 2B | 8 | VRSTG_B_L | -5.6 | -1147 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | ON | ON | OFF | OFF | OFF | ON | ON | ON | ON | OFF | ON | OFF
6 | 2B | 9 | VIREF_B | -1.8 | -369 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | ON | OFF | OFF | ON | OFF | ON | ON | ON | OFF | OFF | OFF | OFF
6 | 2B | 10 | VDDOUT_B | -1.3 | -267 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | ON | OFF | OFF | ON | OFF | OFF | OFF | OFF | ON | OFF | ON | OFF
6 | 2B | 11 | VDETCOM_B | -3 | -615 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | ON | OFF | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF
6 | 2B | 12 | VDDUC_B | -3.7 | -758 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | ON | OFF | ON | OFF | ON | ON | ON | ON | OFF | ON | OFF | ON
6 | 2B | 13 | VNCOL_B | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
6 | 2B | 14 | VNROW_B | -5.6 | -1147 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | ON | ON | OFF | OFF | OFF | ON | ON | ON | ON | OFF | ON | OFF
6 | 2B | 15 | PSS_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
6 | 2B | 16 | PSS_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
13 | 1T | 1 | VDDCL_T_H | -3.4 | -697 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | ON | OFF | ON | OFF | ON | OFF | ON | ON | ON | OFF | OFF | OFF
13 | 1T | 2 | VDDCL_T_L | -3.4 | -697 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | ON | OFF | ON | OFF | ON | OFF | ON | ON | ON | OFF | OFF | OFF
13 | 1T | 3 | PS1_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
13 | 1T | 4 | PS1_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
13 | 1T | 5 | PS2_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
13 | 1T | 6 | PS2_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
13 | 1T | 7 | PROE_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
13 | 1T | 8 | PROE_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
13 | 1T | 9 | PDES_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
13 | 1T | 10 | PDES_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
13 | 1T | 11 | PFS_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
13 | 1T | 12 | PFS_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
13 | 1T | 13 | PF1_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
13 | 1T | 14 | PF1_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
13 | 1T | 15 | PF2_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
13 | 1T | 16 | PF2_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
14 | 2T | 1 | VGGCL_T_H | -2.7 | -553 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | ON | OFF | ON | OFF | OFF | OFF | ON | OFF | ON | OFF | OFF | OFF
14 | 2T | 2 | VGGCL_T_L | -4.9 | -1004 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | ON | OFF | ON | ON | ON | ON | ON | OFF | ON | OFF | ON | ON
14 | 2T | 3 | VROWON_T_H | -5.6 | -1147 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | ON | ON | OFF | OFF | OFF | ON | ON | ON | ON | OFF | ON | OFF
14 | 2T | 4 | VROWON_T_L | -5.6 | -1147 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | ON | ON | OFF | OFF | OFF | ON | ON | ON | ON | OFF | ON | OFF
14 | 2T | 5 | VRSTR_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
14 | 2T | 6 | VRSTR_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
14 | 2T | 7 | VRSTG_T_H | -3.5 | -717 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | OFF | ON | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
14 | 2T | 8 | VRSTG_T_L | -5.6 | -1147 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | ON | ON | OFF | OFF | OFF | ON | ON | ON | ON | OFF | ON | OFF
14 | 2T | 9 | VIREF_T | -1.8 | -369 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | ON | OFF | OFF | ON | OFF | ON | ON | ON | OFF | OFF | OFF | OFF
14 | 2T | 10 | VDDOUT_T | -1.3 | -267 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | ON | OFF | OFF | ON | OFF | OFF | OFF | OFF | ON | OFF | ON | OFF
14 | 2T | 11 | VDETCOM_T | -3 | -615 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | ON | OFF | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF
14 | 2T | 12 | VDDUC_T | -3.7 | -758 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | ON | OFF | ON | OFF | ON | ON | ON | ON | OFF | ON | OFF | ON
14 | 2T | 13 | VNCOL_T | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
14 | 2T | 14 | VNROW_T | -5.6 | -1147 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | ON | ON | OFF | OFF | OFF | ON | ON | ON | ON | OFF | ON | OFF
14 | 2T | 15 | PSS_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
14 | 2T | 16 | PSS_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF

## BIAS2 (2)
SPEX BIAS VOLTAGES AS OF 3/20/98 USING some FOWLER VOLTAGES
|  |  |  | 1 LSB = | 0.0048828 | DAC INPUT BIT CODE | [11 empty] | DIP SWITCH POSITION
| STD BIAS |  |  | VOLTS | # OF LSB'S | 12 | 11 | 10 | 9 | 8 | 7 | 6 | 5 | 4 | 3 | 2 | 1 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | JHX | JHXX
| VDDUC | VDET |  | 9.9951 | 2047 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF | OFF
| -3.7 | -3.375 |  | 9 | 1843 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | OFF | OFF | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
|  |  | signal | 2 | 409 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF
slot# | Board # | DAC# | * program | 1 | 204 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | OFF | ON | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON
|  |  |  | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 1B | 0 | VDDCL_B_H | -4.5 | -922 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | ON | OFF | ON | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON
5 | 1B | 1 | VDDCL_B_L | -1.3 | -267 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | ON | OFF | OFF | ON | OFF | OFF | OFF | OFF | ON | OFF | ON | OFF
5 | 1B | 2 | PS1_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 1B | 3 | PS1_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
5 | 1B | 4 | PS2_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 1B | 5 | PS2_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
5 | 1B | 6 | PROE_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 1B | 7 | PROE_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
5 | 1B | 8 | PDES_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 1B | 9 | PDES_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
5 | 1B | 10 | PFS_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 1B | 11 | PFS_B_L | -4 | -820 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | ON | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON
5 | 1B | 12 | PF1_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 1B | 13 | PF1_B_L | -4 | -820 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | ON | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON
5 | 1B | 14 | PF2_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
5 | 1B | 15 | PF2_B_L | -4 | -820 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | ON | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON
6 | 1 | 0 | VGGCL_B_H | -3.75 | -768 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ON | OFF | ON | OFF | ON | ON | ON | ON | ON | ON | ON | ON
6 | 1 | 1 | VGGCL_B_L | -4.5 | -922 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | ON | OFF | ON | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON
6 | 1 | 2 | VROWON_B_H | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
6 | 1 | 3 | VROWON_B_L | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
6 | 1 | 4 | VRSTR_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
6 | 1 | 5 | VRSTR_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
6 | 1 | 6 | VRSTG_B_H | -2 | -410 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON
6 | 1 | 7 | VRSTG_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
6 | 1 | 8 | VIREF_B | -1.8 | -369 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | ON | OFF | OFF | ON | OFF | ON | ON | ON | OFF | OFF | OFF | OFF
6 | 1 | 9 | VDDOUT_B | -1.3 | -267 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | ON | OFF | OFF | ON | OFF | OFF | OFF | OFF | ON | OFF | ON | OFF
6 | 1 | 10 | VDETCOM_B | -3.75 | -768 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ON | OFF | ON | OFF | ON | ON | ON | ON | ON | ON | ON | ON
6 | 1 | 11 | VDDUC_B | -3.75 | -768 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ON | OFF | ON | OFF | ON | ON | ON | ON | ON | ON | ON | ON
6 | 1 | 12 | VNCOL_B | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
6 | 1 | 13 | VNROW_B | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
6 | 1 | 14 | PSS_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
6 | 1 | 15 | PSS_B_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
13 | 1T | 0 | VDDCL_T_H | -4.5 | -922 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | ON | OFF | ON | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON
13 | 1T | 1 | VDDCL_T_L | -1.3 | -267 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | ON | OFF | OFF | ON | OFF | OFF | OFF | OFF | ON | OFF | ON | OFF
13 | 1T | 2 | PS1_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
13 | 1T | 3 | PS1_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
13 | 1T | 4 | PS2_B_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
13 | 1T | 5 | PS2_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
13 | 1T | 6 | PROE_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
13 | 1T | 7 | PROE_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
13 | 1T | 8 | PDES_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
13 | 1T | 9 | PDES_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
13 | 1T | 10 | PFS_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
13 | 1T | 11 | PFS_T_L | -4 | -820 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | ON | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON
13 | 1T | 12 | PF1_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
13 | 1T | 13 | PF1_T_L | -4 | -820 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | ON | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON
13 | 1T | 14 | PF2_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
13 | 1T | 15 | PF2_T_L | -4 | -820 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | ON | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON
14 | 1 | 0 | VGGCL_T_H | -3.75 | -768 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ON | OFF | ON | OFF | ON | ON | ON | ON | ON | ON | ON | ON
14 | 1 | 1 | VGGCL_T_L | -4.5 | -922 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | ON | OFF | ON | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON
14 | 1 | 2 | VROWON_T_H | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
14 | 1 | 3 | VROWON_T_L | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
14 | 1 | 4 | VRSTR_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
14 | 1 | 5 | VRSTR_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
14 | 1 | 6 | VRSTG_T_H | -2 | -410 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON
14 | 1 | 7 | VRSTG_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
14 | 1 | 8 | VIREF_T | -1.8 | -369 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | ON | OFF | OFF | ON | OFF | ON | ON | ON | OFF | OFF | OFF | OFF
14 | 1 | 9 | VDDOUT_T | -1.3 | -267 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | ON | OFF | OFF | ON | OFF | OFF | OFF | OFF | ON | OFF | ON | OFF
14 | 1 | 10 | VDETCOM_T | -3.75 | -768 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ON | OFF | ON | OFF | ON | ON | ON | ON | ON | ON | ON | ON
14 | 1 | 11 | VDDUC_T | -3.75 | -768 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | ON | OFF | ON | OFF | ON | ON | ON | ON | ON | ON | ON | ON
14 | 1 | 12 | VNCOL_T | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
14 | 1 | 13 | VNROW_T | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF
14 | 1 | 14 | PSS_T_H | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | OFF | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON | ON
14 | 1 | 15 | PSS_T_L | -6 | -1229 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | ON | ON | OFF | OFF | ON | ON | OFF | OFF | ON | ON | OFF | OFF

## backplane
[12 empty] | VIEW FROM BACK | [83 empty] | VIEW FROM BACK
CLOCK/BIAS BOARD |  |  |  |  | SLOT14 | [7 empty] | CLOCK/BIAS BOARD |  |  |  |  | SLOT13 | [6 empty] | ADC 8 CHL BOARD PINOUT |  |  |  |  | SLOT12 | [6 empty] | PREAMPLIFIER BOARD PINOUT |  |  |  |  | SLOT11 | [6 empty] | ADC 8 CHL BOARD PINOUT |  |  |  |  | SLOT10 | [6 empty] | PREAMPLIFIER BOARD PINOUT |  |  |  |  | SLOT9 | [5 empty] | FIBRE CHANNEL BOARD |  |  |  |  | SLOT7 | [6 empty] | CLOCK/BIAS BOARD |  |  |  |  | SLOT6 | [7 empty] | CLOCK/BIAS BOARD |  |  |  |  | SLOT5 | [6 empty] | ADC 8 CHL BOARD PINOUT |  |  |  |  | SLOT4 | [6 empty] | PREAMPLIFIER BOARD PINOUT |  |  |  |  | SLOT3 | [6 empty] | ADC 8 CHL BOARD PINOUT |  |  |  |  | SLOT2 | [6 empty] | PREAMPLIFIER BOARD PINOUT |  |  |  |  | SLOT1
| SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY
P1 | NAME | ROW C | VOLTAGE | P1 | NAME | ROW B | VOLTAGE | P1 | NAME | ROW A | VOLTAGE |  | P1 | NAME | ROW C | VOLTAGE | P1 | NAME | ROW B | VOLTAGE | P1 | NAME | ROW A | VOLTAGE | P1 | NAME | ROW C | VOLTAGE | P1 | NAME | ROW B | VOLTAGE | P1 | NAME | ROW A | VOLTAGE | P1 | NAME | ROW C | VOLTAGE | P1 | NAME | ROW B | VOLTAGE | P1 | NAME | ROW A | VOLTAGE | P1 | NAME | ROW C | VOLTAGE | P1 | NAME | ROW B | VOLTAGE | P1 | NAME | ROW A | VOLTAGE | P1 | NAME | ROW C | VOLTAGE | P1 | NAME | ROW B | VOLTAGE | P1 | NAME | ROW A | P1 | NAME | ROW C | VOLTAGE | P1 | NAME | ROW B | VOLTAGE | P1 | NAME | ROW A | VOLTAGE | P1 | NAME | ROW C | VOLTAGE | P1 | NAME | ROW B | VOLTAGE | P1 | NAME | ROW A | VOLTAGE |  | P1 | NAME | ROW C | VOLTAGE | P1 | NAME | ROW B | VOLTAGE | P1 | NAME | ROW A | VOLTAGE | P1 | NAME | ROW C | VOLTAGE | P1 | NAME | ROW B | VOLTAGE | P1 | NAME | ROW A | VOLTAGE | P1 | NAME | ROW C | VOLTAGE | P1 | NAME | ROW B | VOLTAGE | P1 | NAME | ROW A | VOLTAGE | P1 | NAME | ROW C | VOLTAGE | P1 | NAME | ROW B | VOLTAGE | P1 | NAME | ROW A | VOLTAGE | P1 | NAME | ROW C | VOLTAGE | P1 | NAME | ROW B | VOLTAGE | P1 | NAME | ROW A | VOLTAGE
C1 |  |  |  | B1 |  | [+5V] |  | A1 |  | [+15V] |  |  | C1 |  |  |  | B1 |  | [+5V] |  | A1 |  | [+15V] |  | C1 |  |  |  | B1 |  | [+5V] |  | A1 |  | [+15V] |  | C1 |  |  |  | B1 |  |  |  | A1 |  | [+15V] |  | C1 |  |  |  | B1 |  | [+5V] |  | A1 |  | [+15V] |  | C1 |  |  |  | B1 |  |  |  | A1 |  | [+15V] | C1 |  |  |  | B1 |  | [+5V] |  | A1 |  | [+15V] |  | C1 |  |  |  | B1 |  | [+5V] |  | A1 |  | [+15V] |  |  | C1 |  |  |  | B1 |  | [+5V] |  | A1 |  | [+15V] |  | C1 |  |  |  | B1 |  | [+5V] |  | A1 |  | [+15V] |  | C1 |  |  |  | B1 |  |  |  | A1 |  | [+15V] |  | C1 |  |  |  | B1 |  | [+5V] |  | A1 |  | [+15V] |  | C1 |  |  |  | B1 |  |  |  | A1 |  | [+15V]
C2 |  |  |  | B2 |  | [DGND] |  | A2 |  | [AGND] |  |  | C2 |  |  |  | B2 |  | [DGND] |  | A2 |  | [AGND] |  | C2 |  |  |  | B2 |  | [DGND] |  | A2 |  | [AGND] |  | C2 |  |  |  | B2 |  |  |  | A2 |  | [AGND] |  | C2 |  |  |  | B2 |  | [DGND] |  | A2 |  | [AGND] |  | C2 |  |  |  | B2 |  |  |  | A2 |  | [AGND] | C2 |  |  |  | B2 |  | [DGND] |  | A2 |  | [AGND] |  | C2 |  |  |  | B2 |  | [DGND] |  | A2 |  | [AGND] |  |  | C2 |  |  |  | B2 |  | [DGND] |  | A2 |  | [AGND] |  | C2 |  |  |  | B2 |  | [DGND] |  | A2 |  | [AGND] |  | C2 |  |  |  | B2 |  |  |  | A2 |  | [AGND] |  | C2 |  |  |  | B2 |  | [DGND] |  | A2 |  | [AGND] |  | C2 |  |  |  | B2 |  |  |  | A2 |  | [AGND]
C3 |  |  |  | B3 |  | RB_RD! |  | A3 |  |  |  |  | C3 |  |  |  | B3 |  | RB_RD! |  | A3 |  |  |  | C3 |  |  |  | B3 |  | RB_RD! |  | A3 |  |  |  | C3 | Q1OUT1+ | SIG1+LN |  | B3 |  |  |  | A3 |  | SIG1-LN |  | C3 |  |  |  | B3 |  | RB_RD! |  | A3 |  |  |  | C3 | Q2OUT1+ | SIG1+LN |  | B3 |  |  |  | A3 |  | SIG1-LN | C3 |  |  |  | B3 |  | RB_RD! |  | A3 |  |  |  | C3 |  |  |  | B3 |  | RB_RD! |  | A3 |  |  |  |  | C3 |  |  |  | B3 |  | RB_RD! |  | A3 |  |  |  | C3 |  |  |  | B3 |  | RB_RD! |  | A3 |  |  |  | C3 | Q3OUT1+ | SIG1+LN |  | B3 |  |  |  | A3 |  | SIG1-LN |  | C3 |  |  |  | B3 |  | RB_RD! |  | A3 |  |  |  | C3 | Q4OUT1+ | SIG1+LN |  | B3 |  |  |  | A3 |  | SIG1-LN
C4 |  |  |  | B4 |  | RB_D0 |  | A4 |  |  |  |  | C4 |  |  |  | B4 |  | RB_D0 |  | A4 |  |  |  | C4 |  |  |  | B4 |  | RB_D0 |  | A4 |  |  |  | C4 |  | AGND |  | B4 |  |  |  | A4 |  | AGND |  | C4 |  |  |  | B4 |  | RB_D16 |  | A4 |  |  |  | C4 |  | AGND |  | B4 |  |  |  | A4 |  | AGND | C4 |  | RB_D16 |  | B4 |  | RB_D0 |  | A4 |  |  |  | C4 |  |  |  | B4 |  | RB_D0 |  | A4 |  |  |  |  | C4 |  |  |  | B4 |  | RB_D0 |  | A4 |  |  |  | C4 |  |  |  | B4 |  | RB_D0 |  | A4 |  |  |  | C4 |  | AGND |  | B4 |  |  |  | A4 |  | AGND |  | C4 |  |  |  | B4 |  | RB_D0 |  | A4 |  |  |  | C4 |  | AGND |  | B4 |  |  |  | A4 |  | AGND
C5 |  |  |  | B5 |  | RB_D1 |  | A5 |  |  |  |  | C5 |  |  |  | B5 |  | RB_D1 |  | A5 |  |  |  | C5 |  |  |  | B5 |  | RB_D1 |  | A5 |  |  |  | C5 | Q1OUT2+ | SIG2+LN |  | B5 |  |  |  | A5 |  | SIG2-LN |  | C5 |  |  |  | B5 |  | RB_D17 |  | A5 |  |  |  | C5 | Q2OUT2+ | SIG2+LN |  | B5 |  |  |  | A5 |  | SIG2-LN | C5 |  | RB_D17 |  | B5 |  | RB_D1 |  | A5 |  |  |  | C5 |  |  |  | B5 |  | RB_D1 |  | A5 |  |  |  |  | C5 |  |  |  | B5 |  | RB_D1 |  | A5 |  |  |  | C5 |  |  |  | B5 |  | RB_D1 |  | A5 |  |  |  | C5 | Q3OUT2+ | SIG2+LN |  | B5 |  |  |  | A5 |  | SIG2-LN |  | C5 |  |  |  | B5 |  | RB_D1 |  | A5 |  |  |  | C5 | Q4OUT2+ | SIG2+LN |  | B5 |  |  |  | A5 |  | SIG2-LN
C6 |  |  |  | B6 |  | RB_D2 |  | A6 |  |  |  |  | C6 |  |  |  | B6 |  | RB_D2 |  | A6 |  |  |  | C6 |  |  |  | B6 |  | RB_D2 |  | A6 |  |  |  | C6 |  | AGND |  | B6 |  |  |  | A6 |  | AGND |  | C6 |  |  |  | B6 |  | RB_D18 |  | A6 |  |  |  | C6 |  | AGND |  | B6 |  |  |  | A6 |  | AGND | C6 |  | RB_D18 |  | B6 |  | RB_D2 |  | A6 |  |  |  | C6 |  |  |  | B6 |  | RB_D2 |  | A6 |  |  |  |  | C6 |  |  |  | B6 |  | RB_D2 |  | A6 |  |  |  | C6 |  |  |  | B6 |  | RB_D2 |  | A6 |  |  |  | C6 |  | AGND |  | B6 |  |  |  | A6 |  | AGND |  | C6 |  |  |  | B6 |  | RB_D2 |  | A6 |  |  |  | C6 |  | AGND |  | B6 |  |  |  | A6 |  | AGND
C7 |  |  |  | B7 |  | RB_D3 |  | A7 |  |  |  |  | C7 |  |  |  | B7 |  | RB_D3 |  | A7 |  |  |  | C7 |  |  |  | B7 |  | RB_D3 |  | A7 |  |  |  | C7 | Q1OUT3+ | SIG3+LN |  | B7 |  |  |  | A7 |  | SIG3-LN |  | C7 |  |  |  | B7 |  | RB_D19 |  | A7 |  |  |  | C7 | Q2OUT3+ | SIG3+LN |  | B7 |  |  |  | A7 |  | SIG3-LN | C7 |  | RB_D19 |  | B7 |  | RB_D3 |  | A7 |  |  |  | C7 |  |  |  | B7 |  | RB_D3 |  | A7 |  |  |  |  | C7 |  |  |  | B7 |  | RB_D3 |  | A7 |  |  |  | C7 |  |  |  | B7 |  | RB_D3 |  | A7 |  |  |  | C7 | Q3OUT3+ | SIG3+LN |  | B7 |  |  |  | A7 |  | SIG3-LN |  | C7 |  |  |  | B7 |  | RB_D3 |  | A7 |  |  |  | C7 | Q4OUT3+ | SIG3+LN |  | B7 |  |  |  | A7 |  | SIG3-LN
C8 |  |  |  | B8 |  | RB_D4 |  | A8 |  |  |  |  | C8 |  |  |  | B8 |  | RB_D4 |  | A8 |  |  |  | C8 |  |  |  | B8 |  | RB_D4 |  | A8 |  |  |  | C8 |  | AGND |  | B8 |  |  |  | A8 |  | AGND |  | C8 |  |  |  | B8 |  | RB_D20 |  | A8 |  |  |  | C8 |  | AGND |  | B8 |  |  |  | A8 |  | AGND | C8 |  | RB_D20 |  | B8 |  | RB_D4 |  | A8 |  |  |  | C8 |  |  |  | B8 |  | RB_D4 |  | A8 |  |  |  |  | C8 |  |  |  | B8 |  | RB_D4 |  | A8 |  |  |  | C8 |  |  |  | B8 |  | RB_D4 |  | A8 |  |  |  | C8 |  | AGND |  | B8 |  |  |  | A8 |  | AGND |  | C8 |  |  |  | B8 |  | RB_D4 |  | A8 |  |  |  | C8 |  | AGND |  | B8 |  |  |  | A8 |  | AGND
C9 |  |  |  | B9 |  | RB_D5 |  | A9 |  |  |  |  | C9 |  |  |  | B9 |  | RB_D5 |  | A9 |  |  |  | C9 |  |  |  | B9 |  | RB_D5 |  | A9 |  |  |  | C9 | Q1OUT4+ | SIG4+LN |  | B9 |  |  |  | A9 |  | SIG4-LN |  | C9 |  |  |  | B9 |  | RB_D21 |  | A9 |  |  |  | C9 | Q2OUT4+ | SIG4+LN |  | B9 |  |  |  | A9 |  | SIG4-LN | C9 |  | RB_D21 |  | B9 |  | RB_D5 |  | A9 |  |  |  | C9 |  |  |  | B9 |  | RB_D5 |  | A9 |  |  |  |  | C9 |  |  |  | B9 |  | RB_D5 |  | A9 |  |  |  | C9 |  |  |  | B9 |  | RB_D5 |  | A9 |  |  |  | C9 | Q3OUT4+ | SIG4+LN |  | B9 |  |  |  | A9 |  | SIG4-LN |  | C9 |  |  |  | B9 |  | RB_D5 |  | A9 |  |  |  | C9 | Q4OUT4+ | SIG4+LN |  | B9 |  |  |  | A9 |  | SIG4-LN
C10 |  |  |  | B10 |  | RB_D6 |  | A10 |  |  |  |  | C10 |  |  |  | B10 |  | RB_D6 |  | A10 |  |  |  | C10 |  |  |  | B10 |  | RB_D6 |  | A10 |  |  |  | C10 |  | AGND |  | B10 |  |  |  | A10 |  | AGND |  | C10 |  |  |  | B10 |  | RB_D22 |  | A10 |  |  |  | C10 |  | AGND |  | B10 |  |  |  | A10 |  | AGND | C10 |  | RB_D22 |  | B10 |  | RB_D6 |  | A10 |  |  |  | C10 |  |  |  | B10 |  | RB_D6 |  | A10 |  |  |  |  | C10 |  |  |  | B10 |  | RB_D6 |  | A10 |  |  |  | C10 |  |  |  | B10 |  | RB_D6 |  | A10 |  |  |  | C10 |  | AGND |  | B10 |  |  |  | A10 |  | AGND |  | C10 |  |  |  | B10 |  | RB_D6 |  | A10 |  |  |  | C10 |  | AGND |  | B10 |  |  |  | A10 |  | AGND
C11 |  |  |  | B11 |  | RB_D7 |  | A11 |  |  |  |  | C11 |  |  |  | B11 |  | RB_D7 |  | A11 |  |  |  | C11 |  |  |  | B11 |  | RB_D7 |  | A11 |  |  |  | C11 | Q1OUT5+ | SIG5+LN |  | B11 |  |  |  | A11 |  | SIG5-LN |  | C11 |  |  |  | B11 |  | RB_D23 |  | A11 |  |  |  | C11 | Q2OUT5+ | SIG5+LN |  | B11 |  |  |  | A11 |  | SIG5-LN | C11 |  | RB_D23 |  | B11 |  | RB_D7 |  | A11 |  |  |  | C11 |  |  |  | B11 |  | RB_D7 |  | A11 |  |  |  |  | C11 |  |  |  | B11 |  | RB_D7 |  | A11 |  |  |  | C11 |  |  |  | B11 |  | RB_D7 |  | A11 |  |  |  | C11 | Q3OUT5+ | SIG5+LN |  | B11 |  |  |  | A11 |  | SIG5-LN |  | C11 |  |  |  | B11 |  | RB_D7 |  | A11 |  |  |  | C11 | Q4OUT5+ | SIG5+LN |  | B11 |  |  |  | A11 |  | SIG5-LN
C12 |  |  |  | B12 |  | [GND] |  | A12 |  | [AGND] |  |  | C12 |  |  |  | B12 |  | [DGND] |  | A12 |  | [AGND] |  | C12 |  |  |  | B12 |  | [DGND] |  | A12 |  | [AGND] |  | C12 |  | AGND |  | B12 |  |  |  | A12 |  | [AGND] |  | C12 |  |  |  | B12 |  | DGND |  | A12 |  | [AGND] |  | C12 |  | AGND |  | B12 |  |  |  | A12 |  | [AGND] | C12 |  | DGND |  | B12 |  | [DGND] |  | A12 |  | [AGND] |  | C12 |  |  |  | B12 |  | [GND] |  | A12 |  | [AGND] |  |  | C12 |  |  |  | B12 |  | [DGND] |  | A12 |  | [AGND] |  | C12 |  |  |  | B12 |  | [DGND] |  | A12 |  | [AGND] |  | C12 |  | AGND |  | B12 |  |  |  | A12 |  | [AGND] |  | C12 |  |  |  | B12 |  | [DGND] |  | A12 |  | [AGND] |  | C12 |  | AGND |  | B12 |  |  |  | A12 |  | [AGND]
C13 |  |  |  | B13 |  | [+5V] |  | A13 |  |  |  |  | C13 |  |  |  | B13 |  | [+5V] |  | A13 |  |  |  | C13 |  |  |  | B13 |  | [+5V] |  | A13 |  |  |  | C13 | Q1OUT6+ | SIG6+LN |  | B13 |  |  |  | A13 |  | SIG6-LN |  | C13 |  |  |  | B13 |  | [+5V] |  | A13 |  |  |  | C13 | Q2OUT6+ | SIG6+LN |  | B13 |  |  |  | A13 |  | SIG6-LN | C13 |  | +5V |  | B13 |  | [+5V] |  | A13 |  |  |  | C13 |  |  |  | B13 |  | [+5V] |  | A13 |  |  |  |  | C13 |  |  |  | B13 |  | [+5V] |  | A13 |  |  |  | C13 |  |  |  | B13 |  | [+5V] |  | A13 |  |  |  | C13 | Q3OUT6+ | SIG6+LN |  | B13 |  |  |  | A13 |  | SIG6-LN |  | C13 |  |  |  | B13 |  | [+5V] |  | A13 |  |  |  | C13 | Q4OUT6+ | SIG6+LN |  | B13 |  |  |  | A13 |  | SIG6-LN
C14 |  |  |  | B14 |  | RB_D8 |  | A14 |  |  |  |  | C14 |  |  |  | B14 |  | RB_D8 |  | A14 |  |  |  | C14 |  |  |  | B14 |  | RB_D8 |  | A14 |  |  |  | C14 |  | AGND |  | B14 |  |  |  | A14 |  | AGND |  | C14 |  |  |  | B14 |  | RB_D24 |  | A14 |  |  |  | C14 |  | AGND |  | B14 |  |  |  | A14 |  | AGND | C14 |  | RB_D24 |  | B14 |  | RB_D8 |  | A14 |  |  |  | C14 |  |  |  | B14 |  | RB_D8 |  | A14 |  |  |  |  | C14 |  |  |  | B14 |  | RB_D8 |  | A14 |  |  |  | C14 |  |  |  | B14 |  | RB_D8 |  | A14 |  |  |  | C14 |  | AGND |  | B14 |  |  |  | A14 |  | AGND |  | C14 |  |  |  | B14 |  | RB_D8 |  | A14 |  |  |  | C14 |  | AGND |  | B14 |  |  |  | A14 |  | AGND
C15 |  |  |  | B15 |  | RB_D9 |  | A15 |  |  |  |  | C15 |  |  |  | B15 |  | RB_D9 |  | A15 |  |  |  | C15 |  |  |  | B15 |  | RB_D9 |  | A15 |  |  |  | C15 | Q1OUT7+ | SIG7+LN |  | B15 |  |  |  | A15 |  | SIG7-LN |  | C15 |  |  |  | B15 |  | RB_D25 |  | A15 |  |  |  | C15 | Q2OUT7+ | SIG7+LN |  | B15 |  |  |  | A15 |  | SIG7-LN | C15 |  | RB_D25 |  | B15 |  | RB_D9 |  | A15 |  |  |  | C15 |  |  |  | B15 |  | RB_D9 |  | A15 |  |  |  |  | C15 |  |  |  | B15 |  | RB_D9 |  | A15 |  |  |  | C15 |  |  |  | B15 |  | RB_D9 |  | A15 |  |  |  | C15 | Q3OUT7+ | SIG7+LN |  | B15 |  |  |  | A15 |  | SIG7-LN |  | C15 |  |  |  | B15 |  | RB_D9 |  | A15 |  |  |  | C15 | Q4OUT7+ | SIG7+LN |  | B15 |  |  |  | A15 |  | SIG7-LN
C16 |  |  |  | B16 |  | RB_D10 |  | A16 |  |  |  |  | C16 |  |  |  | B16 |  | RB_D10 |  | A16 |  |  |  | C16 |  |  |  | B16 |  | RB_D10 |  | A16 |  |  |  | C16 |  | AGND |  | B16 |  |  |  | A16 |  | AGND |  | C16 |  |  |  | B16 |  | RB_D26 |  | A16 |  |  |  | C16 |  | AGND |  | B16 |  |  |  | A16 |  | AGND | C16 |  | RB_D26 |  | B16 |  | RB_D10 |  | A16 |  |  |  | C16 |  |  |  | B16 |  | RB_D10 |  | A16 |  |  |  |  | C16 |  |  |  | B16 |  | RB_D10 |  | A16 |  |  |  | C16 |  |  |  | B16 |  | RB_D10 |  | A16 |  |  |  | C16 |  | AGND |  | B16 |  |  |  | A16 |  | AGND |  | C16 |  |  |  | B16 |  | RB_D10 |  | A16 |  |  |  | C16 |  | AGND |  | B16 |  |  |  | A16 |  | AGND
C17 | VGGCL_T | CK0BS0 |  | B17 |  | RB_D11 |  | A17 |  | AGND |  |  | C17 | VDDCL_T | CK0BS0 |  | B17 |  | RB_D11 |  | A17 |  | AGND |  | C17 |  |  |  | B17 |  | RB_D11 |  | A17 |  |  |  | C17 | Q1OUT8+ | SIG8+LN |  | B17 |  |  |  | A17 |  | SIG8-LN |  | C17 |  |  |  | B17 |  | RB_D27 |  | A17 |  |  |  | C17 | Q2OUT8+ | SIG8+LN |  | B17 |  |  |  | A17 |  | SIG8-LN | C17 |  | RB_D27 |  | B17 |  | RB_D11 |  | A17 |  |  |  | C17 | VGGCL_B | CK0BS0 |  | B17 |  | RB_D11 |  | A17 |  | AGND |  |  | C17 | VDDCL_B | CK0BS0 |  | B17 |  | RB_D11 |  | A17 |  | AGND |  | C17 |  |  |  | B17 |  | RB_D11 |  | A17 |  |  |  | C17 | Q3OUT8+ | SIG8+LN |  | B17 |  |  |  | A17 |  | SIG8-LN |  | C17 |  |  |  | B17 |  | RB_D11 |  | A17 |  |  |  | C17 | Q4OUT8+ | SIG8+LN |  | B17 |  |  |  | A17 |  | SIG8-LN
C18 |  | BIAS1 |  | B18 |  | RB_D12 |  | A18 |  | AGND |  |  | C18 |  | BIAS1 |  | B18 |  | RB_D12 |  | A18 |  | AGND |  | C18 |  |  |  | B18 |  | RB_D12 |  | A18 |  |  |  | C18 |  | AGND |  | B18 |  |  |  | A18 |  | AGND |  | C18 |  |  |  | B18 |  | RB_D28 |  | A18 |  |  |  | C18 |  | AGND |  | B18 |  |  |  | A18 |  | AGND | C18 |  | RB_D28 |  | B18 |  | RB_D12 |  | A18 |  |  |  | C18 |  | BIAS1 |  | B18 |  | RB_D12 |  | A18 |  | AGND |  |  | C18 |  | BIAS1 |  | B18 |  | RB_D12 |  | A18 |  | AGND |  | C18 |  |  |  | B18 |  | RB_D12 |  | A18 |  |  |  | C18 |  | AGND |  | B18 |  |  |  | A18 |  | AGND |  | C18 |  |  |  | B18 |  | RB_D12 |  | A18 |  |  |  | C18 |  | AGND |  | B18 |  |  |  | A18 |  | AGND
C19 | VROWON_T | CK1BS2 |  | B19 |  | RB_D13 |  | A19 |  | AGND |  |  | C19 | PS1_T | CK1BS2 |  | B19 |  | RB_D13 |  | A19 |  | AGND |  | C19 |  |  |  | B19 |  | RB_D13 |  | A19 |  |  |  | C19 |  |  |  | B19 |  |  |  | A19 |  |  |  | C19 |  |  |  | B19 |  | RB_D29 |  | A19 |  |  |  | C19 |  |  |  | B19 |  |  |  | A19 |  |  | C19 |  | RB_D29 |  | B19 |  | RB_D13 |  | A19 |  |  |  | C19 | VROWON_B | CK1BS2 |  | B19 |  | RB_D13 |  | A19 |  | AGND |  |  | C19 | PS1_B | CK1BS2 |  | B19 |  | RB_D13 |  | A19 |  | AGND |  | C19 |  |  |  | B19 |  | RB_D13 |  | A19 |  |  |  | C19 |  |  |  | B19 |  |  |  | A19 |  |  |  | C19 |  |  |  | B19 |  | RB_D13 |  | A19 |  |  |  | C19 |  |  |  | B19 |  |  |  | A19
C20 |  | BIAS3 |  | B20 |  | RB_D14 |  | A20 |  | AGND |  |  | C20 |  | BIAS3 |  | B20 |  | RB_D14 |  | A20 |  | AGND |  | C20 |  |  |  | B20 |  | RB_D14 |  | A20 |  |  |  | C20 |  |  |  | B20 |  |  |  | A20 |  |  |  | C20 |  |  |  | B20 |  | RB_D30 |  | A20 |  |  |  | C20 |  |  |  | B20 |  |  |  | A20 |  |  | C20 |  | RB_D30 |  | B20 |  | RB_D14 |  | A20 |  |  |  | C20 |  | BIAS3 |  | B20 |  | RB_D14 |  | A20 |  | AGND |  |  | C20 |  | BIAS3 |  | B20 |  | RB_D14 |  | A20 |  | AGND |  | C20 |  |  |  | B20 |  | RB_D14 |  | A20 |  |  |  | C20 |  |  |  | B20 |  |  |  | A20 |  |  |  | C20 |  |  |  | B20 |  | RB_D14 |  | A20 |  |  |  | C20 |  |  |  | B20 |  |  |  | A20
C21 | VRSTR_T | CK2BS4 |  | B21 |  | RB_D15 |  | A21 |  | AGND |  |  | C21 | PS2_T | CK2BS4 |  | B21 |  | RB_D15 |  | A21 |  | AGND |  | C21 |  |  |  | B21 |  | RB_D15 |  | A21 |  |  |  | C21 |  |  |  | B21 |  |  |  | A21 |  |  |  | C21 |  |  |  | B21 |  | RB_D31 |  | A21 |  |  |  | C21 |  |  |  | B21 |  |  |  | A21 |  |  | C21 |  | RB_D31 |  | B21 |  | RB_D15 |  | A21 |  |  |  | C21 | VRSTR_B | CK2BS4 |  | B21 |  | RB_D15 |  | A21 |  | AGND |  |  | C21 | PS2_B | CK2BS4 |  | B21 |  | RB_D15 |  | A21 |  | AGND |  | C21 |  |  |  | B21 |  | RB_D15 |  | A21 |  |  |  | C21 |  |  |  | B21 |  |  |  | A21 |  |  |  | C21 |  |  |  | B21 |  | RB_D15 |  | A21 |  |  |  | C21 |  |  |  | B21 |  |  |  | A21
C22 |  | BIAS5 |  | B22 |  | [GND] |  | A22 |  | [AGND] |  |  | C22 |  | BIAS5 |  | B22 |  | [DGND] |  | A22 |  | [AGND] |  | C22 |  |  |  | B22 |  | [DGND] |  | A22 |  | [AGND] |  | C22 |  |  |  | B22 |  |  |  | A22 |  | [AGND] |  | C22 |  |  |  | B22 |  | [DGND] |  | A22 |  | [AGND] |  | C22 |  |  |  | B22 |  |  |  | A22 |  | [AGND] | C22 |  |  |  | B22 |  | [DGND] |  | A22 |  | [AGND] |  | C22 |  | BIAS5 |  | B22 |  | [GND] |  | A22 |  | [AGND] |  |  | C22 |  | BIAS5 |  | B22 |  | [DGND] |  | A22 |  | [AGND] |  | C22 |  |  |  | B22 |  | [DGND] |  | A22 |  | [AGND] |  | C22 |  |  |  | B22 |  |  |  | A22 |  | [AGND] |  | C22 |  |  |  | B22 |  | [DGND] |  | A22 |  | [AGND] |  | C22 |  |  |  | B22 |  |  |  | A22 |  | [AGND]
C23 | VRSTG_T | CK3BS6 |  | B23 |  | RB_SD0 |  | A23 |  | AGND |  |  | C23 | PROE_T | CK3BS6 |  | B23 |  | RB_SD0 |  | A23 |  | AGND |  | C23 |  |  |  | B23 |  | RB_SD0 |  | A23 |  |  |  | C23 |  | VREF+ |  | B23 |  |  |  | A23 |  | AGND |  | C23 |  |  |  | B23 |  | RB_SD8 |  | A23 |  |  |  | C23 |  | VREF+ |  | B23 |  |  |  | A23 |  | AGND | C23 |  |  |  | B23 |  | RB_SD0 |  | A23 |  |  |  | C23 | VRSTG_B | CK3BS6 |  | B23 |  | RB_SD0 |  | A23 |  | AGND |  |  | C23 | PROE_B | CK3BS6 |  | B23 |  | RB_SD0 |  | A23 |  | AGND |  | C23 |  |  |  | B23 |  | RB_SD0 |  | A23 |  |  |  | C23 |  | VREF+ |  | B23 |  |  |  | A23 |  | AGND |  | C23 |  |  |  | B23 |  | RB_SD0 |  | A23 |  |  |  | C23 |  | VREF+ |  | B23 |  |  |  | A23 |  | AGND
C24 |  | BIAS7 |  | B24 |  | RB_SD1 |  | A24 |  | AGND |  |  | C24 |  | BIAS7 |  | B24 |  | RB_SD1 |  | A24 |  | AGND |  | C24 |  |  |  | B24 |  | RB_SD1 |  | A24 |  |  |  | C24 |  | OFFSET1 |  | B24 |  |  |  | A24 |  | AGND |  | C24 |  |  |  | B24 |  | RB_SD9 |  | A24 |  |  |  | C24 |  | OFFSET1 |  | B24 |  |  |  | A24 |  | AGND | C24 |  |  |  | B24 |  | RB_SD1 |  | A24 |  |  |  | C24 |  | BIAS7 |  | B24 |  | RB_SD1 |  | A24 |  | AGND |  |  | C24 |  | BIAS7 |  | B24 |  | RB_SD1 |  | A24 |  | AGND |  | C24 |  |  |  | B24 |  | RB_SD1 |  | A24 |  |  |  | C24 |  | OFFSET1 |  | B24 |  |  |  | A24 |  | AGND |  | C24 |  |  |  | B24 |  | RB_SD1 |  | A24 |  |  |  | C24 |  | OFFSET1 |  | B24 |  |  |  | A24 |  | AGND
C25 | VREF_T | CK4BS8 |  | B25 |  | RB_SD2 |  | A25 |  | AGND |  |  | C25 | PDES_T | CK4BS8 |  | B25 |  | RB_SD2 |  | A25 |  | AGND |  | C25 |  |  |  | B25 |  | RB_SD2 |  | A25 |  |  |  | C25 |  | OFFSET2 |  | B25 |  |  |  | A25 |  | AGND |  | C25 |  |  |  | B25 |  | RB_SD10 |  | A25 |  |  |  | C25 |  | OFFSET2 |  | B25 |  |  |  | A25 |  | AGND | C25 |  |  |  | B25 |  | RB_SD2 |  | A25 |  |  |  | C25 | VREF_B | CK4BS8 |  | B25 |  | RB_SD2 |  | A25 |  | AGND |  |  | C25 | PDES_B | CK4BS8 |  | B25 |  | RB_SD2 |  | A25 |  | AGND |  | C25 |  |  |  | B25 |  | RB_SD2 |  | A25 |  |  |  | C25 |  | OFFSET2 |  | B25 |  |  |  | A25 |  | AGND |  | C25 |  |  |  | B25 |  | RB_SD2 |  | A25 |  |  |  | C25 |  | OFFSET2 |  | B25 |  |  |  | A25 |  | AGND
C26 | VDDOUT_T | BIAS9 |  | B26 |  | RB_SD3 |  | A26 |  | AGND |  |  | C26 |  | BIAS9 |  | B26 |  | RB_SD3 |  | A26 |  | AGND |  | C26 |  |  |  | B26 |  | RB_SD3 |  | A26 |  |  |  | C26 |  | OFFSET3 |  | B26 |  |  |  | A26 |  | AGND |  | C26 |  |  |  | B26 |  | RB_SD11 |  | A26 |  |  |  | C26 |  | OFFSET3 |  | B26 |  |  |  | A26 |  | AGND | C26 |  |  |  | B26 |  | RB_SD3 |  | A26 |  |  |  | C26 | VDDOUT_B | BIAS9 |  | B26 |  | RB_SD3 |  | A26 |  | AGND |  |  | C26 |  | BIAS9 |  | B26 |  | RB_SD3 |  | A26 |  | AGND |  | C26 |  |  |  | B26 |  | RB_SD3 |  | A26 |  |  |  | C26 |  | OFFSET3 |  | B26 |  |  |  | A26 |  | AGND |  | C26 |  |  |  | B26 |  | RB_SD3 |  | A26 |  |  |  | C26 |  | OFFSET3 |  | B26 |  |  |  | A26 |  | AGND
C27 | VDETCOM_T | CK5BS10 |  | B27 |  | RB_SD4 |  | A27 |  | AGND |  |  | C27 | PFS_T | CK5BS10 |  | B27 |  | RB_SD4 |  | A27 |  | AGND |  | C27 |  |  |  | B27 |  | RB_SD4 |  | A27 |  |  |  | C27 |  | OFFSET4 |  | B27 |  |  |  | A27 |  | AGND |  | C27 |  |  |  | B27 |  | RB_SD12 |  | A27 |  |  |  | C27 |  | OFFSET4 |  | B27 |  |  |  | A27 |  | AGND | C27 |  |  |  | B27 |  | RB_SD4 |  | A27 |  |  |  | C27 | VDETCOM_B | CK5BS10 |  | B27 |  | RB_SD4 |  | A27 |  | AGND |  |  | C27 | PFS_B | CK5BS10 |  | B27 |  | RB_SD4 |  | A27 |  | AGND |  | C27 |  |  |  | B27 |  | RB_SD4 |  | A27 |  |  |  | C27 |  | OFFSET4 |  | B27 |  |  |  | A27 |  | AGND |  | C27 |  |  |  | B27 |  | RB_SD4 |  | A27 |  |  |  | C27 |  | OFFSET4 |  | B27 |  |  |  | A27 |  | AGND
C28 | VDDUC_T | BIAS11 |  | B28 |  | RB_SD5 |  | A28 |  | AGND |  |  | C28 |  | BIAS11 |  | B28 |  | RB_SD5 |  | A28 |  | AGND |  | C28 |  |  |  | B28 |  | RB_SD5 |  | A28 |  |  |  | C28 |  | OFFSET5 |  | B28 |  |  |  | A28 |  | AGND |  | C28 |  |  |  | B28 |  | RB_SD13 |  | A28 |  |  |  | C28 |  | OFFSET5 |  | B28 |  |  |  | A28 |  | AGND | C28 |  |  |  | B28 |  | RB_SD5 |  | A28 |  |  |  | C28 | VDDUC_B | BIAS11 |  | B28 |  | RB_SD5 |  | A28 |  | AGND |  |  | C28 |  | BIAS11 |  | B28 |  | RB_SD5 |  | A28 |  | AGND |  | C28 |  |  |  | B28 |  | RB_SD5 |  | A28 |  |  |  | C28 |  | OFFSET5 |  | B28 |  |  |  | A28 |  | AGND |  | C28 |  |  |  | B28 |  | RB_SD5 |  | A28 |  |  |  | C28 |  | OFFSET5 |  | B28 |  |  |  | A28 |  | AGND
C29 | VNCOL_T | CK6BS12 |  | B29 |  | RB_SD6 |  | A29 |  | AGND |  |  | C29 | PF1_T | CK6BS12 |  | B29 |  | RB_SD6 |  | A29 |  | AGND |  | C29 |  |  |  | B29 |  | RB_SD6 |  | A29 |  |  |  | C29 |  | OFFSET6 |  | B29 |  |  |  | A29 |  | AGND |  | C29 |  |  |  | B29 |  | RB_SD14 |  | A29 |  |  |  | C29 |  | OFFSET6 |  | B29 |  |  |  | A29 |  | AGND | C29 |  |  |  | B29 |  | RB_SD6 |  | A29 |  |  |  | C29 | VNCOL_B | CK6BS12 |  | B29 |  | RB_SD6 |  | A29 |  | AGND |  |  | C29 | PF1_B | CK6BS12 |  | B29 |  | RB_SD6 |  | A29 |  | AGND |  | C29 |  |  |  | B29 |  | RB_SD6 |  | A29 |  |  |  | C29 |  | OFFSET6 |  | B29 |  |  |  | A29 |  | AGND |  | C29 |  |  |  | B29 |  | RB_SD6 |  | A29 |  |  |  | C29 |  | OFFSET6 |  | B29 |  |  |  | A29 |  | AGND
C30 | VNROW_T | BIAS13 |  | B30 |  | RB_SD7 |  | A30 |  | AGND |  |  | C30 |  | BIAS13 |  | B30 |  | RB_SD7 |  | A30 |  | AGND |  | C30 |  |  |  | B30 |  | RB_SD7 |  | A30 |  |  |  | C30 |  | OFFSET7 |  | B30 |  |  |  | A30 |  | AGND |  | C30 |  |  |  | B30 |  | RB_SD15 |  | A30 |  |  |  | C30 |  | OFFSET7 |  | B30 |  |  |  | A30 |  | AGND | C30 |  |  |  | B30 |  | RB_SD7 |  | A30 |  |  |  | C30 | VNROW_B | BIAS13 |  | B30 |  | RB_SD7 |  | A30 |  | AGND |  |  | C30 |  | BIAS13 |  | B30 |  | RB_SD7 |  | A30 |  | AGND |  | C30 |  |  |  | B30 |  | RB_SD7 |  | A30 |  |  |  | C30 |  | OFFSET7 |  | B30 |  |  |  | A30 |  | AGND |  | C30 |  |  |  | B30 |  | RB_SD7 |  | A30 |  |  |  | C30 |  | OFFSET7 |  | B30 |  |  |  | A30 |  | AGND
C31 | PSS_T | CK7BS14 |  | B31 |  | [GND] |  | A31 |  | [AGND] |  |  | C31 | PF2_T | CK7BS14 |  | B31 |  | [DGND] |  | A31 |  | [AGND] |  | C31 |  |  |  | B31 |  | [DGND] |  | A31 |  | [AGND] |  | C31 |  | OFFSET8 |  | B31 |  |  |  | A31 |  | [AGND] |  | C31 |  |  |  | B31 |  | [DGND] |  | A31 |  | [AGND] |  | C31 |  | OFFSET8 |  | B31 |  |  |  | A31 |  | [AGND] | C31 |  |  |  | B31 |  | [GND] |  | A31 |  | [AGND] |  | C31 | PSS_B | CK7BS14 |  | B31 |  | [GND] |  | A31 |  | [AGND] |  |  | C31 | PF2_B | CK7BS14 |  | B31 |  | [DGND] |  | A31 |  | [AGND] |  | C31 |  |  |  | B31 |  | [DGND] |  | A31 |  | [AGND] |  | C31 |  | OFFSET8 |  | B31 |  |  |  | A31 |  | [AGND] |  | C31 |  |  |  | B31 |  | [DGND] |  | A31 |  | [AGND] |  | C31 |  | OFFSET8 |  | B31 |  |  |  | A31 |  | [AGND]
C32 |  | BIAS15 |  | B32 |  | [+5V] |  | A32 |  | [-15V] |  |  | C32 |  | BIAS15 |  | B32 |  | [+5V] |  | A32 |  | [-15V] |  | C32 |  |  |  | B32 |  | [+5V] |  | A32 |  | [-15V] |  | C32 |  | VREF+ |  | B32 |  |  |  | A32 |  | [-15V] |  | C32 |  |  |  | B32 |  | [+5V] |  | A32 |  | [-15V] |  | C32 |  | VREF+ |  | B32 |  |  |  | A32 |  | [-15V] | C32 |  |  |  | B32 |  | [+5V] |  | A32 |  | [-15V] |  | C32 |  | BIAS15 |  | B32 |  | [+5V] |  | A32 |  | [-15V] |  |  | C32 |  | BIAS15 |  | B32 |  | [+5V] |  | A32 |  | [-15V] |  | C32 |  |  |  | B32 |  | [+5V] |  | A32 |  | [-15V] |  | C32 |  | VREF+ |  | B32 |  |  |  | A32 |  | [-15V] |  | C32 |  |  |  | B32 |  | [+5V] |  | A32 |  | [-15V] |  | C32 |  | VREF+ |  | B32 |  |  |  | A32 |  | [-15V]
| SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | ASSY | SIGNAL |  |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | ASSY | SIGNAL |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | ASSY | SIGNAL |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | ASSY | SIGNAL |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | ASSY | SIGNAL |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | ASSY | SIGNAL |  |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | ASSY | SIGNAL |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | ASSY | SIGNAL |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | ASSY | SIGNAL |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY |  |  | SYSTEM | ASSY
P2 | NAME | ROW C | VOLTAGE | P2 | NAME | ROW B | VOLTAGE | P2 | NAME | ROW A | VOLTAGE |  | P2 | NAME | ROW C | VOLTAGE | P2 | NAME | ROW B | VOLTAGE | P2 | NAME | ROW A | VOLTAGE | P2 | NAME | ROW C | VOLTAGE | P2 | NAME | ROW B | VOLTAGE | P2 | NAME | ROW A | NAME | P2 | NAME | ROW C | VOLTAGE | P2 | NAME | ROW B | VOLTAGE | P2 | NAME | ROW A | VOLTAGE | P2 | NAME | ROW C | VOLTAGE | P2 | NAME | ROW B | VOLTAGE | P2 | NAME | ROW A | VOLTAGE | P2 | NAME | ROW C | VOLTAGE | P2 | NAME | ROW B | VOLTAGE | P2 | NAME | ROW A | P2 | NAME | ROW C | VOLTAGE | P2 | NAME | ROW B | VOLTAGE | P2 | NAME | ROW A | VOLTAGE | P2 | NAME | ROW C | VOLTAGE | P2 | NAME | ROW B | VOLTAGE | P2 | NAME | ROW A | VOLTAGE |  | P2 | NAME | ROW C | VOLTAGE | P2 | NAME | ROW B | VOLTAGE | P2 | NAME | ROW A | VOLTAGE | P2 | NAME | ROW C | VOLTAGE | P2 | NAME | ROW B | VOLTAGE | P2 | NAME | ROW A | NAME | P2 | NAME | ROW C | VOLTAGE | P2 | NAME | ROW B | VOLTAGE | P2 | NAME | ROW A | VOLTAGE | P2 | NAME | ROW C | VOLTAGE | P2 | NAME | ROW B | VOLTAGE | P2 | NAME | ROW A | VOLTAGE | P2 | NAME | ROW C | VOLTAGE | P2 | NAME | ROW B | VOLTAGE | P2 | NAME | ROW A | VOLTAGE
C1 |  |  |  | B1 |  | [+5V] |  | A1 |  | [+15V] |  |  | C1 |  |  |  | B1 |  | [+5V] |  | A1 |  | [+15V] |  | C1 |  |  |  | B1 |  | [+5V] |  | A1 |  | [+15V] |  | C1 |  |  |  | B1 |  | [+5V] |  | A1 |  | [+15V] |  | C1 |  |  |  | B1 |  | [+5V] |  | A1 |  | [+15V] |  | C1 |  |  |  | B1 |  | [+5V] |  | A1 |  | [+15V] | C1 |  |  |  | B1 |  | [+5V] |  | A1 |  | [+15V] |  | C1 |  |  |  | B1 |  | [+5V] |  | A1 |  | [+15V] |  |  | C1 |  |  |  | B1 |  | [+5V] |  | A1 |  | [+15V] |  | C1 |  |  |  | B1 |  | [+5V] |  | A1 |  | [+15V] |  | C1 |  |  |  | B1 |  | [+5V] |  | A1 |  | [+15V] |  | C1 |  |  |  | B1 |  | [+5V] |  | A1 |  | [+15V] |  | C1 |  |  |  | B1 |  | [+5V] |  | A1 |  | [+15V]
C2 |  |  |  | B2 |  | [DGND] |  | A2 |  | [AGND] |  |  | C2 |  |  |  | B2 |  | [DGND] |  | A2 |  | [AGND] |  | C2 |  | AGND |  | B2 |  | [DGND] |  | A2 |  | [AGND] |  | C2 |  | AGND |  | B2 |  | [DGND] |  | A2 |  | [AGND] |  | C2 |  | AGND |  | B2 |  | [DGND] |  | A2 |  | [AGND] |  | C2 |  | AGND |  | B2 |  | [DGND] |  | A2 |  | [AGND] | C2 |  |  |  | B2 |  | [DGND] |  | A2 |  | [AGND] |  | C2 |  |  |  | B2 |  | [DGND] |  | A2 |  | [AGND] |  |  | C2 |  |  |  | B2 |  | [DGND] |  | A2 |  | [AGND] |  | C2 |  | AGND |  | B2 |  | [DGND] |  | A2 |  | [AGND] |  | C2 |  | AGND |  | B2 |  | [DGND] |  | A2 |  | [AGND] |  | C2 |  | AGND |  | B2 |  | [DGND] |  | A2 |  | [AGND] |  | C2 |  | AGND |  | B2 |  | [DGND] |  | A2 |  | [AGND]
C3 |  |  |  | B3 |  | CB_WR! |  | A3 |  |  |  |  | C3 |  |  |  | B3 |  | CB_WR! |  | A3 |  |  |  | C3 |  | POST1 |  | B3 |  | CB_WR! |  | A3 |  | POST1RTN |  | C3 |  | POST1 |  | B3 |  | CB_WR! |  | A3 |  | POST1RTN |  | C3 |  | POST1 |  | B3 |  | CB_WR! |  | A3 |  | POST1RTN |  | C3 |  | POST1 |  | B3 |  | CB_WR! |  | A3 |  | POST1RTN | C3 |  |  |  | B3 |  | CB_WR! |  | A3 |  |  |  | C3 |  |  |  | B3 |  | CB_WR! |  | A3 |  |  |  |  | C3 |  |  |  | B3 |  | CB_WR! |  | A3 |  |  |  | C3 |  | POST1 |  | B3 |  | CB_WR! |  | A3 |  | POST1RTN |  | C3 |  | POST1 |  | B3 |  | CB_WR! |  | A3 |  | POST1RTN |  | C3 |  | POST1 |  | B3 |  | CB_WR! |  | A3 |  | POST1RTN |  | C3 |  | POST1 |  | B3 |  | CB_WR! |  | A3 |  | POST1RTN
C4 |  |  |  | B4 | vggcl_t_clk | CB_D0 |  | A4 |  |  |  |  | C4 |  |  |  | B4 | vggcl_t_clk | CB_D0 |  | A4 |  |  |  | C4 |  | AGND |  | B4 | vggcl_t_clk | CB_D0 |  | A4 |  | AGND |  | C4 |  | AGND |  | B4 | vggcl_t_clk | CB_D0 |  | A4 |  | AGND |  | C4 |  | AGND |  | B4 | vggcl_t_clk | CB_D0 |  | A4 |  | AGND |  | C4 |  | AGND |  | B4 | vggcl_t_clk | CB_D16 |  | A4 |  | AGND | C4 | vggcl_t_clk | CB_D16 |  | B4 | vggcl_b_clk | CB_D0 |  | A4 |  |  |  | C4 |  |  |  | B4 | vggcl_b_clk | CB_D0 |  | A4 |  |  |  |  | C4 |  |  |  | B4 | vggcl_b_clk | CB_D0 |  | A4 |  |  |  | C4 |  | AGND |  | B4 | vggcl_b_clk | CB_D0 |  | A4 |  | AGND |  | C4 |  | AGND |  | B4 | vggcl_b_clk | CB_D0 |  | A4 |  | AGND |  | C4 |  | AGND |  | B4 | vggcl_b_clk | CB_D0 |  | A4 |  | AGND |  | C4 |  | AGND |  | B4 | vggcl_b_clk | CB_D0 |  | A4 |  | AGND
C5 |  |  |  | B5 | vrowon_t_clk | CB_D1 |  | A5 |  |  |  |  | C5 |  |  |  | B5 | vrowon_t_clk | CB_D1 |  | A5 |  |  |  | C5 |  | POST2 |  | B5 | vrowon_t_clk | CB_D1 |  | A5 |  | POST2RTN |  | C5 |  | POST2 |  | B5 | vrowon_t_clk | CB_D1 |  | A5 |  | POST2RTN |  | C5 |  | POST2 |  | B5 | vrowon_t_clk | CB_D1 |  | A5 |  | POST2RTN |  | C5 |  | POST2 |  | B5 | vrowon_t_clk | CB_D17 |  | A5 |  | POST2RTN | C5 | vrowon_t_clk | CB_D17 |  | B5 | vrowon_b_clk | CB_D1 |  | A5 |  |  |  | C5 |  |  |  | B5 | vrowon_b_clk | CB_D1 |  | A5 |  |  |  |  | C5 |  |  |  | B5 | vrowon_b_clk | CB_D1 |  | A5 |  |  |  | C5 |  | POST2 |  | B5 | vrowon_b_clk | CB_D1 |  | A5 |  | POST2RTN |  | C5 |  | POST2 |  | B5 | vrowon_b_clk | CB_D1 |  | A5 |  | POST2RTN |  | C5 |  | POST2 |  | B5 | vrowon_b_clk | CB_D1 |  | A5 |  | POST2RTN |  | C5 |  | POST2 |  | B5 | vrowon_b_clk | CB_D1 |  | A5 |  | POST2RTN
C6 |  |  |  | B6 | vrstr_t_clk | CB_D2 |  | A6 |  |  |  |  | C6 |  |  |  | B6 | vrstr_t_clk | CB_D2 |  | A6 |  |  |  | C6 |  | AGND |  | B6 | vrstr_t_clk | CB_D2 |  | A6 |  | AGND |  | C6 |  | AGND |  | B6 | vrstr_t_clk | CB_D2 |  | A6 |  | AGND |  | C6 |  | AGND |  | B6 | vrstr_t_clk | CB_D2 |  | A6 |  | AGND |  | C6 |  | AGND |  | B6 | vrstr_t_clk | CB_D18 |  | A6 |  | AGND | C6 | vrstr_t_clk | CB_D18 |  | B6 | vrstr_b_clk | CB_D2 |  | A6 |  |  |  | C6 |  |  |  | B6 | vrstr_b_clk | CB_D2 |  | A6 |  |  |  |  | C6 |  |  |  | B6 | vrstr_b_clk | CB_D2 |  | A6 |  |  |  | C6 |  | AGND |  | B6 | vrstr_b_clk | CB_D2 |  | A6 |  | AGND |  | C6 |  | AGND |  | B6 | vrstr_b_clk | CB_D2 |  | A6 |  | AGND |  | C6 |  | AGND |  | B6 | vrstr_b_clk | CB_D2 |  | A6 |  | AGND |  | C6 |  | AGND |  | B6 | vrstr_b_clk | CB_D2 |  | A6 |  | AGND
C7 |  |  |  | B7 | vrstg_t_clk | CB_D3 |  | A7 |  |  |  |  | C7 |  |  |  | B7 | vrstg_t_clk | CB_D3 |  | A7 |  |  |  | C7 |  | POST3 |  | B7 | vrstg_t_clk | CB_D3 |  | A7 |  | POST3RTN |  | C7 |  | POST3 |  | B7 | vrstg_t_clk | CB_D3 |  | A7 |  | POST3RTN |  | C7 |  | POST3 |  | B7 | vrstg_t_clk | CB_D3 |  | A7 |  | POST3RTN |  | C7 |  | POST3 |  | B7 | vrstg_t_clk | CB_D19 |  | A7 |  | POST3RTN | C7 | vrstg_t_clk | CB_D19 |  | B7 | vrstg_b_clk | CB_D3 |  | A7 |  |  |  | C7 |  |  |  | B7 | vrstg_b_clk | CB_D3 |  | A7 |  |  |  |  | C7 |  |  |  | B7 | vrstg_b_clk | CB_D3 |  | A7 |  |  |  | C7 |  | POST3 |  | B7 | vrstg_b_clk | CB_D3 |  | A7 |  | POST3RTN |  | C7 |  | POST3 |  | B7 | vrstg_b_clk | CB_D3 |  | A7 |  | POST3RTN |  | C7 |  | POST3 |  | B7 | vrstg_b_clk | CB_D3 |  | A7 |  | POST3RTN |  | C7 |  | POST3 |  | B7 | vrstg_b_clk | CB_D3 |  | A7 |  | POST3RTN
C8 |  |  |  | B8 | read_data_t_clk | CB_D4 |  | A8 |  |  |  |  | C8 |  |  |  | B8 | read_data_t_clk | CB_D4 |  | A8 |  |  |  | C8 |  | AGND |  | B8 | read_data_t_clk | CB_D4 |  | A8 |  | AGND |  | C8 |  | AGND |  | B8 | read_data_t_clk | CB_D4 |  | A8 |  | AGND |  | C8 |  | AGND |  | B8 | read_data_t_clk | CB_D4 |  | A8 |  | AGND |  | C8 |  | AGND |  | B8 | read_data_t_clk | CB_D20 |  | A8 |  | AGND | C8 | read_data_t_clk | CB_D20 |  | B8 | read_data_b_clk | CB_D4 |  | A8 |  |  |  | C8 |  |  |  | B8 | read_data_b_clk | CB_D4 |  | A8 |  |  |  |  | C8 |  |  |  | B8 | read_data_b_clk | CB_D4 |  | A8 |  |  |  | C8 |  | AGND |  | B8 | read_data_b_clk | CB_D4 |  | A8 |  | AGND |  | C8 |  | AGND |  | B8 | read_data_b_clk | CB_D4 |  | A8 |  | AGND |  | C8 |  | AGND |  | B8 | read_data_b_clk | CB_D4 |  | A8 |  | AGND |  | C8 |  | AGND |  | B8 | read_data_b_clk | CB_D4 |  | A8 |  | AGND
C9 |  |  |  | B9 | convert_t_clk | CB_D5 |  | A9 |  |  |  |  | C9 |  |  |  | B9 | convert_t_clk | CB_D5 |  | A9 |  |  |  | C9 |  | POST4 |  | B9 | convert_t_clk | CB_D5 |  | A9 |  | POST4RTN |  | C9 |  | POST4 |  | B9 | convert_t_clk | CB_D5 |  | A9 |  | POST4RTN |  | C9 |  | POST4 |  | B9 | convert_t_clk | CB_D5 |  | A9 |  | POST4RTN |  | C9 |  | POST4 |  | B9 | convert_t_clk | CB_D21 |  | A9 |  | POST4RTN | C9 | convert_t_clk | CB_D21 |  | B9 | convert_b_clk | CB_D5 |  | A9 |  |  |  | C9 |  |  |  | B9 | convert_b_clk | CB_D5 |  | A9 |  |  |  |  | C9 |  |  |  | B9 | convert_b_clk | CB_D5 |  | A9 |  |  |  | C9 |  | POST4 |  | B9 | convert_b_clk | CB_D5 |  | A9 |  | POST4RTN |  | C9 |  | POST4 |  | B9 | convert_b_clk | CB_D5 |  | A9 |  | POST4RTN |  | C9 |  | POST4 |  | B9 | convert_b_clk | CB_D5 |  | A9 |  | POST4RTN |  | C9 |  | POST4 |  | B9 | convert_b_clk | CB_D5 |  | A9 |  | POST4RTN
C10 |  |  |  | B10 | frame_t_clk | CB_D6 |  | A10 |  |  |  |  | C10 |  |  |  | B10 | frame_t_clk | CB_D6 |  | A10 |  |  |  | C10 |  | AGND |  | B10 | frame_t_clk | CB_D6 |  | A10 |  | AGND |  | C10 |  | AGND |  | B10 | frame_t_clk | CB_D6 |  | A10 |  | AGND |  | C10 |  | AGND |  | B10 | frame_t_clk | CB_D6 |  | A10 |  | AGND |  | C10 |  | AGND |  | B10 | frame_t_clk | CB_D22 |  | A10 |  | AGND | C10 | frame_t_clk | CB_D22 |  | B10 | frame_b_clk | CB_D6 |  | A10 |  |  |  | C10 |  |  |  | B10 | frame_b_clk | CB_D6 |  | A10 |  |  |  |  | C10 |  |  |  | B10 | frame_b_clk | CB_D6 |  | A10 |  |  |  | C10 |  | AGND |  | B10 | frame_b_clk | CB_D6 |  | A10 |  | AGND |  | C10 |  | AGND |  | B10 | frame_b_clk | CB_D6 |  | A10 |  | AGND |  | C10 |  | AGND |  | B10 | frame_b_clk | CB_D6 |  | A10 |  | AGND |  | C10 |  | AGND |  | B10 | frame_b_clk | CB_D6 |  | A10 |  | AGND
C11 |  |  |  | B11 | phiss_t_clk | CB_D7 |  | A11 |  |  |  |  | C11 |  |  |  | B11 | phiss_t_clk | CB_D7 |  | A11 |  |  |  | C11 |  | POST5 |  | B11 | phiss_t_clk | CB_D7 |  | A11 |  | POST5RTN |  | C11 |  | POST5 |  | B11 | phiss_t_clk | CB_D7 |  | A11 |  | POST5RTN |  | C11 |  | POST5 |  | B11 | phiss_t_clk | CB_D7 |  | A11 |  | POST5RTN |  | C11 |  | POST5 |  | B11 | phiss_t_clk | CB_D23 |  | A11 |  | POST5RTN | C11 | phiss_t_clk | CB_D23 |  | B11 | phiss_b_clk | CB_D7 |  | A11 |  |  |  | C11 |  |  |  | B11 | phiss_b_clk | CB_D7 |  | A11 |  |  |  |  | C11 |  |  |  | B11 | phiss_b_clk | CB_D7 |  | A11 |  |  |  | C11 |  | POST5 |  | B11 | phiss_b_clk | CB_D7 |  | A11 |  | POST5RTN |  | C11 |  | POST5 |  | B11 | phiss_b_clk | CB_D7 |  | A11 |  | POST5RTN |  | C11 |  | POST5 |  | B11 | phiss_b_clk | CB_D7 |  | A11 |  | POST5RTN |  | C11 |  | POST5 |  | B11 | phiss_b_clk | CB_D7 |  | A11 |  | POST5RTN
C12 |  |  |  | B12 |  | [DGND] |  | A12 |  | [AGND] |  |  | C12 |  |  |  | B12 |  | [DGND] |  | A12 |  | [AGND] |  | C12 |  | AGND |  | B12 |  | [DGND] |  | A12 |  | [AGND] |  | C12 |  | AGND |  | B12 |  | [DGND] |  | A12 |  | [AGND] |  | C12 |  | AGND |  | B12 |  | [DGND] |  | A12 |  | [AGND] |  | C12 |  | AGND |  | B12 |  | [DGND] |  | A12 |  | [AGND] | C12 |  | DGND |  | B12 |  | [DGND] |  | A12 |  | [AGND] |  | C12 |  |  |  | B12 |  | [DGND] |  | A12 |  | [AGND] |  |  | C12 |  |  |  | B12 |  | [DGND] |  | A12 |  | [AGND] |  | C12 |  | AGND |  | B12 |  | [DGND] |  | A12 |  | [AGND] |  | C12 |  | AGND |  | B12 |  | [DGND] |  | A12 |  | [AGND] |  | C12 |  | AGND |  | B12 |  | [DGND] |  | A12 |  | [AGND] |  | C12 |  | AGND |  | B12 |  | [DGND] |  | A12 |  | [AGND]
C13 |  |  |  | B13 |  | [+5V] |  | A13 |  |  |  |  | C13 |  |  |  | B13 |  | [+5V] |  | A13 |  |  |  | C13 |  | POST6 |  | B13 |  | [+5V] |  | A13 |  | POST6RTN |  | C13 |  | POST6 |  | B13 |  | [+5V] |  | A13 |  | POST6RTN |  | C13 |  | POST6 |  | B13 |  | [+5V] |  | A13 |  | POST6RTN |  | C13 |  | POST6 |  | B13 |  | [+5V] |  | A13 |  | POST6RTN | C13 |  | +5V |  | B13 |  | [+5V] |  | A13 |  |  |  | C13 |  |  |  | B13 |  | [+5V] |  | A13 |  |  |  |  | C13 |  |  |  | B13 |  | [+5V] |  | A13 |  |  |  | C13 |  | POST6 |  | B13 |  | [+5V] |  | A13 |  | POST6RTN |  | C13 |  | POST6 |  | B13 |  | [+5V] |  | A13 |  | POST6RTN |  | C13 |  | POST6 |  | B13 |  | [+5V] |  | A13 |  | POST6RTN |  | C13 |  | POST6 |  | B13 |  | [+5V] |  | A13 |  | POST6RTN
C14 |  |  |  | B14 | vddcl_t_clk | CB_D8 |  | A14 |  |  |  |  | C14 |  |  |  | B14 | vddcl_t_clk | CB_D8 |  | A14 |  |  |  | C14 |  | AGND |  | B14 | vddcl_t_clk | CB_D8 |  | A14 |  | AGND |  | C14 |  | AGND |  | B14 | vddcl_t_clk | CB_D8 |  | A14 |  | AGND |  | C14 |  | AGND |  | B14 | vddcl_t_clk | CB_D8 |  | A14 |  | AGND |  | C14 |  | AGND |  | B14 | vddcl_t_clk | CB_D24 |  | A14 |  | AGND | C14 | vddcl_t_clk | CB_D24 |  | B14 | vddcl_b_clk | CB_D8 |  | A14 |  |  |  | C14 |  |  |  | B14 | vddcl_b_clk | CB_D8 |  | A14 |  |  |  |  | C14 |  |  |  | B14 | vddcl_b_clk | CB_D8 |  | A14 |  |  |  | C14 |  | AGND |  | B14 | vddcl_b_clk | CB_D8 |  | A14 |  | AGND |  | C14 |  | AGND |  | B14 | vddcl_b_clk | CB_D8 |  | A14 |  | AGND |  | C14 |  | AGND |  | B14 | vddcl_b_clk | CB_D8 |  | A14 |  | AGND |  | C14 |  | AGND |  | B14 | vddcl_b_clk | CB_D8 |  | A14 |  | AGND
C15 |  |  |  | B15 | phis1_t_clk | CB_D9 |  | A15 |  |  |  |  | C15 |  |  |  | B15 | phis1_t_clk | CB_D9 |  | A15 |  |  |  | C15 |  | POST7 |  | B15 | phis1_t_clk | CB_D9 |  | A15 |  | POST7RTN |  | C15 |  | POST7 |  | B15 | phis1_t_clk | CB_D9 |  | A15 |  | POST7RTN |  | C15 |  | POST7 |  | B15 | phis1_t_clk | CB_D9 |  | A15 |  | POST7RTN |  | C15 |  | POST7 |  | B15 | phis1_t_clk | CB_D25 |  | A15 |  | POST7RTN | C15 | phis1_t_clk | CB_D25 |  | B15 | phis1_b_clk | CB_D9 |  | A15 |  |  |  | C15 |  |  |  | B15 | phis1_b_clk | CB_D9 |  | A15 |  |  |  |  | C15 |  |  |  | B15 | phis1_b_clk | CB_D9 |  | A15 |  |  |  | C15 |  | POST7 |  | B15 | phis1_b_clk | CB_D9 |  | A15 |  | POST7RTN |  | C15 |  | POST7 |  | B15 | phis1_b_clk | CB_D9 |  | A15 |  | POST7RTN |  | C15 |  | POST7 |  | B15 | phis1_b_clk | CB_D9 |  | A15 |  | POST7RTN |  | C15 |  | POST7 |  | B15 | phis1_b_clk | CB_D9 |  | A15 |  | POST7RTN
C16 |  |  |  | B16 | phis2_t_clk | CB_D10 |  | A16 |  |  |  |  | C16 |  |  |  | B16 | phis2_t_clk | CB_D10 |  | A16 |  |  |  | C16 |  | AGND |  | B16 | phis2_t_clk | CB_D10 |  | A16 |  | AGND |  | C16 |  | AGND |  | B16 | phis2_t_clk | CB_D10 |  | A16 |  | AGND |  | C16 |  | AGND |  | B16 | phis2_t_clk | CB_D10 |  | A16 |  | AGND |  | C16 |  | AGND |  | B16 | phis2_t_clk | CB_D26 |  | A16 |  | AGND | C16 | phis2_t_clk | CB_D26 |  | B16 | phis2_b_clk | CB_D10 |  | A16 |  |  |  | C16 |  |  |  | B16 | phis2_b_clk | CB_D10 |  | A16 |  |  |  |  | C16 |  |  |  | B16 | phis2_b_clk | CB_D10 |  | A16 |  |  |  | C16 |  | AGND |  | B16 | phis2_b_clk | CB_D10 |  | A16 |  | AGND |  | C16 |  | AGND |  | B16 | phis2_b_clk | CB_D10 |  | A16 |  | AGND |  | C16 |  | AGND |  | B16 | phis2_b_clk | CB_D10 |  | A16 |  | AGND |  | C16 |  | AGND |  | B16 | phis2_b_clk | CB_D10 |  | A16 |  | AGND
C17 |  |  |  | B17 | phiroe_t_clk | CB_D11 |  | A17 |  |  |  |  | C17 |  |  |  | B17 | phiroe_t_clk | CB_D11 |  | A17 |  |  |  | C17 |  | POST8 |  | B17 | phiroe_t_clk | CB_D11 |  | A17 |  | POST8RTN |  | C17 |  | POST8 |  | B17 | phiroe_t_clk | CB_D11 |  | A17 |  | POST8RTN |  | C17 |  | POST8 |  | B17 | phiroe_t_clk | CB_D11 |  | A17 |  | POST8RTN |  | C17 |  | POST8 |  | B17 | phiroe_t_clk | CB_D27 |  | A17 |  | POST8RTN | C17 | phiroe_t_clk | CB_D27 |  | B17 | phiroe_b_clk | CB_D11 |  | A17 |  |  |  | C17 |  |  |  | B17 | phiroe_b_clk | CB_D11 |  | A17 |  |  |  |  | C17 |  |  |  | B17 | phiroe_b_clk | CB_D11 |  | A17 |  |  |  | C17 |  | POST8 |  | B17 | phiroe_b_clk | CB_D11 |  | A17 |  | POST8RTN |  | C17 |  | POST8 |  | B17 | phiroe_b_clk | CB_D11 |  | A17 |  | POST8RTN |  | C17 |  | POST8 |  | B17 | phiroe_b_clk | CB_D11 |  | A17 |  | POST8RTN |  | C17 |  | POST8 |  | B17 | phiroe_b_clk | CB_D11 |  | A17 |  | POST8RTN
C18 |  |  |  | B18 | phides_t_clk | CB_D12 |  | A18 |  |  |  |  | C18 |  |  |  | B18 | phides_t_clk | CB_D12 |  | A18 |  |  |  | C18 |  | AGND |  | B18 | phides_t_clk | CB_D12 |  | A18 |  | AGND |  | C18 |  | AGND |  | B18 | phides_t_clk | CB_D12 |  | A18 |  | AGND |  | C18 |  | AGND |  | B18 | phides_t_clk | CB_D12 |  | A18 |  | AGND |  | C18 |  | AGND |  | B18 | phides_t_clk | CB_D28 |  | A18 |  | AGND | C18 | phides_t_clk | CB_D28 |  | B18 | phides_b_clk | CB_D12 |  | A18 |  |  |  | C18 |  |  |  | B18 | phides_b_clk | CB_D12 |  | A18 |  |  |  |  | C18 |  |  |  | B18 | phides_b_clk | CB_D12 |  | A18 |  |  |  | C18 |  | AGND |  | B18 | phides_b_clk | CB_D12 |  | A18 |  | AGND |  | C18 |  | AGND |  | B18 | phides_b_clk | CB_D12 |  | A18 |  | AGND |  | C18 |  | AGND |  | B18 | phides_b_clk | CB_D12 |  | A18 |  | AGND |  | C18 |  | AGND |  | B18 | phides_b_clk | CB_D12 |  | A18 |  | AGND
C19 |  |  |  | B19 | phifs_t_clk | CB_D13 |  | A19 |  |  |  |  | C19 |  |  |  | B19 | phifs_t_clk | CB_D13 |  | A19 |  |  |  | C19 |  |  |  | B19 | phifs_t_clk | CB_D13 |  | A19 |  |  |  | C19 |  |  |  | B19 | phifs_t_clk | CB_D13 |  | A19 |  |  |  | C19 |  |  |  | B19 | phifs_t_clk | CB_D13 |  | A19 |  |  |  | C19 |  |  |  | B19 | phifs_t_clk | CB_D29 |  | A19 |  |  | C19 | phifs_t_clk | CB_D29 |  | B19 | phifs_b_clk | CB_D13 |  | A19 |  |  |  | C19 |  |  |  | B19 | phifs_b_clk | CB_D13 |  | A19 |  |  |  |  | C19 |  |  |  | B19 | phifs_b_clk | CB_D13 |  | A19 |  |  |  | C19 |  |  |  | B19 | phifs_b_clk | CB_D13 |  | A19 |  |  |  | C19 |  |  |  | B19 | phifs_b_clk | CB_D13 |  | A19 |  |  |  | C19 |  |  |  | B19 | phifs_b_clk | CB_D13 |  | A19 |  |  |  | C19 |  |  |  | B19 | phifs_b_clk | CB_D13 |  | A19
C20 |  |  |  | B20 | phif1_t_clk | CB_D14 |  | A20 |  |  |  |  | C20 |  |  |  | B20 | phif1_t_clk | CB_D14 |  | A20 |  |  |  | C20 |  |  |  | B20 | phif1_t_clk | CB_D14 |  | A20 |  |  |  | C20 |  |  |  | B20 | phif1_t_clk | CB_D14 |  | A20 |  |  |  | C20 |  |  |  | B20 | phif1_t_clk | CB_D14 |  | A20 |  |  |  | C20 |  |  |  | B20 | phif1_t_clk | CB_D30 |  | A20 |  |  | C20 | phif1_t_clk | CB_D30 |  | B20 | phif1_b_clk | CB_D14 |  | A20 |  |  |  | C20 |  |  |  | B20 | phif1_b_clk | CB_D14 |  | A20 |  |  |  |  | C20 |  |  |  | B20 | phif1_b_clk | CB_D14 |  | A20 |  |  |  | C20 |  |  |  | B20 | phif1_b_clk | CB_D14 |  | A20 |  |  |  | C20 |  |  |  | B20 | phif1_b_clk | CB_D14 |  | A20 |  |  |  | C20 |  |  |  | B20 | phif1_b_clk | CB_D14 |  | A20 |  |  |  | C20 |  |  |  | B20 | phif1_b_clk | CB_D14 |  | A20
C21 |  |  |  | B21 | phif2_t_cl;k | CB_D15 |  | A21 |  |  |  |  | C21 |  |  |  | B21 | phif2_t_cl;k | CB_D15 |  | A21 |  |  |  | C21 |  |  |  | B21 | phif2_t_cl;k | CB_D15 |  | A21 |  |  |  | C21 |  |  |  | B21 | phif2_t_cl;k | CB_D15 |  | A21 |  |  |  | C21 |  |  |  | B21 | phif2_t_cl;k | CB_D15 |  | A21 |  |  |  | C21 |  |  |  | B21 | phif2_t_cl;k | CB_D31 |  | A21 |  |  | C21 | phif2_t_cl;k | CB_D31 |  | B21 | phif2_b_clk | CB_D15 |  | A21 |  |  |  | C21 |  |  |  | B21 | phif2_b_clk | CB_D15 |  | A21 |  |  |  |  | C21 |  |  |  | B21 | phif2_b_clk | CB_D15 |  | A21 |  |  |  | C21 |  |  |  | B21 | phif2_b_clk | CB_D15 |  | A21 |  |  |  | C21 |  |  |  | B21 | phif2_b_clk | CB_D15 |  | A21 |  |  |  | C21 |  |  |  | B21 | phif2_b_clk | CB_D15 |  | A21 |  |  |  | C21 |  |  |  | B21 | phif2_b_clk | CB_D15 |  | A21
C22 |  |  |  | B22 |  | [DGND] |  | A22 |  | [AGND] |  |  | C22 |  |  |  | B22 |  | [DGND] |  | A22 |  | [AGND] |  | C22 |  |  |  | B22 |  | [DGND] |  | A22 |  | [AGND] |  | C22 |  |  |  | B22 |  | [DGND] |  | A22 |  | [AGND] |  | C22 |  |  |  | B22 |  | [DGND] |  | A22 |  | [AGND] |  | C22 |  |  |  | B22 |  | [DGND] |  | A22 |  | [AGND] | C22 |  | DGND |  | B22 |  | [DGND] |  | A22 |  | [AGND] |  | C22 |  |  |  | B22 |  | [DGND] |  | A22 |  | [AGND] |  |  | C22 |  |  |  | B22 |  | [DGND] |  | A22 |  | [AGND] |  | C22 |  |  |  | B22 |  | [DGND] |  | A22 |  | [AGND] |  | C22 |  |  |  | B22 |  | [DGND] |  | A22 |  | [AGND] |  | C22 |  |  |  | B22 |  | [DGND] |  | A22 |  | [AGND] |  | C22 |  |  |  | B22 |  | [DGND] |  | A22 |  | [AGND]
C23 |  |  |  | B23 |  | CB_CW0 |  | A23 |  |  |  |  | C23 |  |  |  | B23 |  | CB_CW0 |  | A23 |  |  |  | C23 |  |  |  | B23 |  | CB_CW0 |  | A23 |  |  |  | C23 |  |  |  | B23 |  | CB_CW0 |  | A23 |  |  |  | C23 |  |  |  | B23 |  | CB_CW0 |  | A23 |  |  |  | C23 |  |  |  | B23 |  | CB_CW8 |  | A23 |  |  | C23 |  | CB_CW8 |  | B23 |  | CB_CW0 |  | A23 |  |  |  | C23 |  |  |  | B23 |  | CB_CW0 |  | A23 |  |  |  |  | C23 |  |  |  | B23 |  | CB_CW0 |  | A23 |  |  |  | C23 |  |  |  | B23 |  | CB_CW0 |  | A23 |  |  |  | C23 |  |  |  | B23 |  | CB_CW0 |  | A23 |  |  |  | C23 |  |  |  | B23 |  | CB_CW0 |  | A23 |  |  |  | C23 |  |  |  | B23 |  | CB_CW0 |  | A23
C24 |  |  |  | B24 |  | CB_CW1 |  | A24 |  |  |  |  | C24 |  |  |  | B24 |  | CB_CW1 |  | A24 |  |  |  | C24 |  |  |  | B24 |  | CB_CW1 |  | A24 |  |  |  | C24 |  |  |  | B24 |  | CB_CW1 |  | A24 |  |  |  | C24 |  |  |  | B24 |  | CB_CW1 |  | A24 |  |  |  | C24 |  |  |  | B24 |  | CB_CW9 |  | A24 |  |  | C24 |  | CB_CW9 |  | B24 |  | CB_CW1 |  | A24 |  |  |  | C24 |  |  |  | B24 |  | CB_CW1 |  | A24 |  |  |  |  | C24 |  |  |  | B24 |  | CB_CW1 |  | A24 |  |  |  | C24 |  |  |  | B24 |  | CB_CW1 |  | A24 |  |  |  | C24 |  |  |  | B24 |  | CB_CW1 |  | A24 |  |  |  | C24 |  |  |  | B24 |  | CB_CW1 |  | A24 |  |  |  | C24 |  |  |  | B24 |  | CB_CW1 |  | A24
C25 |  |  |  | B25 |  | CB_CW2 |  | A25 |  |  |  |  | C25 |  |  |  | B25 |  | CB_CW2 |  | A25 |  |  |  | C25 |  |  |  | B25 |  | CB_CW2 |  | A25 |  |  |  | C25 |  |  |  | B25 |  | CB_CW2 |  | A25 |  |  |  | C25 |  |  |  | B25 |  | CB_CW2 |  | A25 |  |  |  | C25 |  |  |  | B25 |  | CB_CW10 |  | A25 |  |  | C25 |  | CB_CW10 |  | B25 |  | CB_CW2 |  | A25 |  |  |  | C25 |  |  |  | B25 |  | CB_CW2 |  | A25 |  |  |  |  | C25 |  |  |  | B25 |  | CB_CW2 |  | A25 |  |  |  | C25 |  |  |  | B25 |  | CB_CW2 |  | A25 |  |  |  | C25 |  |  |  | B25 |  | CB_CW2 |  | A25 |  |  |  | C25 |  |  |  | B25 |  | CB_CW2 |  | A25 |  |  |  | C25 |  |  |  | B25 |  | CB_CW2 |  | A25
C26 |  |  |  | B26 |  | CB_CW3 |  | A26 |  |  |  |  | C26 |  |  |  | B26 |  | CB_CW3 |  | A26 |  |  |  | C26 |  |  |  | B26 |  | CB_CW3 |  | A26 |  |  |  | C26 |  |  |  | B26 |  | CB_CW3 |  | A26 |  |  |  | C26 |  |  |  | B26 |  | CB_CW3 |  | A26 |  |  |  | C26 |  |  |  | B26 |  | CB_CW11 |  | A26 |  |  | C26 |  | CB_CW11 |  | B26 |  | CB_CW3 |  | A26 |  |  |  | C26 |  |  |  | B26 |  | CB_CW3 |  | A26 |  |  |  |  | C26 |  |  |  | B26 |  | CB_CW3 |  | A26 |  |  |  | C26 |  |  |  | B26 |  | CB_CW3 |  | A26 |  |  |  | C26 |  |  |  | B26 |  | CB_CW3 |  | A26 |  |  |  | C26 |  |  |  | B26 |  | CB_CW3 |  | A26 |  |  |  | C26 |  |  |  | B26 |  | CB_CW3 |  | A26
C27 |  |  |  | B27 |  | CB_CW4 |  | A27 |  |  |  |  | C27 |  |  |  | B27 |  | CB_CW4 |  | A27 |  |  |  | C27 |  |  |  | B27 |  | CB_CW4 |  | A27 |  |  |  | C27 |  |  |  | B27 |  | CB_CW4 |  | A27 |  |  |  | C27 |  |  |  | B27 |  | CB_CW4 |  | A27 |  |  |  | C27 |  |  |  | B27 |  | CB_CW12 |  | A27 |  |  | C27 |  | CB_CW12 |  | B27 |  | CB_CW4 |  | A27 |  |  |  | C27 |  |  |  | B27 |  | CB_CW4 |  | A27 |  |  |  |  | C27 |  |  |  | B27 |  | CB_CW4 |  | A27 |  |  |  | C27 |  |  |  | B27 |  | CB_CW4 |  | A27 |  |  |  | C27 |  |  |  | B27 |  | CB_CW4 |  | A27 |  |  |  | C27 |  |  |  | B27 |  | CB_CW4 |  | A27 |  |  |  | C27 |  |  |  | B27 |  | CB_CW4 |  | A27
C28 |  |  |  | B28 |  | CB_CW5 |  | A28 |  |  |  |  | C28 |  |  |  | B28 |  | CB_CW5 |  | A28 |  |  |  | C28 |  |  |  | B28 |  | CB_CW5 |  | A28 |  |  |  | C28 |  |  |  | B28 |  | CB_CW5 |  | A28 |  |  |  | C28 |  |  |  | B28 |  | CB_CW5 |  | A28 |  |  |  | C28 |  |  |  | B28 |  | CB_CW13 |  | A28 |  |  | C28 |  | CB_CW13 |  | B28 |  | CB_CW5 |  | A28 |  |  |  | C28 |  |  |  | B28 |  | CB_CW5 |  | A28 |  |  |  |  | C28 |  |  |  | B28 |  | CB_CW5 |  | A28 |  |  |  | C28 |  |  |  | B28 |  | CB_CW5 |  | A28 |  |  |  | C28 |  |  |  | B28 |  | CB_CW5 |  | A28 |  |  |  | C28 |  |  |  | B28 |  | CB_CW5 |  | A28 |  |  |  | C28 |  |  |  | B28 |  | CB_CW5 |  | A28
C29 |  |  |  | B29 |  | CB_CW6 |  | A29 |  |  |  |  | C29 |  |  |  | B29 |  | CB_CW6 |  | A29 |  |  |  | C29 |  |  |  | B29 |  | CB_CW6 |  | A29 |  |  |  | C29 |  |  |  | B29 |  | CB_CW6 |  | A29 |  |  |  | C29 |  |  |  | B29 |  | CB_CW6 |  | A29 |  |  |  | C29 |  |  |  | B29 |  | CB_CW14 |  | A29 |  |  | C29 |  | CB_CW14 |  | B29 |  | CB_CW6 |  | A29 |  |  |  | C29 |  |  |  | B29 |  | CB_CW6 |  | A29 |  |  |  |  | C29 |  |  |  | B29 |  | CB_CW6 |  | A29 |  |  |  | C29 |  |  |  | B29 |  | CB_CW6 |  | A29 |  |  |  | C29 |  |  |  | B29 |  | CB_CW6 |  | A29 |  |  |  | C29 |  |  |  | B29 |  | CB_CW6 |  | A29 |  |  |  | C29 |  |  |  | B29 |  | CB_CW6 |  | A29
C30 |  |  |  | B30 |  | CB_CW7 |  | A30 |  |  |  |  | C30 |  |  |  | B30 |  | CB_CW7 |  | A30 |  |  |  | C30 |  |  |  | B30 |  | CB_CW7 |  | A30 |  |  |  | C30 |  |  |  | B30 |  | CB_CW7 |  | A30 |  |  |  | C30 |  |  |  | B30 |  | CB_CW7 |  | A30 |  |  |  | C30 |  |  |  | B30 |  | CB_CW15 |  | A30 |  |  | C30 |  | CB_CW15 |  | B30 |  | CB_CW7 |  | A30 |  |  |  | C30 |  |  |  | B30 |  | CB_CW7 |  | A30 |  |  |  |  | C30 |  |  |  | B30 |  | CB_CW7 |  | A30 |  |  |  | C30 |  |  |  | B30 |  | CB_CW7 |  | A30 |  |  |  | C30 |  |  |  | B30 |  | CB_CW7 |  | A30 |  |  |  | C30 |  |  |  | B30 |  | CB_CW7 |  | A30 |  |  |  | C30 |  |  |  | B30 |  | CB_CW7 |  | A30
C31 |  |  |  | B31 |  | [DGND] |  | A31 |  | [AGND] |  |  | C31 |  |  |  | B31 |  | [DGND] |  | A31 |  | [AGND] |  | C31 |  |  |  | B31 |  | [DGND] |  | A31 |  | [AGND] |  | C31 |  |  |  | B31 |  | [DGND] |  | A31 |  | [AGND] |  | C31 |  |  |  | B31 |  | [DGND] |  | A31 |  | [AGND] |  | C31 |  |  |  | B31 |  | [DGND] |  | A31 |  | [AGND] | C31 |  | DGND |  | B31 |  | [DGND] |  | A31 |  | [AGND] |  | C31 |  |  |  | B31 |  | [DGND] |  | A31 |  | [AGND] |  |  | C31 |  |  |  | B31 |  | [DGND] |  | A31 |  | [AGND] |  | C31 |  |  |  | B31 |  | [DGND] |  | A31 |  | [AGND] |  | C31 |  |  |  | B31 |  | [DGND] |  | A31 |  | [AGND] |  | C31 |  |  |  | B31 |  | [DGND] |  | A31 |  | [AGND] |  | C31 |  |  |  | B31 |  | [DGND] |  | A31 |  | [AGND]
C32 |  |  |  | B32 |  | [+5V] |  | A32 |  | [-15V] |  |  | C32 |  |  |  | B32 |  | [+5V] |  | A32 |  | [-15V] |  | C32 |  |  |  | B32 |  | [+5V] |  | A32 |  | [-15V] |  | C32 |  |  |  | B32 |  | [+5V] |  | A32 |  | [-15V] |  | C32 |  |  |  | B32 |  | [+5V] |  | A32 |  | [-15V] |  | C32 |  |  |  | B32 |  | [+5V] |  | A32 |  | [-15V] | C32 |  | +5V |  | B32 |  | [+5V] |  | A32 |  | [-15V] |  | C32 |  |  |  | B32 |  | [+5V] |  | A32 |  | [-15V] |  |  | C32 |  |  |  | B32 |  | [+5V] |  | A32 |  | [-15V] |  | C32 |  |  |  | B32 |  | [+5V] |  | A32 |  | [-15V] |  | C32 |  |  |  | B32 |  | [+5V] |  | A32 |  | [-15V] |  | C32 |  |  |  | B32 |  | [+5V] |  | A32 |  | [-15V] |  | C32 |  |  |  | B32 |  | [+5V] |  | A32 |  | [-15V]

## 61 Cond Cable Assy
61 PIN |  | 61 PIN |  | 61 PIN |  | 61 PIN |  |  | 61 PIN |  | 61 PIN |  | 61 PIN |  | 61 PIN | 61 PIN |  | 61 PIN |  | 61 PIN |  | 61 PIN |  | IDC JH2 | BENDIX
MS3116-E24-61S | Belden 9221 Coax | MS3116-E24-61S | SIGNAL NAME | MS3116-E24-61S | Belden 9221 Coax | MS3116-E24-61S | SIGNAL NAME | coax1-16 critical | MS3116-E24-61S | Belden 9221 Coax | MS3116-E24-61S | Notes: | MS3116-E24-61S | Belden 9221 Coax | MS3116-E24-61S | MS3116-E24-61S | Belden 9221 Coax | MS3116-E24-61S | Notes: | MS3116-E24-61S | Belden 9221 Coax | MS3116-E24-61S | Notes: | Pin Num | 61PIN
B | SHIELD | B | VDETCOM | A | COND | A | VDDOUT_T |  | A | COND | A |  | f | COND | f | L | COND | L | (+ Outer Braided Shield | x | COND | x |  | 1 | Z | VP_T | COND
A | COND | A | VDDOUT_T | B | SHIELD | B | VDETCOM |  | B | SHIELD | B |  | g | SHIELD | g | M | SHIELD | M | connected on both ends) | y | SHIELD | y |  | 2 | Y | AGND | SHIELD
c | SHIELD | c | Q1OUT8- | C | COND | C | Q1OUT3+ | coax1 | C | COND | C |  | h | COND | h | N | COND | N |  | f | COND | f |  | 3 | u | VGGCL_T | SHIELD
b | COND | b | Q1OUT8+ | D | SHIELD | D | Q1OUT3- |  | D | SHIELD | D |  | i | SHIELD | i | m | SHIELD | m |  | g | SHIELD | g |  | 4 | a | VSSCM_T | COND
w | SHIELD | w | Q1OUT7- | E | COND | E | Q1OUT1+ | coax2 | E | COND | E |  | j | COND | j | P | COND | P |  | E | COND | E |  | 5 | GG | VRWOFF_T | SHIELD
v | COND | v | Q1OUT7+ | F | SHIELD | F | Q1OUT1- |  | F | SHIELD | F |  | k | SHIELD | k | R | SHIELD | R |  | F | SHIELD | F |  | 6 | FF | VRTSG_T | COND
e | SHIELD | e | Q1OUT6- | G | COND | G | PFS_T |  | G | COND | G |  | n | COND | n | J | COND | J |  | NN | COND | NN |  | 7 | t | VROWON_T | SHIELD
d | COND | d | Q1OUT6+ | H | SHIELD | H | PDES_T |  | H | SHIELD | H |  | p | SHIELD | p | K | SHIELD | K |  | MM | SHIELD | MM |  | 8 | s | VIREF_T | COND
JJ | SHIELD | JJ | Q1OUT5- | J | COND | J | PF2_T |  | J | COND | J |  | q | COND | q | j | COND | j |  | FF | COND | FF |  | 9 | MM | VRTSR_T | SHIELD
HH | COND | HH | Q1OUT5+ | K | SHIELD | K | PF1_T |  | K | SHIELD | K |  | r | SHIELD | r | k | SHIELD | k |  | GG | SHIELD | GG |  | 10 | NN | VTEND_T | COND
y | SHIELD | y | Q1OUT4- | L | COND | L |  | (+ Outer Braided Shield | L | COND | L | (+ Outer Braided Shield | s | COND | s | G | COND | G |  | s | COND | s |  | 11 | EE | Q1OUT1- | SHIELD
x | COND | x | Q1OUT4+ | M | SHIELD | M |  | connected on both ends ) | M | SHIELD | M | connected on both ends ) | t | SHIELD | t | H | SHIELD | H |  | t | SHIELD | t |  | 12 | DD | Q1OUT1+ | COND
D | SHIELD | D | Q1OUT3- | N | COND | N | VDDOUT_T |  | N | COND | N |  | v | COND | v | h | COND | h |  | W | COND | W |  | 13 | X | Q1OUT2- | SHIELD
C | COND | C | Q1OUT3+ | m | SHIELD | m |  |  | m | SHIELD | m |  | w | SHIELD | w | i | SHIELD | i |  | X | SHIELD | X |  | 14 | W | Q1OUT2+ | COND
g | SHIELD | g | Q1OUT2- | P | COND | P | Q1OUT8+ | coax3 | P | COND | P |  | x | COND | x | z | COND | z |  | HH | COND | HH |  | 15 | r | Q1OUT3- | SHIELD
f | COND | f | Q1OUT2+ | R | SHIELD | R | Q1OUT8- |  | R | SHIELD | R |  | y | SHIELD | y | AA | SHIELD | AA |  | JJ | SHIELD | JJ |  | 16 | q | Q1OUT3+ | COND
F | SHIELD | F | Q1OUT1- | S | COND | S | Q1OUT5+ | coax4 | S | COND | S |  | z | COND | z | n | COND | n |  | v | COND | v |  | 17 | V | Q1OUT4- | SHIELD
E | COND | E | Q1OUT1+ | T | SHIELD | T | Q1OUT5- |  | T | SHIELD | T |  | AA | SHIELD | AA | p | SHIELD | p |  | w | SHIELD | w |  | 18 | U | Q1OUT4+ | COND
LL | SHIELD | LL | VDDCL_T | U | COND | U | Q1OUT4+ | coax5 | U | COND | U |  | BB | COND | BB | BB | COND | BB |  | d | COND | d |  | 19 | T | Q1OUT5- | SHIELD
KK | COND | KK | VDDUC_T | V | SHIELD | V | Q1OUT4- |  | V | SHIELD | V |  | CC | SHIELD | CC | CC | SHIELD | CC |  | e | SHIELD | e |  | 20 | S | Q1OUT5+ | COND
AA | SHIELD | AA | PSS_T | W | COND | W | Q1OUT2+ | coax6 | W | COND | W |  | DD | COND | DD | S | COND | S |  | C | COND | C |  | 21 | p | Q1OUT6- | SHIELD
z | COND | z | PS1_T | X | SHIELD | X | Q1OUT2- |  | X | SHIELD | X |  | EE | SHIELD | EE | T | SHIELD | T |  | D | SHIELD | D |  | 22 | n | Q1OUT6+ | COND
i | SHIELD | i | PS2_T | Y | SHIELD | Y | AGND | *** Note change in | Y | SHIELD | Y | *** Note change in | FF | COND | FF | U | COND | U |  | b | COND | b |  | 23 | CC | Q1OUT7- | SHIELD
h | COND | h | PROE_T | Z | COND | Z | VP_T | pattern | Z | COND | Z | pattern | GG | SHIELD | GG | V | SHIELD | V |  | c | SHIELD | c |  | 24 | BB | Q1OUT7+ | COND
H | SHIELD | H | PDES_T | a | COND | a | VSSCM_T |  | a | COND | a |  | HH | COND | HH | q | COND | q |  | A | COND | A |  | 25 | R | Q1OUT8- | SHIELD
G | COND | G | PFS_T | u | SHIELD | u | VGGCL_T |  | u | SHIELD | u |  | JJ | SHIELD | JJ | r | SHIELD | r |  | B | SHIELD | B |  | 26 | P | Q1OUT8+ | COND
K | SHIELD | K | PF1_T | b | COND | b | Q1OUT8+ | coax7 | b | COND | b |  | KK | COND | KK | DD | COND | DD |  | a | COND | a |  | 27 | m |  | SHIELD
J | COND | J | PF2_T | c | SHIELD | c | Q1OUT8- |  | c | SHIELD | c |  | LL | SHIELD | LL | EE | SHIELD | EE |  | u | SHIELD | u |  | 28 | N | VDDOUT_T | COND
k | SHIELD | k | VNROW_T | d | COND | d | Q1OUT6+ | coax8 | d | COND | d |  | NN | COND | NN | KK | COND | KK |  | Y | SHIELD | Y | Note change in | 29 | PP
j | COND | j | VNCOL_T | e | SHIELD | e | Q1OUT6- |  | e | SHIELD | e |  | MM | SHIELD | MM | LL | SHIELD | LL |  | Z | COND | Z | pattern | 30 | NC | Q1TEND
M | SHIELD | M |  | f | COND | f | Q1OUT2+ | coax9 |  |  |  |  | PP | NC | PP |  |  |  |  | PP | NC | PP |  | 31 | NC
L | COND | L |  | g | SHIELD | g | Q1OUT2- | [16 empty] | 32 | NC
Z | COND | Z | VP_T | h | COND | h | PROE_T | [16 empty] | 33 | NC
Y | SHIELD | Y | AGND | i | SHIELD | i | PS2_T | [16 empty] | 34 | NC
u | SHIELD | u | VGGCL_T | j | COND | j | VNCOL_T | [16 empty] | Pin Num | 61PIN
a | COND | a | VSSCM_T | k | SHIELD | k | VNROW_T | [16 empty] | 1 | B | VDETCOM | SHIELD
GG | SHIELD | GG | VRWOFF_T | n | COND | n | Q1OUT6+ | coax10 | [15 empty] | 2 | A | VDDOUT_T | COND
FF | COND | FF | VRTSG_T | p | SHIELD | p | Q1OUT6- | [16 empty] | 3 | c | Q1OUT8- | SHIELD
t | SHIELD | t | VROWON_T | q | COND | q | Q1OUT3+ | coax11 | [15 empty] | 4 | b | Q1OUT8+ | COND
s | COND | s | VIREF_T | r | SHIELD | r | Q1OUT3- | [16 empty] | 5 | w | Q1OUT7- | SHIELD
MM | SHIELD | MM | VRTSR_T | s | COND | s | VIREF_T | [16 empty] | 6 | v | Q1OUT7+ | COND
NN | COND | NN | VTEND_T | t | SHIELD | t | VROWON_T | [16 empty] | 7 | e | Q1OUT6- | SHIELD
EE | SHIELD | EE | Q1OUT1- | v | COND | v | Q1OUT7+ | coax12 | [15 empty] | 8 | d | Q1OUT6+ | COND
DD | COND | DD | Q1OUT1+ | w | SHIELD | w | Q1OUT7- | [16 empty] | 9 | JJ | Q1OUT5- | SHIELD
X | SHIELD | X | Q1OUT2- | x | COND | x | Q1OUT4+ | coax13 | [15 empty] | 10 | HH | Q1OUT5+ | COND
W | COND | W | Q1OUT2+ | y | SHIELD | y | Q1OUT4- | [16 empty] | 11 | y | Q1OUT4- | SHIELD
r | SHIELD | r | Q1OUT3- | z | COND | z | PS1_T | [16 empty] | 12 | x | Q1OUT4+ | COND
q | COND | q | Q1OUT3+ | AA | SHIELD | AA | PSS_T | [16 empty] | 13 | D | Q1OUT3- | SHIELD
V | SHIELD | V | Q1OUT4- | BB | COND | BB | Q1OUT7+ | coax14 | [15 empty] | 14 | C | Q1OUT3+ | COND
U | COND | U | Q1OUT4+ | CC | SHIELD | CC | Q1OUT7- | [16 empty] | 15 | g | Q1OUT2- | SHIELD
T | SHIELD | T | Q1OUT5- | DD | COND | DD | Q1OUT1+ | coax15 | [15 empty] | 16 | f | Q1OUT2+ | COND
S | COND | S | Q1OUT5+ | EE | SHIELD | EE | Q1OUT1- | [16 empty] | 17 | F | Q1OUT1- | SHIELD
p | SHIELD | p | Q1OUT6- | FF | COND | FF | VRTSG_T | [16 empty] | 18 | E | Q1OUT1+ | COND
n | COND | n | Q1OUT6+ | GG | SHIELD | GG | VRWOFF_T | [16 empty] | 19 | LL | VDDCL_T | SHIELD
CC | SHIELD | CC | Q1OUT7- | HH | COND | HH | Q1OUT5+ | coax16 | [15 empty] | 20 | KK | VDDUC_T | COND
BB | COND | BB | Q1OUT7+ | JJ | SHIELD | JJ | Q1OUT5- | [16 empty] | 21 | AA | PSS_T | SHIELD
R | SHIELD | R | Q1OUT8- | KK | COND | KK | VDDUC_T | [16 empty] | 22 | z | PS1_T | COND
P | COND | P | Q1OUT8+ | LL | SHIELD | LL | VDDCL_T | [16 empty] | 23 | i | PS2_T | SHIELD
m | SHIELD | m |  | NN | COND | NN | VTEND_T | [16 empty] | 24 | h | PROE_T | COND
N | COND | N | VDDOUT_T | MM | SHIELD | MM | VRTSR_T | [16 empty] | 25 | H | PDES_T | SHIELD
PP | NC | PP |  | PP | NC | PP | [17 empty] | 26 | G | PFS_T | COND
NC |  | NC | Q1TEND | NC |  | NC | Q1TEND | [16 empty] | 27 | K | PF1_T | SHIELD
[24 empty] | 28 | J | PF2_T | COND
[24 empty] | 29 | k | VNROW_T | SHIELD
[24 empty] | 30 | j | VNCOL_T | COND
[24 empty] | 31 | M |  | SHIELD
[24 empty] | 32 | L |  | COND
[24 empty] | 33 | NC
[24 empty] | 34 | NC
Cable layout configuration 1
Cable layout configuration 2
Build from bottom of connector

## 61 PIN-34IDC
IDC JH2 | BENDIX
Pin Num | 61PIN | Signal | Belden 9221
1 | Z | VP_T | COND
2 | Y | AGND | SHIELD
3 | u | VGGCL_T | SHIELD
4 | a | VSSCM_T | COND
5 | GG | VRWOFF_T | SHIELD
6 | FF | VRTSG_T | COND
7 | t | VROWON_T | SHIELD
8 | s | VIREF_T | COND
9 | MM | VRTSR_T | SHIELD
10 | NN | VTEND_T | COND
11 | EE | Q1OUT1- | SHIELD
12 | DD | Q1OUT1+ | COND
13 | X | Q1OUT2- | SHIELD
14 | W | Q1OUT2+ | COND
15 | r | Q1OUT3- | SHIELD
16 | q | Q1OUT3+ | COND
17 | V | Q1OUT4- | SHIELD
18 | U | Q1OUT4+ | COND
19 | T | Q1OUT5- | SHIELD
20 | S | Q1OUT5+ | COND
21 | p | Q1OUT6- | SHIELD
22 | n | Q1OUT6+ | COND
23 | CC | Q1OUT7- | SHIELD
24 | BB | Q1OUT7+ | COND
25 | R | Q1OUT8- | SHIELD
26 | P | Q1OUT8+ | COND
27 | m |  | SHIELD
28 | N | VDDOUT_T | COND
29 | PP
30 | NC | Q1TEND
31 | NC
32 | NC
33 | NC
34 | NC
IDC JH1
Pin Num | 61PIN
1 | B | VDETCOM | SHIELD
2 | A | VDDOUT_T | COND
3 | c | Q1OUT8- | SHIELD
4 | b | Q1OUT8+ | COND
5 | w | Q1OUT7- | SHIELD
6 | v | Q1OUT7+ | COND
7 | e | Q1OUT6- | SHIELD
8 | d | Q1OUT6+ | COND
9 | JJ | Q1OUT5- | SHIELD
10 | HH | Q1OUT5+ | COND
11 | y | Q1OUT4- | SHIELD
12 | x | Q1OUT4+ | COND
13 | D | Q1OUT3- | SHIELD
14 | C | Q1OUT3+ | COND
15 | g | Q1OUT2- | SHIELD
16 | f | Q1OUT2+ | COND
17 | F | Q1OUT1- | SHIELD
18 | E | Q1OUT1+ | COND
19 | LL | VDDCL_T | SHIELD
20 | KK | VDDUC_T | COND
21 | AA | PSS_T | SHIELD
22 | z | PS1_T | COND
23 | i | PS2_T | SHIELD
24 | h | PROE_T | COND
25 | H | PDES_T | SHIELD
26 | G | PFS_T | COND
27 | K | PF1_T | SHIELD
28 | J | PF2_T | COND
29 | k | VNROW_T | SHIELD
30 | j | VNCOL_T | COND
31 | M |  | SHIELD
32 | L |  | COND
33 | NC
34 | NC

## Sheet2
EXTRAPOLATION FOR 3RD IXTHOS DSP BOARD
| # pixels | outputs | tmin | T per address | T per pixel | #DSPs | X speed | hz
256X256 NSFCAM | 65536 | 4 | 0.076 | 4.6387e-06 | 1.1597e-06 | 1 | 1 | 13.1579
using the NSFCAM T per pixel rate=
512x512 Ixthos bds=2 | 262144 | 32 | 0.152 | 1.8555e-05 |  | 1 | 2 | 6.5789
512x512 Ixthos bds=3 | 262144 | 32 | 0.035765 | 4.3658e-06 |  | 4.25 | 2 | 27.9605

## T_READ4 (2)
ARRAY READOUT ESTIMATES
T_frame = (N_num_pixels / N_num_outputs) * ((T_pixel_address+ T_preamp_settle) + ((N_num_outputs / N_num_ADCs) * ( (T_mux_switch*K_mux) + T_ADC_convert + T_ship_data)))
T_2quad=(N_num_pixels_2quad/N_num_outputs_2quad)*((T_pixel_address+T_preamp_settle)+((N_num_outputs_2quad/N_num_adcs_2quad)*((T_mux_switch*K_mux_2quad)+T_ADC_convert+T_ship_data_2quad)))
T_quadrant = (N_num_pixels_quad / N_num_outputs_quad) * ((T_pixel_address+ T_preamp_settle) + ((N_num_outputs_quad / N_num_adcs_quad) * ( (T_mux_switch*K_mux_quad) + T_ADC_convert +T_ship_data_quad)))
| T_frame | 0.062259 | sec |  | T_2quad | 0.098304 | sec |  | T_quadrant | 0.055706 | sec
| R_frame_rate | 16.0619 | Hz |  | R_2quad_rate | 10.1725 | Hz |  | R_quad_rate | 17.9515 | Hz
| FRAME VARIABLES |  |  |  | HALF FRAME VARIABLES |  |  |  | QUADRANT VARIABLES
| N_num_pixels | 1048576 | pixels |  | N_num_pixels_2quad | 524288 | pixels |  | N_num_pixels_quad | 262144 | pixels
| N_num_outputs | 32 |  |  | N_num_outputs_2quad | 16 |  |  | N_num_outputs_quad | 8
SIGNAL CHAIN VARIABLES
| T_pixel_address | 1e-06 | sec
| T_preamp_settle | 0 | sec
| T_mux_switch | 5e-08 | sec
| K_mux | 0 |  |  | K_mux_2quad | 2 |  |  | K_mux_quad | 0
DATA PROCESS VARIABLES
| N_num_ADCs | 32 |  |  | N_num_adcs_2quad | 8 |  |  | N_num_adcs_quad | 8
| T_ADC_convert | 5e-07 | sec
| T_ship_data | 4e-07 | sec |  | T_ship_data_2quad | 4e-07 | sec |  | T_ship_data_quad | 2e-07 | sec
| N_ship_channels | 2 |  |  | N_ship_channels_2quad | 1 |  |  | N_ship_channels_quad | 1
| T_ship_one | 2.5e-08 | sec
FRAME DSP CALCULATIONS
DSP | N_instr_per_addr | N_num_DSPs | N_DSP_clk | T_inst | pixels_per_DSP | instr_per_pixel
SHARC | 76 | 4 | 40000000 | 2.5e-08 | 8 | 9.5
C40* | 47.5 | 4 | 50000000 | 4e-08 | 8 | 5.9375 |  |  | *=2 cycles/inst
96002* | 38 | 4 | 40000000 | 5e-08 | 8 | 4.75 |  |  | *=2 cycles/inst
56301 | 152 | 4 | 80000000 | 1.25e-08 | 8 | 19
QUAD DSP CALCULATIONS
DSP | N_instr_per_addr | N_num_DSPs | N_DSP_clk | T_inst | pixels_per_DSP | instr_per_pixel
SHARC | 68 | 4 | 40000000 | 2.5e-08 | 2 | 34
C40* | 42.5 | 4 | 50000000 | 4e-08 | 2 | 21.25 |  |  | *=2 cycles/inst
96002* | 34 | 4 | 40000000 | 5e-08 | 2 | 17 |  |  | *=2 cycles/inst
56301 | 136 | 4 | 80000000 | 1.25e-08 | 2 | 68
EXTRAPOLATION FOR 3RD IXTHOS DSP BOARD
|  | # pixels | outputs | tmin | T per address | T per pixel | #DSPs | speed factor | hz
| 256X256 NSFCAM | 65536 | 4 | 0.076 | 4.6387e-06 | 1.1597e-06 | 1 | 1 | 13.1579
| 512x512 Ixthos bds=2 | 262144 | 32 | 0.152 | 1.8555e-05 |  | 1 | 2 | 6.5789
| 512x512 Ixthos bds=3 | 262144 | 32 | 0.035765 | 4.3658e-06 |  | 4.25 | 2 | 27.9605

## DSPs
DSP | clk | data/addr | MIPS | MOPS | MFLOPS | Tinstruct | Registers | Internal RAM | buses-ext | Com Ports | Com speed | Power | IDD
Motorola
DSP96002 | 33/40Mhz | 32/32 bits | 20 | 200 | 60 | 50nsec | 10 | 1K+512Words | 2
*DSP96002 | *66Mhz | 32/32 bits | *33 | *300 | *90 | 30nsec | 10 | 1K+512Words | 2
*DSP56301 | *66Mhz | 24/24 bits | 80 |  |  | 12.5nsec |  | 1K+512Words | 1+1PCI
Analog Devices
ADSP21060 | 33/40Mhz | 40/32 bits | 40 |  | 80 | 25nsec | 16 | 80K pgmWords | 1^ | 6 | 40MByte/s | 4 | 0.8
ADSP21062 | 33/40Mhz | 40/32 bits | 40 |  | 80 | 25nsec | 16 | 40K pgmWords | 1^ | 6 | 40MByte/s | 4 | 0.8
Texas Inst
TMS320C40 | 40/50Mhz | 32/31 bits | 25 | 275 | 50 | 40nsec | 12 | 2K Words | 2 | 6 | 20MByte/s
TMS320C44 | 50Mhz | 32/23 bits | 25 | 275 | 50 | 40nsec | 12 | 2K Words | 2 | 4
TMS320C80 |  | 64 |  | 2GOPS |  |  |  |  | 1

## SPX_COST
A = Outside Fabrication, B = Purchased Item, C = Simple, D = Moderately Complex, E = Complex, F = Major component
SPEX ELECTRONICS COST ESTIMATE |  |  |  | Item Code | Number of Parts | Number of Spares | Item Cost | Purchase Price | Lead Time Weeks | Spares Price | Design Hours | Test/Int Hours | Tech Hours | [6 empty] | Actual Cost | PO number | Vendor & part number
CRYOSTAT ASSEMBLIES
1. Array Electronics
| 1.1 Array Fanout Assembly |  |  | D | [7 empty] | 2
|  | 1.1.1 Printed Circuit Board |  | B | 10 |  | 240 | 2400 |  | 0 | 40 |  | 60 | [6 empty] | -2400
|  | 1.1.2 Hermetic Connectors |  | B | 4 | 4 | 80 | 320 |  | 320 | 8 |  | 36 | [6 empty] | -640
|  | 1.1.3 Shielding |  | C | 2 |  | 80 | 160 |  | 0 | 8 |  | 24
|  | 1.1.4 Yamaichi socket |  |  | 2 | 2 | 100 | 200 |  | 200 | 8 |  | 12 | [6 empty] | 273.6 | P722004 | Yamaichi socket
|  |  | Subtotal |  | 18 | 6 |  | 3080 |  | 520 | 64 | 2 | 132 | 9804 | [5 empty] | -2766.4
| 1.2 Configuration Board |  |  | D
|  | 1.3.1 Printed Circuit Board |  | B | 2 | 1 | 400 | 1200 |  | 0 | [9 empty] | -1200
|  | 1.3.2 Electronic Components |  | B | 2 | 1 | 50 | 100 |  | 50 | [9 empty] | -150
|  |  | Subtotal |  | 4 | 2 |  | 1300 |  | 50 | 0 | 0 | 0 | [6 empty] | -1350
| 1.3 Array Cryostat Wiring |  |  | D
|  | 1.2.1 Cryo Coax wire |  | B | 100ft |  |  | 450 |  |  | 8 |  | 40 | [6 empty] | 1297 | P722375 | WL GORE/275ft custom cable
|  | 1.2.2 Constantan wire |  | B | 100ft | [7 empty] | 24 | [6 empty] | 1430 | P722374 | MWS wire/ 2000ft manganin
|  | 1.2.3 Flexwire assy |  | A | 4 | 2 | 400 | 1600 |  | 800 | 40 |  | 40 | [6 empty] | 0
|  | 1.2.4 Feedthrough Assy |  | C | 4 |  | 100 | 400 |  | 0 | 8 |  | 8
|  |  | Subtotal |  | 8 | 2 |  | 2450 |  | 800 | 56 | 0 | 112 | 8514 | [5 empty] | 2727
2. Motor Assemblies
| 2.1 Stepper Motor Assembly |  |  | D | 8 | [6 empty] | 32
|  | 2.1.1 Stepper Motor |  |  | 8 | 2 | 400 | 3200 |  | 800 | 24 | 8 | 60 | [6 empty] | 513 | P722339 | Stepper #M2-2232-D&drive
|  | 2.1.1 Motor Wiring |  | C | 8 | [5 empty] | 8 |  | 16 | [6 empty] | 2129 | P817587 | 6 Stepper #M2-2232-D&drive
|  | 2.1.2. Cold Test support |  | D | [6 empty] | 16 | 2 | 24
|  | 2.1.3 Hermetic connectors |  | B | 0 | 0 | 100 | 0 |  | 0 | 8 |  | 16
|  |  | Subtotal |  | 24 | 2 |  | 3200 |  | 800 | 56 | 42 | 116 | [6 empty] | 2642
3. Position Sensing Assemblies
| 3.1 Absolute Encoder Assembly |  |  | D | [7 empty] | 20
|  | 3.1.1 Inductosyn |  | B | 0 | 0 | 2000 | 0 |  | 0 | 16 |  | 40
|  | 3.1.2 Wiring |  |  | 2 | [5 empty] | 16 |  | 40
|  | 3.1.3 Hermetic connector |  | B | 2 | 0 | 100 | 200 |  | 0 | 8 |  | 40
|  |  | Subtotal |  | 4 | 0 |  | 200 |  | 0 | 40 | 20 | 120 | [6 empty] | 0
| 3.2 Limit switch/detent Assemblies |  |  | D | [7 empty] | 16
|  | 3.2.1 Cryogenic sensors? |  | B | 8 | 2 | 525 | 4200 |  | 1050 | 40 |  | 40
|  | 3.2.2 Wiring |  | C | 8 | [5 empty] | 24 |  | 40
|  | 3.2.4 Signal Conditioning |  | C | 8 | 1 | 100 | 800 |  | 100 | 40 |  | 40
|  |  | Subtotal |  | 24 | 3 |  | 5000 |  | 1150 | 104 | 16 | 120 | [6 empty] | 0
320
CRYOSTAT MOUNTED ELECTRONICS ASSEMBLIES
4. Array readout electronics
| 4.1 Eight Channel Preamplifier |  |  | D | [7 empty] | 32
|  | 4.1.1 Printed Circuit Board |  | B | 6 | 1 | 400 | 2400 |  | 400 | 80 |  | 24
|  | 4.1.2 Component kit |  | B | 6 | 1 | 400 | 2400 |  | 400 | 8 |  | 24 | [6 empty] | -2800 |  | ESTIMATED
|  | 4.1.3 Fabrication |  | B |  |  | 300 | 0 |  | 0 | 8 |  | 120
|  |  | Subtotal |  | 12 | 2 |  | 4800 |  | 800 | 96 | 32 | 168 | 13496 | [5 empty] | -2800
| 4.2 Eight Channel A/D converter |  |  | D | [7 empty] | 24
|  | 4.2.1 Printed Circuit Board |  | B | 6 | 1 | 400 | 2400 |  | 400 | 40 |  | 16
|  | 4.2.2 Component kit |  | B | 6 | 1 | 400 | 2400 |  | 400 | 8 |  | 24 | [6 empty] | -2800 |  | ESIMATED
|  | 4.2.3 2Mhz 16 bit ADC conveters |  | B | 40 | 8 | 971.1 | 38844 |  | 7768.8 | 8 |  | 16 | [6 empty] | 68127 |  | Analogic, ADC4322
|  | 4.2.3 Fabrication |  | B |  |  | 300 | 0 |  | 0 | 8 |  | 60
|  |  | Subtotal |  | 52 | 10 |  | 43644 |  | 8568.8 | 64 | 24 | 116 | 57664.8 | [5 empty] | 65327
| 4.3 8/16 Channel Clock/Bias Driver |  |  | D | [7 empty] | 16
|  | 4.3.1 Printed Circuit Board |  | B | 6 | 0 | 400 | 2400 |  | 0 | 60 |  | 8
|  | 4.3.2 Component kit |  | B | 6 | 0 | 750 | 4500 |  | 0 | 8 |  | 30 | [6 empty] | -4500
|  | 4.3.3 Fabrication |  | B |  |  | 300 | 0 |  | 0 | 8 |  | 80
|  |  | Subtotal |  | 12 | 0 |  | 6900 |  | 0 | 76 | 16 | 118 | 12446 | [5 empty] | -4500
| 4.3.1 Clock/Bias Daughter board |  |  | D | [7 empty] | 16
|  | 4.3.1 Printed Circuit Board |  | B | 6 | 0 | 75 | 450 |  | 0 | 60 |  | 8
|  | 4.3.2 Component kit |  | B | 6 | 0 | 100 | 600 |  | 0 | 8 |  | 30 | [6 empty] | -600
|  | 4.3.3 Fabrication |  | B |  |  | 100 | 0 |  | 0 | 8 |  | 80
|  |  | Subtotal |  | 12 | 0 |  | 1050 |  | 0 | 76 | 16 | 118 | 6596 | [5 empty] | -600
| 4.4 Fibre Optic Interface (1Gbps/sec) |  |  | E | [7 empty] | 80
|  | 4.4.1 Printed Circuit Board |  | B | 6 | 2 | 600 | 3600 |  | 1200 | 60 |  | 24
|  | 4.4.2 Component kit |  | B | 6 | 2 | 1600 | 9600 |  | 3200 | 24 |  | 24 | [6 empty] | 16260 | P72318&? | Vitesse and Finisar
|  | 4.4.3 Fabrication |  | B |  |  | 300 | 0 |  | 0 | 8 |  | 80
|  |  | Subtotal |  | 12 | 4 |  | 13200 |  | 4400 | 92 | 80 | 128 | 23616 | [5 empty] | 16260
| 4.4.1 Fibre Optic Daughter Board |  |  | E | [7 empty] | 80
|  | 4.4.1 Printed Circuit Board |  | B | 6 | 0 | 75 | 450 |  | 0 | 60 |  | 24
|  | 4.4.2 Component kit |  | B | 6 | 0 | 100 | 600 |  | 0 | 24 |  | 24 | [6 empty] | -600
|  | 4.4.3 Fabrication |  | B |  |  | 100 | 0 |  | 0 | 8 |  | 80
|  |  | Subtotal |  | 24 | 4 |  | 1050 |  | 0 | 184 | 160 | 256 | 13082 | [5 empty] | -600
| 4.5 Chassis & backplane |  |  | E | [7 empty] | 20
|  | 4.5.1 Custom Eurocard backplane |  | B | 5 | 1 | 400 | 2000 |  | 400 | 40 |  | 64 | [6 empty] | -2800
|  | 4.5.2 VME J2 backplane |  | B | 1 |  | 1000 | 1000 |  | 0 | 16 |  | 24
|  | 4.5.3 Chassis |  | B | 1 |  | 3000 | 3000 |  | 0 | 8 |  | 16
|  |  | Subtotal |  | 7 | 1 |  | 6000 |  | 400 | 64 | 20 | 104 | 11288 | [5 empty] | -2800
4.6 Cables (attached to cryostat)
|  | 4.6.1 Array Cables |  | C | 5 | [5 empty] | 16 | 4 | 40
|  | 4.6.2 Array cable connectors |  | B | 8 | 2 | 80 | 640 |  | 160 |  |  | 8
|  | 4.6.3 Motor Cables |  | C | ? | [5 empty] | 16 | 4 | 24
|  | 4.6.4 Motor cable connectors |  | B | 4 | 2 | 80 | 320 |  | 160 |  |  | 8
|  | 4.6.5 Position sense cables |  | C | ? | [5 empty] | 16 | 4 | 24
|  | 4.6.6 Pos sense cable connectors |  | B | 4 | 2 | 80 | 320 |  | 160 |  |  | 8
|  |  | Subtotal |  | 21 | 6 |  | 1280 |  | 480 | 48 | 12 | 112 | 7024 | [5 empty] | 0
[8 empty] | 75824 | 0 | 14648.8 | 440 |  |  | 143852.8
ELECTRONICS POWER SUPPLIES
5. Array Power Supply |  |  |  | C | [7 empty] | 8
| 5.1.1 Chassis |  |  | B | 1 |  | 720 | 720 |  | 0 | 16 |  | 40
| 5.1.2 Serial control interface |  |  | B | 1 |  | 1500 | 1500 |  | 0 | 4 |  | 4
| 5.1.3 +5V 20amp supply |  |  | B | 1 | 1 | 1700 | 1700 |  | 1700 | 2 |  | 4
| 5.1.4 15V 12amp supply |  |  | B | 2 | 1 | 1700 | 3400 |  | 1700 | 2 |  | 4
| 5.1.5 Cabling/connectors |  |  | B | 1 |  | 100 | 100 |  | 0 | 4 |  | 40
|  |  | Subtotal |  | 6 | 2 |  | 7320 |  | 3400 | 28 | 8 | 92 | [6 empty] | 0
6. Motor Driver Power Supply |  |  |  | C
| 6.1.1 Stepper motor drive module |  |  | B | 7 | 2 | 500 | 3500 |  | 1000 | 32 | 16 | 32
| 6.1.1 Chassis |  |  | B | 1 |  | 720 | 720 |  | 0 | 8 | 16 | 24
? | 6.1.2 Serial control interface |  |  | B | 1 |  | 500 | 500 |  | 0 | 8 |  | 8
? | 6.1.3 +5V 20amp supply |  |  | B | 1 | 1 | 500 | 500 |  | 500 | 2 |  | 2
? | 6.1.4 15V 12amp supply |  |  | B | 1 | 1 | 500 | 500 |  | 500 | 2 |  | 2
? | 6.1.5 24V 8amp supply |  |  | B | 1 | 1 | 500 | 500 |  | 500 | 2 |  | 4
| 6.1.5 Cabling/connectors |  |  | B | 1 |  | 100 | 100 |  | 0 | 4 |  | 8
|  |  | Subtotal |  | 13 | 5 |  | 6320 |  | 2500 | 58 | 32 | 80 | [6 empty] | 0
7. AC Power Distribution/Control |  |  |  | C
| 7.1.1 Filtered AC distribution panel |  |  | B | 1 |  | 1700 | 1700 |  | 0 | 16 | 16 | 8
| 7.1.2 Serial Control interface |  |  | B | 1 |  | 1000 | 1000 |  | 0 | 8 |  | 16
|  |  | Subtotal |  | 2 | 0 |  | 2700 | 0 | 0 | 24 | 16 | 24 | [6 empty] | 0
110
INSTRUMENT CONTROL ELECTRONICS ASSEMBLIES
8. VME Control System
8.1 Communications and Processing
|  | 8.1.1 .0 Single Board Computer Specrogarph |  | B | 1 | 1 | 10000 | 10000 |  | 10000 | 8 | 16 | 16 | [6 empty] | 10238 | P990456-00030 | Themis Ultrasparc 2I Processor+accs
|  | 8.1.1.0.1 Graphics card/mouse/keyboard |  | B |  |  | 0 | 0 |  | 0 | 8 | 16 | 16
|  | 8.1.1.0.2 RAM |  | B |  |  | 0 | 0 |  | 0 | 8 | 16 | 16
|  | 8.1.1.0.3 Disk |  | B |  |  | 0 | 0 |  | 0 | 8 | 16 | 16 | [6 empty] | 915 | P722478 | TO IRTF????
|  | 8.1.1.0.4 Monitor |  | B |  |  | 0 | 0 |  | 0 | 8 | 16 | 16
|  | 8.1.1.1 Single Board Computer Imager |  | B | 1 | 1 | 0 | 0 |  | 0 | 8 | 16 | 16 | [6 empty] | 5427.25 | P722913 | Themis SPARC 5/64-32-110
|  | 8.1.1.1.1 Graphics card/mouse/keyboard |  | B |  |  | 0 | 0 |  | 0 | 8 | 16 | 16 | [6 empty] | 788 | ?P723239 | ZZYX/Integrix TGX120,etc
|  | 8.1.1.1.2 RAM |  | B |  |  | 0 | 0 |  | 0 | 8 | 16 | 16 | [6 empty] | 200 | P723154 | Centon CUN32/M Sparc LXE PLUS
|  | 8.1.1.1.3 Disk |  | B |  |  | 0 | 0 |  | 0 | 8 | 16 | 16
|  | 8.1.1.1.4 Monitor |  | B |  |  | 0 | 0 |  | 0 | 8 | 16 | 16 | [6 empty] | 699 | P723000 | USA flex/SONY 17PS
|  | 8.1.2 SCSI disk (>4 Gbyte) |  | B | 1 |  | 1500 | 1500 |  | 0 | 8 |  | 8
|  | 8.1.3 Chassis/power supply |  | B | 1 |  | 4000 | 4000 |  | 0 | 8 |  | 16
|  | 8.1.4 Rackmount keyboard |  | B |  |  | 400 | 0 |  | 0
|  | 8.1.5 Monitor |  | B |  |  | 400 | 0 |  | 0
|  | 8.1.6 Cables |  | B | 1 |  | 100 | 100 |  | 0 |  |  | 8 | [6 empty] | 231 | P722689 | cables
|  |  | Subtotal |  | 5 | 2 |  | 15600 |  | 10000 | 96 | 160 | 192 | [6 empty] | 18498.25
| 8.2 Vendor DSP board |  |  | B | 1 | [14 empty] | 24165 | PZ004852 | Ixthos DSP bd + memory
|  | 8.2.1 Multi DSP board |  | B | 2 | 1 | 8000 | 16000 |  | 8000 | 16 | 16 | 8 | [6 empty] | 49581.5 | P722502 | Ixthos/ qty 2 IXZ4444 DSP boards
|  | 8.2.2 Support software/cables/interface |  | B | 1 |  | 2000 | 2000 |  | 0 | [9 empty] | 2278 | P723288 | Ixthos
|  |  | Subtotal |  | 4 | 1 |  | 18000 |  | 8000 | 16 | 16 | 8 | [6 empty] | 51859.5
8.3 Motor Control TURNED INTO LITTLE DOG PC
|  | 8.3.0 486 PC |  | B | 0 | 0 | 0 | 0 |  | 0 | 24 | 16 | 40 | [6 empty] | 747 | P626440 | Advantech/IPC-6010PC,PCA-6149
|  | 8.3.0.1 Analog I/O |  | B | 0 | 0 | 0 | 0 |  | 0 | 24 | 16 | 40 | [6 empty] | 409 | P626437 | Indus. Comp./ADIO 1600
|  | 8.3.0.2 RAM |  | B | 0 | 0 | 0 | 0 |  | 0 | 24 | 16 | 40 | [6 empty] | 285 | P626439 | Chip merchant/16+8 MB 60ns
|  | 8.3.0.3 Disk/serial card |  | B | 0 | 0 | 0 | 0 |  | 0 | 24 | 16 | 40 | [6 empty] | 802 | P626437 | PromoX 486CPU HD etc
|  | 8.3.1 Eight axis stepper controller |  | B | 2 | 1 | 2000 | 4000 |  | 2000 | 24 | 16 | 40 | [6 empty] | 2119 | P722168 | Oregon Micro PC58-8S etc
|  | 8.3.2 Motor control cabilng |  | B | 1 |  | 200 | 200 |  | 0 | 8 |  | 32
|  |  | Subtotal |  | 3 | 1 |  | 4200 | 0 | 2000 | 128 | 80 | 232 | [6 empty] | 4362
8.4 Timestamp
|  | 8.4.1 Time code processor |  | B | 1 |  | 2000 | 2000 |  | 0 | 8 | 16 | 8
|  |  | Subtotal |  | 1 | 0 |  | 2000 | 0 | 0 | 8 | 16 | 8 | [6 empty] | 0
8.5 Temperature Monitor
|  | 8.5.1. VME 16 Channel Temp Monitor |  | B | 0 |  | 3600 | 0 |  | 0 | 16 | 16 | 40
|  |  | Subtotal |  | 0 | 0 |  | 0 | 0 |  | 16 | 16 | 40
8.6 Analog Input
| ? | 8.6.1. VME 16 Diff Input analog |  | B | 0 |  | 2400 | 0 |  | 0 | 16 | 8 | 24
| ? | 8.6.2 Cabilng |  | B | 0 |  | 100 | 0 |  | 0 | 8 |  | 32
|  |  | Subtotal |  | 0 | 0 |  | 0 | 0 | 0 | 24 | 8 | 56
8.7 Digital I/O
|  | 8.7.1. VME Digital I/O |  | B | 0 |  | 700 | 0 |  | 0 | 16 | 16 | 8
|  | 8.7.2 Cabilng |  | B | 0 |  | 100 | 0 |  | 0 | 8 |  | 32
|  |  | Subtotal |  | 0 | 0 |  | 0 | 0 | 0 | 24 | 16 | 40
8.8 Absolute Encoder Readout
|  | 8.8.1. Synchro/Resolver to Digital Board |  | B | 0 |  | 3000 | 0 |  | 0 | 24 | 32 | 40
|  | 8.8.2 Farrand Amplifier/ocs/converter |  | B | 0 |  | 1000 | 0 |  | 0 | 16 |  | 40
|  |  | Subtotal |  | 0 | 0 |  | 0 | 0 | 0 | 40 | 32 | 80
TEMPERATURE CONTROL
9. Temperature Control
| 9.1.1 Temperature Controller |  |  | B | 2 | 0 | 2600 | 5200 |  | 0 | 16 | 8 | 24
| 9.1.2 Si Diode calibrated |  |  | B | 3 |  | 355 | 1065 |  | 0 | 16 | 8 | 8 | [6 empty] | 0
| 9.1.3 Si Diode standard curve |  |  | B | 3 |  | 125 | 375 |  | 0 | 8 | 8 | 8 | [6 empty] | 0
|  |  | Subtotal |  | 8 | 0 |  | 6640 | 0 | 0 | 40 | 24 | 40 | [6 empty] | 0
REALTIME DISPLAY
10. Realtime Display | [19 empty] | 5979.75 | P722368 | Themis SPARC 5/64-32-110
| 10.1.1 IRTF designed DSP interface/RAM board |  |  | B | 0 |  | 6000 | 0 |  | 0 | 32 | 24 | 32 | [6 empty] | 1279 | P772367 | ZZYX/Integrix TGX120,etc
| 10.1.2 Vendor PCI graphics/video board |  |  | B | 0 | 0 | 1000 | 0 |  | 0 | 4 | 4 | 4 | [6 empty] | 190 | P772367 | Centon CUN32/M Sparc LXE PLUS
| 10.1.3 Video Monitor/amplifier |  |  | B | 0 |  | 1000 | 0 |  | 0 | 4 | [8 empty] | 6355 | P722747 | Lakeshore model 330
|  |  | Subtotal |  | 0 | 0 |  | 0 | 0 | 0 | 40 | 28 | 36 | [6 empty] | 13803.75
[11 empty] | 2520 | 1764
|  |  | Total |  | 240 | 49 |  | 155934 | 0 | 43868.8 | 1562 | 892 | 2648 | 163530.8 | 0 | 0 | 0 | 0 | 0 | 160063.1
[8 empty] | est total |  | 199802.8
11. Remaining Items
| 11.1 15V 12amp supply |  |  | B | 2 | 0 | 1638 | 3276 |  | 0 | 0 |  | 0
| 11.2 Motor and temp cabling |  |  | B | 4 | 0 | 300 | 1200 |  | 0 | 0 | 0 | 0
| 11.3 Utility (Littledog) PC cooling fans and connectors |  |  | B | 1 | 0 | 800 | 800 |  | 0 | 0
| 11.4 Repair of TCS GPS timestamp board |  |  | B | 1 | 0 | 400 | 400 |  | 0 | 0
| 11.5 Cal box power supply |  |  | B | 1 | 0 | 1638 | 1638 |  | 0 | 0
|  |  | Subtotal |  | 9 | 0 |  | 7314 | 0 | 0 | 0 | 0 | 0
[11 empty] | 0 | 0

## power_est
Nominal Power Dissipation Estimate
CRYOSTAT ASSEMBLIES |  |  |  |  | I (mA @ 5V) | I (mA @ +15V) | I (mA @ -15V) | I (mA @ +28V) | +5VDC | +15VDC | -15VDC | +28VDC | qty/bd
1. Array Electronics
1.1 Array Fanout Assembly
1.1.1 Printed Circuit Board
1.1.2 Hermetic Connectors
1.1.3 Shielding
1.1.4 Yamaichi socket
Subtotal
1.2 Array Cryostat Wiring
1.2.1 Cryo Coax wire
1.2.2 Constantan wire
1.2.3 Flexwire assy
1.2.4 Feedthrough Assy
Subtotal
1.3 Cold Preamplifier Assembly
1.3.1 Printed Circuit Board
1.3.2 Electronic Components
Subtotal
2. Motor Assemblies
2.1 Stepper Motor Assembly
2.1.1 Stepper Motor
2.1.1 Motor Wiring
2.1.2. Cold Test support
2.1.3 Hermetic connectors
Subtotal
3. Position Sensing Assemblies
3.1 Absolute Encoder Assembly
3.1.1 Inductosyn
3.1.2 Wiring
3.1.3 Hermetic connector
Subtotal
3.2 Limit switch/detent Assemblies
3.2.1 Cryogenic sensors?
3.2.2 Wiring
3.2.4 Signal Conditioning
Subtotal
CRYOSTAT MOUNTED ELECTRONICS ASSEMBLIES
4. Array readout electronics
7.5W | 4.1 Eight Channel Preamplifier
4.1.1 Printed Circuit Board
4.1.2 Component kit
|  |  | 4.1.2.1 OPAMP AD797 |  | 0
|  |  | 4.1.2.2 MUX HA210HS |  | 0 | 10 | 10
4.1.2.3 CPLD
4.1.3 Fabrication
Subtotal
9W | 4.2 Four Channel A/D converter
4.2.1 Printed Circuit Board
4.2.2 Component kit
4.2.3 2Mhz 16 bit ADC conveters
4.2.3 Fabrication
Subtotal
11W | 4.3 8/16 Channel Clock/Bias Driver
4.3.1 Printed Circuit Board
4.3.2 Component kit
4.3.3 Fabrication
Subtotal
5W | 4.4 Fibre Optic Interface (1Gbps/sec)
4.4.1 Printed Circuit Board
4.4.2 Component kit
4.4.3 Fabrication
Subtotal
2W | 4.5 Chassis & backplane
4.5.1 Custom Eurocard backplane
4.5.2 VME J2 backplane
4.5.3 Chassis
Subtotal
4.6 Cables (attached to cryostat)
4.6.1 Array Cables
4.6.2 Array cable connectors
4.6.3 Motor Cables
4.6.4 Motor cable connectors
4.6.5 Position sense cables
4.6.6 Pos sense cable connectors
Subtotal
ELECTRONICS POWER SUPPLIES
5. Array Power Supply
5.1.1 Chassis
5.1.2 Serial control interface
5.1.3 +5V 20amp supply
5.1.4 15V 12amp supply
5.1.5 Cabling/connectors
Subtotal
6. Motor Driver Power Supply
6.1.1 Stepper motor drive module
6.1.1 Chassis
? | 6.1.2 Serial control interface
? | 6.1.3 +5V 20amp supply
? | 6.1.4 15V 12amp supply
? | 6.1.5 24V 8amp supply
6.1.5 Cabling/connectors
Subtotal
7. AC Power Distribution/Control
7.1.1 Filtered AC distribution panel
7.1.2 Serial Control interface
Subtotal
INSTRUMENT CONTROL ELECTRONICS ASSEMBLIES
8. VME Control System
8.1 Communications and Processing
8.1.1 Single Board Computer
8.1.2 SCSI disk (>4 Gbyte)
8.1.3 Chassis/power supply
8.1.4 Rackmount keyboard
8.1.5 Monitor
8.1.6 Cables
Subtotal
8.2 Vendor DSP board
8.2.1 Multi DSP board
8.2.2 Support software
Subtotal
8.3 Motor Control
8.3.1 Eight axis stepper controller
8.3.2 Motor control cabilng
Subtotal
8.4 Timestamp
8.4.1 Time code processor
Subtotal
8.5 Temperature Monitor
8.5.1. VME 16 Channel Temp Monitor
Subtotal
8.6 Analog Input
| ? | 8.6.1. VME 16 Diff Input analog
| ? | 8.6.2 Cabilng
Subtotal
8.7 Digital I/O
8.7.1. VME Digital I/O
8.7.2 Cabilng
Subtotal
8.8 Absolute Encoder Readout
8.8.1. Synchro/Resolver to Digital Board
8.8.2 Farrand Amplifier/ocs/converter
Subtotal
TEMPERATURE CONTROL
9. Temperature Control
9.1.1 Temperature Controller
9.1.2 Si Diode calibrated
9.1.3 Si Diode standard curve
Subtotal
REALTIME DISPLAY
10. Realtime Display
10.1.1 IRTF designed DSP interface/RAM board
10.1.2 Vendor PCI graphics/video board
10.1.3 Video Monitor/amplifier
Subtotal

## PRTSEL
1 GIGABAUD+ LINK | Data rate | Bidirectional | Package | Input | Output | Parity | CRC | Auto fsync | Auto bsync | Sync | Vsupply | Pwr | I supply | Cost | Telephone | Distributor
AMCC S2039 XMT ENDEC | 1.0625 Gbit/sec | no | 132 LDCC | 32 TTL data,5 cmd | PECL | 0 | no | no | yes | ? | 5V | 3W |  | 342
AMCC S20340 RCVR ENDEC | 1.0625 Gbit/sec | no | 132 LDCC | PECL |  | 0 | no | no | yes | ? | 5V | 3.5W |  | 342
TRIQUINT GA9303 ENDEC | 1.0625 Gbaud | yes | 160 FP | 32 TTL data,8 cmd | 10 TTL | 4 | yes | no | yes | ? | 5V |  |  | 67
TRIQUINT TQ9502 RCVR | 1.0625 Gbaud |  | 44 MQUAD | 10 TTL | PECL | [5 empty] | 5V
TRIQUINT TQ9501 XMT | 1.0625 Gbaud |  | 45 MQUAD | PECL | [6 empty] | 5V
TRIQUINT GA9011 HOT ROD XMT | 1 Gbaud=800Mbit/sec (4B/5D) | no | 68 CLCC | 40 TTL data, 2 cmd | PECL | 0 | no | yes | yes | no | 5V |  | 310t/400mA max | 232
TRIQUINT GA9012 HOT ROD RCVR | 2 Gbaud=800Mbit/sec (4B/5D) | no | 69 CLCC | PECL |  | 0 | no |  |  |  | 5V |  | 360t/450mA max | 422
VITESSE VSC7107 ENDEC | 100MB/sec simplex 200MB Dup | yes | 184 PQFP | 32 TTL data, 6 cmd | 10 TTL | 4 | yes | no | yes | yes | 5V | 1.3W typ | 260mA typ | 55
VITESSE VSC7105 XMTR | up to 1.0625 Gbit/sec |  | 44 PQFP | 10 TTL | PECL | [5 empty] | 3.3V | 1.2W max | 350mA max | 75
VITESSE VSC7106 RCVR | up to 1.0625 Gbit/sec |  | 45 PQFP | PECL | [6 empty] | 3.3V | 1.8W max | 520mA max | 78
FIBER OPTIC TRANSCEIVERS
TRIQUINT
GA9040 FIBER OPTIC SET | [13 empty] | 950
GA9041 XMT |  | no | 2.25X.875X.560
GA9042 RCVR |  | no | 2.25X.875X.561
GA9043 LINK CNTL |  |  | 2.25 X1.5X.425
FINISAR
FTR-8510 XSCVR/LINK CNTLR | 100M-1.5Gbps | yes | 2.0X1.5X.4 | PECL | PECL | [5 empty] | 5V | 0.8W typ |  | 660 | (415)364-2722
FORCE INC.
2667T | 2.5Gbps | no | 1.5x.75x.75 | ECL 10K/100K | [6 empty] | 5V,-5V
METHODE
MRM-8150 RCVR | 1.5Gps | no | [11 empty] | 295 | (800)323-6858
MTM-8150 XMT | 1.5Gps | no | [11 empty] | 235
FIFO | Size | Cycle Time | Pkg | Iol,Ioh | Freq | Mailbox | Parity | Sync/Async | Prog Flags | 1st wd | Prog Req? | Prog Req? | Icc typ. | Cost |  | Distributor
CYPRESS CY7C464/474 | 32K X 9 | 15nsec |  |  | 33MHZ | NO | NO | [5 empty] | 100mA | 58
SHARP LH540215/25 | 512 X 18/1024 X 18 | 20nsec | 68PLCC | 16mA,-8.0mA | 50MHZ | NO | NO | both | PAF,PAE |  | NO |  | 250mA |  |  | Marshall
TEXAS INST SN74ABT7803 | 512 X 18 |  | DL 25mil/300mil | 24mA,-12mA | 67MHz | NO | NO | clocked | PAF/PAE | yes
TEXAS INST SN74ABT7805 | 256 x 18 |  | DL 25mil/300mil | 24mA,-12mA | 67MHz | NO | NO | clocked | PAF/PAE | yes |  |  |  | $26*
TEXAS INST SN74ABT7813 | 64 X 18 |  | DL 25mil/300mil | 24mA,-12mA | 67MHz | NO | NO | clocked | PAF/PAE | yes

## Sheet1 (2)
Pin Num | 61PIN
1 | Z | VP_T | COND
2 | Y | AGND | SHIELD
3 | u | VGGCL_T | SHIELD
4 | a | VSSCM_T | COND
5 | GG | VRWOFF_T | SHIELD
6 | FF | VRTSG_T | COND
7 | t | VROWON_T | SHIELD
8 | s | VIREF_T | COND
9 | MM | VRTSR_T | SHIELD
10 | NN | VTEND_T | COND
11 | EE | Q1OUT1- | SHIELD
12 | DD | Q1OUT1+ | COND
13 | X | Q1OUT2- | SHIELD
14 | W | Q1OUT2+ | COND
15 | r | Q1OUT3- | SHIELD
16 | q | Q1OUT3+ | COND
17 | V | Q1OUT4- | SHIELD
18 | U | Q1OUT4+ | COND
19 | T | Q1OUT5- | SHIELD
20 | S | Q1OUT5+ | COND
21 | p | Q1OUT6- | SHIELD
22 | n | Q1OUT6+ | COND
23 | CC | Q1OUT7- | SHIELD
24 | BB | Q1OUT7+ | COND
25 | R | Q1OUT8- | SHIELD
26 | P | Q1OUT8+ | COND
27 | m |  | SHIELD
28 | N | VDDOUT_T | COND
29 | PP
30 | NC | Q1TEND
31 | NC
32 | NC
33 | NC
34 | NC
Pin Num | 61PIN
1 | B | VDETCOM | SHIELD
2 | A | VDDOUT_T | COND
3 | c | Q1OUT8- | SHIELD
4 | b | Q1OUT8+ | COND
5 | w | Q1OUT7- | SHIELD
6 | v | Q1OUT7+ | COND
7 | e | Q1OUT6- | SHIELD
8 | d | Q1OUT6+ | COND
9 | JJ | Q1OUT5- | SHIELD
10 | HH | Q1OUT5+ | COND
11 | y | Q1OUT4- | SHIELD
12 | x | Q1OUT4+ | COND
13 | D | Q1OUT3- | SHIELD
14 | C | Q1OUT3+ | COND
15 | g | Q1OUT2- | SHIELD
16 | f | Q1OUT2+ | COND
17 | F | Q1OUT1- | SHIELD
18 | E | Q1OUT1+ | COND
19 | LL | VDDCL_T | SHIELD
20 | KK | VDDUC_T | COND
21 | AA | PSS_T | SHIELD
22 | z | PS1_T | COND
23 | i | PS2_T | SHIELD
24 | h | PROE_T | COND
25 | H | PDES_T | SHIELD
26 | G | PFS_T | COND
27 | K | PF1_T | SHIELD
28 | J | PF2_T | COND
29 | k | VNROW_T | SHIELD
30 | j | VNCOL_T | COND
31 | M |  | SHIELD
32 | L |  | COND
33 | NC
34 | NC

## Motors
Vendor | Model No. | Type | Torque (oz.in.) | Gearing | Form | Step Resolution | $ | Notes:
Anaheim Automation | 23D306D | Stepper | 160 Stat. | N/A | NEMA 23 | 400 Steps/Rev | 150 | Used on NSFCam
(714) 922-6990
Micro Mo Electronics, Inc. | 2338S-024S 62 + H23-1 43:1 + HEDM 5500J 06 + X0434 + X0550C | DC Motor | 100 Dyn. | 43:1 | 1" Dia | 1024 | 305 | 5-24 Pcs. Plntry Gearhd. Opt Enc.
(813) 572-0131 | 2338S-024S 62 + H23-1 246:1 + HEDM 5500J 06 + X0434 + X0550C | DC Motor | 100 Dyn. | 246:1 | 1" Dia | 1024 | 330 | 5-24 Pcs. Plntry Gearhd. Opt Enc.
Paula Mangan
Animatics Corporation | SM2310 SmartMotor - Short Stack | Servo | 70 Peak 15 Cont. |  | NEMA 23 | 2000 | 1520 | 5-9 $1400
3050 Tasman Drive | SM2330 SmartMotor - Long Stack | Servo | 150 Peak 35 Cont. |  | NEMA 23 | 2000 | 1725 | 5-9 $1600
Santa Clara, CA
95054-1116
(408) 748-8721
(408) 748-8725
Al Hastings
David - Engineer
Intelligent Motion Systems | M2-2220-S | Stepper | 72 Hold |  | NEMA 23 | 1.8 Deg | 83 | Hybrid
(860) 295-6102 | M2-2232-S | Stepper | 150 Hold |  | NEMA 23 | 1.8 Deg | 150 | Hybrid
FAX (860) 295-6107

## Motor Drivers
Vendor | Model No. | Type | Form | Step Resolution | Feedback | Amps/Phase | $ | Features
Oregon Micro Sys. | MD10A | Stepper | Brick | 25K/Rev |  | .72-3.6 | 425
(503) 629-8081 | MD125 | Stepper | Brick | 10-125 steps/step |  | .5-7.0 | 600
| MD400 | Servo | Brick | N/A |  | 0-20 | 600
Micro Mo. Electronics | BLD453 | Servo | 2x3.5x1" Open Frm. | N/A |  | 3 | 208 | PWM
| BLD568 | Servo | 2.5x4x1" Open Frm. | N/A |  | 8 | 244 | PWM
| LC3002A | Servo | 3U | N/A |  | 3 | 300 | Linear
Maxon Motor | MMC-QR-030 024-02 LE 00A | Servo | Big Brick | N/A | Opt. Enc. | 2 | 400 | Linear
| MMC-QR-030 024-06 PE00A | Servo | 3U | N/A | Opt. Enc. | 6 | 600 | PWM
Anaheim Automation | MBL536 | Stepper | Brick | 400 Steps/Rev |  | 5 | 175 | Bi-level, outputs-off
(714) 922-6990 | (replaced MBL500)
Galil | MBA 6-60 | Servo | ? |  |  | 6A/60V | 400 | For brushless motors
(415) 967-1700
Marlene - Sales
John - App. Eng.
Intelligent Motion Systems | IM483 | Stepper | 3x3x1.2" Open Frm | 400-51K Steps/Rev |  | .4-3 | 200 | Micro Stepper 20KHz. Chop
(860) 295-6102 | IM1007 | Stepper | 3x6x1" Brick | 400-51K Steps/Rev |  | 2.0-7 | 450 | Micro Stepper 20KHz. Chop
FAX (860) 295-6107

## Motor Controllers
Vendor | Model No. | Type | Form | # Axis | #User I/O | Intrfc. | $ | Notes:
Oregon Micro Sys. | PC58-8 | Stepper | PC | 8 | 22 | ISA PC Bus | 1700 | Can specify for both steppers
(503) 629-8081 | PC58-8S | Servo | PC | 8 | 22 | ISA PC Bus | 2000 | and servos on same board in
| VME58-8 | Stepper | VME | 8 | 22 | VME Bus | 2000 | increments of 2's
| VME58-8S | Servo | VME | 8 | 22 | VME Bus | 2500
Galil | DMC-120-10 | Servo | STD | 2 | ? | ? | 600
(415) 967-1700 | DMC-1000 | Servo | PC | 8 | ? | ? | 3200
Marlene - Sales | DMC-1500 | Servo | ? | 8 | ? | ? | 4000 | Stand-alone - no motor drive
John - App. Eng.
Polytech PI, Inc. | C-808.00 | Servo | 3U | 2 | ? | VME | 1400 | Motor Amp. built on board
(800) 832-3452 | C-804.10 | Servo | 19" Rack | 4 | 16 | RS-232,488 | 5000 | Stand-alone
| C-812.00 | Servo | PC | 4 | 16 | RS-232 | 2900 | 4 Analog inputs (8 Bits)
| C-812.10 | Servo | PC | 4 | 16 | RS-232,488 | 3100 | 4 Analog inputs (8 Bits)
Animatics Corporation
(408) 748-8721

## Hall-distance
| Dist. In. | Dist. mm | Output (mV) |  | Dist. In. | Dist. mm | Output (mV) |  | Dist. In. | Dist. mm | Output (mV) |  | Dist. In. | Dist. mm | Output (mV)
10 | 0 | 0 | 65 | 10 | 0 | 0 | 102.5 | 10 | 0 | 0 | 37.8 | 10 | 0 | 0 | 123.4
9.95 | 0.0019685 | 0.05 | 63.2 | 9.95 | 0.0019685 | 0.05 | 100.9 | 9.95 | 0.0019685 | 0.05 | 36.9 | 9.95 | 0.0019685 | 0.05 | 119.3
9.85 | 0.0059055 | 0.15 | 59 | 9.85 | 0.0059055 | 0.15 | 98.9 | 9.85 | 0.0059055 | 0.15 | 36.1 | 9.85 | 0.0059055 | 0.15 | 119.1
9.75 | 0.0098425 | 0.25 | 54.4 | 9.75 | 0.0098425 | 0.25 | 90.9 | 9.75 | 0.0098425 | 0.25 | 33.2 | 9.75 | 0.0098425 | 0.25 | 109.7
9.65 | 0.01378 | 0.35 | 48.8 | 9.65 | 0.01378 | 0.35 | 81.4 | 9.65 | 0.01378 | 0.35 | 29.9 | 9.65 | 0.01378 | 0.35 | 97.8
9.55 | 0.017717 | 0.45 | 42.9 | 9.55 | 0.017717 | 0.45 | 72.3 | 9.55 | 0.017717 | 0.45 | 25.9 | 9.55 | 0.017717 | 0.45 | 87
9.45 | 0.021654 | 0.55 | 37.1 | 9.45 | 0.021654 | 0.55 | 62.8 | 9.45 | 0.021654 | 0.55 | 22.9 | 9.45 | 0.021654 | 0.55 | 75.8
9.25 | 0.029528 | 0.75 | 32 | 9.25 | 0.029528 | 0.75 | 53.6 | 9.25 | 0.029528 | 0.75 | 19.5 | 9.25 | 0.029528 | 0.75 | 64.6
9.05 | 0.037402 | 0.95 | 25.6 | 9.05 | 0.037402 | 0.95 | 42.7 | 9.05 | 0.037402 | 0.95 | 15.5 | 9.05 | 0.037402 | 0.95 | 51.6
8.9 | 0.043307 | 1.1 | 21.7 | 8.9 | 0.043307 | 1.1 | 36.3 | 8.9 | 0.043307 | 1.1 | 13.3 | 8.9 | 0.043307 | 1.1 | 43.8
8.6 | 0.055118 | 1.4 | 17.3 | 8.6 | 0.055118 | 1.4 | 28.3 | 8.6 | 0.055118 | 1.4 | 10.4 | 8.6 | 0.055118 | 1.4 | 34.5
8.3 | 0.066929 | 1.7 | 13.5 | 8.3 | 0.066929 | 1.7 | 22.3 | 8.3 | 0.066929 | 1.7 | 8.2 | 8.3 | 0.066929 | 1.7 | 26.8
8 | 0.07874 | 2 | 10.7 | 8 | 0.07874 | 2 | 17.5 | 8 | 0.07874 | 2 | 6.4 | 8 | 0.07874 | 2 | 21.1
7.8 | 0.086614 | 2.2 | 9.5 | 7.8 | 0.086614 | 2.2 | 15.5 | 7.8 | 0.086614 | 2.2 | 5.7 | 7.8 | 0.086614 | 2.2 | 18.7
7.6 | 0.094488 | 2.4 | 8.4 | 7.6 | 0.094488 | 2.4 | 13.7 | 7.6 | 0.094488 | 2.4 | 5 | 7.6 | 0.094488 | 2.4 | 16.6
7.4 | 0.10236 | 2.6 | 7.4 | 7.4 | 0.10236 | 2.6 | 12 | 7.4 | 0.10236 | 2.6 | 4.4 | 7.4 | 0.10236 | 2.6 | 14.5
7.2 | 0.11024 | 2.8 | 6.7 | 7.2 | 0.11024 | 2.8 | 10.9 | 7.2 | 0.11024 | 2.8 | 4 | 7.2 | 0.11024 | 2.8 | 13.2
7 | 0.11811 | 3 | 6 | 7 | 0.11811 | 3 | 9.7 | 7 | 0.11811 | 3 | 3.6 | 7 | 0.11811 | 3 | 11.7
6.75 | 0.12795 | 3.25 | 5.7 | 6.75 | 0.12795 | 3.25 | 8.7 | 6.75 | 0.12795 | 3.25 | 3.2 | 6.75 | 0.12795 | 3.25 | 10.5
6.5 | 0.1378 | 3.5 | 4.8 | 6.5 | 0.1378 | 3.5 | 7.7 | 6.5 | 0.1378 | 3.5 | 2.8 | 6.5 | 0.1378 | 3.5 | 9.3
6.25 | 0.14764 | 3.75 | 4.4 | 6.25 | 0.14764 | 3.75 | 7 | 6.25 | 0.14764 | 3.75 | 2.6 | 6.25 | 0.14764 | 3.75 | 8.5
6 | 0.15748 | 4 | 4 | 6 | 0.15748 | 4 | 6.4 | 6 | 0.15748 | 4 | 2.3 | 6 | 0.15748 | 4 | 7.7
Re = 475 ohms |  |  |  | 301 ohms |  |  |  | 824 ohms |  |  |  | 250 ohms
9V sup @ 10.16mA |  |  |  | 9V sup @ 15.92mA |  |  |  | 9V sup @ 5.86mA |  |  |  | 9V sup @ 19.11mA
sarmarium cobalt magnets 0.098" diameter, PN SMC0056
| Dist. In. | Dist. mm | Output (mV)
| 0 | 0 | 65
| 0.0019685 | 0.05 | 63.2
| 0.0059055 | 0.15 | 59
| 0.0098425 | 0.25 | 54.4
| 0.01378 | 0.35 | 48.8
| 0.017717 | 0.45 | 42.9
| 0.021654 | 0.55 | 37.1
| 0.029528 | 0.75 | 32
| 0.037402 | 0.95 | 25.6
| 0.043307 | 1.1 | 21.7
| 0.055118 | 1.4 | 17.3
| 0.066929 | 1.7 | 13.5
| 0.07874 | 2 | 10.7
| 0.086614 | 2.2 | 9.5
| 0.094488 | 2.4 | 8.4
| 0.10236 | 2.6 | 7.4
| 0.11024 | 2.8 | 6.7
| 0.11811 | 3 | 6
| 0.12795 | 3.25 | 5.7
| 0.1378 | 3.5 | 4.8
| 0.14764 | 3.75 | 4.4
| 0.15748 | 4 | 4

## Sheet1 (3)
CLOCKING MODES:
1 | Multiplexer Optical Imaging |  | clock to each pixel reset/sample PEDESTAL/short ITIME/sample SIGNAL - clock to next pixel
USE: optical alignment of instrument
2 | Fowler Single Sampling/Global Reset |  | global reset whole array/ sample entire frame of PEDESTAL/variable ITIME/sample entire frame of SIGNAL
3 | Fowler Single Sampling/Row Reset
4 | Fowler Multiple Sampling/Endpoints
5 | Fowler Multiple Sampling/Sample up Ramp
6 | CDS/Global
7 | CDS/row
8 | Correlated Multiple Sampling
9 | Subarray readout
10 | Single Sample Mode (raw signal or pedestal)
11 | Clock into the integration at the same cadence as the background resets
Desired Engineering Tools
1 | Graphical Editor for clocking pattern
2 | Method for uniformly slowing down clocking sequence (slow counts)
3 | Automatic parametric testing (vary a DAC voltage or clock speed over a # of tests and display results)
|  | noise vs detector bias |  | well depth vs. reset duration
|  | noise vs readout speed |  | well depth vs. IIDLE
|  | noise vs temperature |  | well depth vs ISLEW
|  | noise vs. IIDLE |  | well depth vs. column clocking speed
|  | noise vs. ISLEW |  | persistance vs. reset time
|  | noise vs. VRSTG |  | readout speed vs. IIDLE
|  | noise vs. VRSTR |  | readout speed vs ISLEW
|  | noise vs. output# |  | readout speed vs VROWON
|  | dark current vs. temperature |  | output gain per channel
|  | dark current vs. detector bias |  | quantum efficiency vs temperature
|  | linearity vs. well depth |  | array offset level vs detector bias voltages
|  | linearity vs. detector bias |  | bad pixels vs. detector bias
well depth vs. detector bias
4 | Method for chaining different patterns together
5 | Method for picking between Fowler/CDS, Global/Row readout and single/multiple sample modes
6 | Method for excluding bad pixels from vf box statistics
7 | Method for saving all programmable voltages in named file
8 | Method for checking if extra pixels were shipped down fiber link
9 | Method for checking on progress of pixel readout

## Temp_cryo_wirelist
CRYOSTAT ASSEMBLIES | [16 empty] | TEMPERATURE CONTROL RACK/BOX
TEMPERATURE CONTROL | [17 empty] | DIN6 | DIN6-REAR |  |  |  |  | RJ45 | RJ45 | RJ45 | RJ45
FUNCTION |  |  |  | MOLEX | MOLEX |  | .1X2X7 | .1X2X7 |  |  | 55 PIN | 55 PIN | TEMP | 15FT | TEMP | TEMP |  | LAKESHORE | LAKESHORE |  | BANANA |  | BANANA | LAKESHORE | SPEC | SPEC | SERIAL
| CRYO | WIRE | LOCAL | CGRID | CGRID | ? INCH | 14PIN HDR | 14PIN HDR | ? INCH | IDC30 | HERMETIC | HERMETIC | 55 PIN | WIRE | 55 PIN | RACK | WIRE | 330 | 330 |  | LAKESHORE |  | LAKESHORE | 330 | SERIAL | SERIAL | FEED-
| NUM/ | COLOR | SIGNAL | 10 PIN | 10 PIN | *(CONST)ANTAN | FEED- | FEED- | MANGANIN | CONN | INTRF BD | CONN | CABLE | COLOR/ | CABLE | 55 PIN | COLOR/ | CABLE | SENSOR |  | 330 | WIRE | FILTER | RS232 | CABLE | CABLE | THROUGH
| NAME | CODE | NAME | CONN | CONN | CABLE | THROUGH | THROUGH | CABLE | PIN# | PIN # | PIN # | 55 CONN | AWG | 55 CONN | CONN | AWG | CONN | CONN |  | HEATER | COLOR/ | INPUT | CONN | CONN | CONN | CONN
|  |  |  | 1 | 1 | [15 empty] | CONN | AWG | CONN
|  |  |  | 2 | 2 | [15 empty] | HI | RED/14 | HI
SPEC ARRAY | SA_TS | BLK/BLU | I+ | 3 | 3 | *10mil CONST | 2 | 2 | 1 | 1 | JH1-1 | p | p |  | p | p |  | 5 [I+] CHNL A | 5 [I+] CHNL A |  | GND | WHT/14 | GND
TEMP SENSOR | SA_TS | YEL/CLR | V+ | 4 | 4 | *10mil CONST | 1 | 1 | 2 | 2 | JH1-2 | BB | BB |  | BB | BB |  | 4 [V+] CHNL A | 4 [V+] CHNL A |  | LO | BLK/14 | LO
| SA_TS | RED | I- | 5 | 5 | *10mil CONST | 4 | 4 | 3 | 3 | JH1-3 | q | q |  | q | q |  | 1 [I-] CHNL A | 1 [I-] CHNL A |  |  |  | OUTPUT
| SA_TS | GRN | V- | 6 | 6 | *10mil CONST | 3 | 3 | 4 | 4 | JH1-4 | V | V |  | V | V |  | 2 [V-] CHNL A | 2 [V-] CHNL A |  |  |  | CONN
SPEC ARRAY | R3¦¦R4+ | N/A | R3¦¦R4+ | 7 | 7 | Cu | 6 | 6 | 5 | 5 | JH1-5 | A | A |  | A | A | [5 empty] | RED/22 | HI
HEATER | R3¦¦R4- | N/A | R3¦¦R4- | 8 | 8 | Cu | 5 | 5 | 6 | 6 | JH1-6 | W | W |  | W | W | [6 empty] | GND
RESISTORS | R1¦¦R2+ | N/A | R1¦¦R2+ | 9 | 9 | Cu | 8 | 8 | 7 | 7 | JH1-7 | X | X |  | X | X
| R1¦¦R2- | N/A | R1¦¦R2- | 10 | 10 | Cu | 7 | 7 | 8 | 8 | JH1-8 | B | B |  | B | B | [5 empty] | BLK/22 | LO
COLD | CS1_TS | BLK/BLU | I+ | - | - | - | 10 | 10 | 9 | 9 | JH1-9 | Y | Y |  | Y | Y |  | 5 [I+] CHNL B | 5 [I+] CHNL B | SPECTROGRAPH
STRUCTURE | CS1_TS | YEL/CLR | V+ | - | - | - | 9 | 9 | 10 | 10 | JH1-10 | C | C |  | C | C |  | 4 [V+] CHNL B | 4 [V+] CHNL B | LAKESHORE 330
TEMP SENSOR | CS1_TS | RED | I- | - | - | - | 12 | 12 | 11 | 11 | JH1-11 | CC | CC |  | CC | CC |  | 1 [I-] CHNL B | 1 [I-] CHNL B | TEMP CONTROLLER
| CS1_TS | GRN | V- | - | - | - | 11 | 11 | 12 | 12 | JH1-12 | r | r |  | r | r |  | 2 [V-] CHNL B | 2 [V-] CHNL B
[8 empty] | N.C. |  |  | JH1-28,29 | J | J |  | J
[18 empty] | DIN6 | DIN6-REAR |  |  |  |  | RJ45 | RJ45 | RJ45 | RJ45
|  |  |  | MOLEX | MOLEX |  | .1X2X7 | .1X2X7 | [9 empty] | LAKESHORE | LAKESHORE |  | BANANA |  | BANANA | LAKESHORE | GUIDER | GUIDER | SERIAL
|  |  |  | CGRID | CGRID | ? INCH | 14PIN HDR | 14PIN HDR | ? INCH | IDC30 | [6 empty] | WIRE | 330 | 330 |  | LAKESHORE |  | LAKESHORE | 330 | SERIAL | SERIAL | FEED-
|  |  |  | 10 PIN | 10 PIN | *(CONST)ANTAN | FEED- | FEED- | MANGANIN | CONN | [6 empty] | COLOR/ | CABLE | SENSOR |  | 330 | WIRE | FILTER | RS232 | CABLE | CABLE | THROUGH
|  |  |  | CONN | CONN | CABLE | THROUGH | THROUGH | CABLE | PIN# | [6 empty] | AWG | CONN | CONN |  | HEATER | COLOR/ | INPUT | CONN | CONN | CONN | CONN
|  |  |  | 1 | 1 | [15 empty] | CONN | AWG | CONN
|  |  |  | 2 | 2 | [15 empty] | HI | RED/14 | HI
GUIDER ARRAY | GA_TS | BLK/BLU | I+ | 3 | 3 | *10mil CONST | 2 | 2 | 1 | 1 | JH2-1 | n | n |  | n | n |  | 5 [I+] CHNL A | 5 [I+] CHNL A |  | GND | WHT/14 | GND
TEMP SENSOR | GA_TS | YEL/CLR | V+ | 4 | 4 | *10mil CONST | 1 | 1 | 2 | 2 | JH2-2 | U | U |  | U | U |  | 4 [V+] CHNL A | 4 [V+] CHNL A |  | LO | BLK/14 | LO
| GA_TS | RED | I- | 5 | 5 | *10mil CONST | 4 | 4 | 3 | 3 | JH2-3 | T | T |  | T | T |  | 1 [I-] CHNL A | 1 [I-] CHNL A |  |  |  | OUTPUT
| GA_TS | GRN | V- | 6 | 6 | *10mil CONST | 3 | 3 | 4 | 4 | JH2-4 | AA | AA |  | AA | AA |  | 2 [V-] CHNL A | 2 [V-] CHNL A |  |  |  | CONN
GUIDER ARRAY | R3¦¦R4+ | N/A | R3¦¦R4+ | 7 | 7 | Cu | 6 | 6 | 5 | 5 | JH2-5 | z | z |  | z | z | [5 empty] | RED/22 | HI
HEATER | R3¦¦R4- | N/A | R3¦¦R4- | 8 | 8 | Cu | 5 | 5 | 6 | 6 | JH2-6 | m | m |  | m | m | [6 empty] | GND
RESISTORS | R1¦¦R2+ | N/A | R1¦¦R2+ | 9 | 9 | Cu | 8 | 8 | 7 | 7 | JH2-7 | k | k |  | k | k
| R1¦¦R2- | N/A | R1¦¦R2- | 10 | 10 | Cu | 7 | 7 | 8 | 8 | JH2-8 | S | S |  | S | S | [5 empty] | BLK/22 | LO
COLD | CS2_TS | BLK/BLU | I+ | - | - | - | 10 | 10 | 9 | 9 | JH2-9 | GG | GG |  | GG | GG |  | 5 [I+] CHNL B | 5 [I+] CHNL B | GUIDER
STRUCTURE | CS2_TS | YEL/CLR | V+ | - | - | - | 9 | 9 | 10 | 10 | JH2-10 | j | j |  | j | j |  | 4 [V+] CHNL B | 4 [V+] CHNL B | LAKESHORE 330
TEMP SENSOR | CS2_TS | RED | I- | - | - | - | 12 | 12 | 11 | 11 | JH2-11 | P | P |  | P | P |  | 1 [I-] CHNL B | 1 [I-] CHNL B | TEMP CONTROLLER
| CS2_TS | GRN | V- | - | - | - | 11 | 11 | 12 | 12 | JH2-12 | R | R |  | R | R |  | 2 [V-] CHNL B | 2 [V-] CHNL B
SEE LITTLEDOG
CABLING
SEE LITTLEDOG
CABLING

## Master_utility
CRYOSTAT ASSEMBLIES | [19 empty] | JBOX | [18 empty] | UTILITY PC (LITTLEDOG)
HALL EFFECT SENSORS |  |  |  | .1X1X4 | NON-STD |  |  |  |  | NON-STD | [14 empty] | HALL EFFECT PREAMPLIFIER BD | [32 empty] | PC IDE |  |  |  | PC_PS_H
FUNCTION |  |  |  | HALL | HALL |  |  |  |  | .1X2X5 | .1X2X5 |  | JH2 | 55 PIN | 55 PIN | HALL SENSOR | 36 INCH | HALL SENSOR | HALL |  | JH2 |  |  | HALL EFFECT | HALL EFFECT |  | HALL EFFECT | HALL EFFECT |  |  | JH1 |  | JBOX | JBOX | 15FT |  | JBOX | UTLILY PC |  | JH1 | [11 empty] | 1 | +5VDC |  |  |  |  | DISK |  |  |  | PC_PS_N |  | AC POWER |  | AC POWER
| CRYO | WIRE | LOCAL | SENSOR | SENSOR | ? INCH |  |  | ? INCH | 10PIN HDR | 10PIN HDR | ? INCH | 30 PIN | HERMETIC | HERMETIC | 55 PIN | WIRE | 55 PIN | SENSOR | 55 PIN | IDC30 | RIBBON | IDC64 | PREAMP | PREAMP |  | PREAMP | PREAMP | 16 PIN |  | 30 PIN | POSITION | POSITION | POSITION | WIRE |  | POSITION | POSITION | POSITION | 30 PIN | [10 empty] | 20 |  | CTR2 OUT | [8 empty] | PC_PS_GND |  | TERM |  | INPUT
| NUM/ | COLOR | SIGNAL | MOUNT | CABLE | MANGANIN |  |  | MANGANIN | FEED- | FEED- | MANGANIN | IDC | INTRF BD | HALL | CABLE | COLOR/ | CABLE | 55 PIN | INTRF BD | CONN | 28AWG | CONN | 96PIN CONN | LOCAL |  | LOCAL | VOUT | IDC | SIGNAL | IDC | INTRF BD | 55 PIN | CABLE | COLOR/ | SIGNAL | CABLE | 55 PIN | INTRF BD | IDC | [11 empty] | 2 | CTR0 OUT | [10 empty] | STRIP |  | EMI
| NAME | CODE | NAME | CONN | CONN | CABLE |  |  | CABLE | THROUGH | THROUGH | CABLE | CONN | PIN # | CONN | 55 CONN | AWG | 55 CONN | CONN | PIN # | PIN# | # | PIN# | PIN# | NAME |  | NAME | 16PIN HDR | CONN | NAME | CONN | PIN # | CONN | 55 CONN | AWG | NAME | 55 CONN | CONN | PIN # | CONN | [7 empty] | ICS ADIO1600 |  |  | 21 |  | CTR0 IN | [8 empty] | SERVO_H |  |  |  | CONN
[44 empty] | 19 |  |  | 19 | ANALOG/DIGITAL |  |  |  | 3 | OP3 | [8 empty] | SERVO_N |  |  | BLK
DICHROIC | DICTRT_HS | RED | I+ | 1 | 1 |  |  |  |  | 2 | 1 | 1 | 1 | JH2-1 | n | n |  | n | n | JH2-1 | 1 | 1 | P1-C2 | P1-C2 | CH1_I+ |  | CH1_OUT | JP2-1 | 1 | DICTRT_POS | 1 | JH1-1 | p | p |  | DICTRT_POS | p | p | JH1-1 | 1 | DICTRT_POS | CH0 HI | 37 |  |  | 37 |  | I/O BOARD |  |  | 22 |  | OP2 | [8 empty] | SERVO_GND |  |  | WHT
TURRET | DICTRT_HS | BLACK | I- | 2 | 2 |  |  |  |  | 1 | 2 | 2 | 2 | JH2-2 | U | U |  | U | U | JH2-2 | 2 | 2 | P1-A2 | P1-A2 | AGND |  | AGND | JP2-2 | 2 | AGND | 2 | JH1-2 | BB | BB |  | AGND | BB | BB | JH1-2 | 2 | AGND | CH0LO/CH8HI |  | 18 |  |  | 18 |  |  |  |  | 4 | OP1 | [8 empty] | STEPPER1_H |  |  | GRN
DETENT | DICTRT_HS | GREEN | V+ | 3 | 3 |  |  |  |  | 4 | 3 | 3 | 3 | JH2-3 | T | T |  | T | T | JH2-3 | 3 | 3 | P1-C3 | P1-C3 | CH1_VIN+ |  | CH2_OUT | JP2-3 | 3 | OSFLTR_POS | 3 | JH1-3 | q | q |  | OSFLTR_POS | q | q | JH1-3 | 3 | OSFLTR_POS | CH1 HI | 36 |  |  | 36 |  |  |  |  | 23 |  | OP0 | [8 empty] | STEPPER1_N
HALL EFFECT | DICTRT_HS | YELLOW | V- | 4 | 4 |  |  |  |  | 3 | 4 | 4 | 4 | JH2-4 | AA | AA |  | AA | AA | JH2-4 | 4 | 4 | P1-A3 | P1-A3 | CH1_VIN- |  | AGND | JP2-4 | 4 | AGND | 4 | JH1-4 | V | V |  | AGND | V | V | JH1-4 | 4 | AGND | CH1LO/CH9HI |  | 17 |  |  | 17 |  |  |  |  | 5 | IP3 | [8 empty] | STEPPER1_GND |  |  |  | UPPER
[5 empty] | Note Chng |  | 16P HDR | 16P HDR | [18 empty] | CH3_OUT | JP2-5 | 5 | SLTWHL_POS | 5 | JH1-5 | A | A |  | SLTWHL_POS | A | A | JH1-5 | 5 | SLTWHL_POS | CH2 HI | 35 |  |  | 35 |  |  |  |  | 24 |  | IP2/CTR0 GATE | [8 empty] | STEPPER2_H |  |  |  | FAN
[5 empty] | **** |  | FEEDTHRU | FEEDTHRU | [18 empty] | AGND | JP2-6 | 6 | AGND | 6 | JH1-6 | W | W |  | AGND | W | W | JH1-6 | 6 | AGND | CH2LO/CH10HI |  | 16 |  |  | 16 |  |  |  |  | 6 | IP1 | [8 empty] | STEPPER2_N |  |  |  | AC
ORDER SORTER | OSFLTR_HS | RED | I+ | 1 | 2 |  | 2 | 1 |  | 2 | 1 | 5 | 5 | JH2-5 | z | z |  | z | z | JH2-5 | 5 | 5 | P1-C4 | P1-C4 | CH2_I+ |  | CH4_OUT | JP2-7 | 7 | GFLWHL_POS | 7 | JH1-7 | X | X |  | GFLWHL_POS | X | X | JH1-7 | 7 | GFLWHL_POS | CH3 HI | 34 |  |  | 34 |  |  |  |  | 25 |  | IP0/TRIG0 | [8 empty] | STEPPER2_GND |  |  |  | CONN
FILTER | OSFLTR_HS | BLACK | I- | 2 | 1 |  | 1 | 2 |  | 1 | 2 | 6 | 6 | JH2-6 | m | m |  | m | m | JH2-6 | 6 | 6 | P1-A4 | P1-A4 | AGND |  | AGND | JP2-8 | 8 | AGND | 8 | JH1-8 | B | B |  | AGND | B | B | JH1-8 | 8 | AGND | CH3LO/CH11HI |  | 15 |  |  | 15 |  |  |  |  | 7 | COM |  |  |  |  | PCA-6149 | [7 empty] | UPPER
DETENT | OSFLTR_HS | GREEN | V+ | 3 | 4 |  | 4 | 3 |  | 4 | 3 | 7 | 7 | JH2-7 | k | k |  | k | k | JH2-7 | 7 | 7 | P1-C5 | P1-C5 | CH2_VIN+ |  | CH5_OUT | JP2-9 | 9 | SFOCUS_POS | 9 | JH1-9 | Y | Y |  | SFOCUS_POS | Y | Y | JH1-9 | 9 | SFOCUS_POS | CH4 HI | 33 |  |  | 33 |  |  |  |  | 26 |  | N.C. |  |  |  |  | 486 PC CPU CARD | [7 empty] | FAN
HALL EFFECT | OSFLTR_HS | YELLOW | V- | 4 | 3 |  | 3 | 4 |  | 3 | 4 | 8 | 8 | JH2-8 | S | S |  | S | S | JH2-8 | 8 | 8 | P1-A5 | P1-A5 | CH2_VIN- |  | AGND | JP2-10 | 10 | AGND | 10 | JH1-10 | C | C |  | AGND | C | C | JH1-10 | 10 | AGND | CH4LO/CH12HI |  | 14 |  |  | 14 |  |  |  |  | 8 | N.C. | [12 empty] | AC
[10 empty] | 14PIN HDR | 14PIN HDR | [15 empty] | CH6_OUT | JP2-11 | 11 | FPM_POS | 11 | JH1-11 | CC | CC |  | FPM_POS | CC | CC | JH1-11 | 11 | FPM_POS | CH5 HI | 32 |  |  | 32 |  |  |  |  | 27 |  | DAC1 OUT | [12 empty] | CONN
[27 empty] | AGND | JP2-12 | 12 | AGND | 12 | JH1-12 | r | r |  | AGND | r | r | JH1-12 | 12 | AGND | CH5LO/CH13HI |  | 13 |  |  | 13 |  |  |  |  | 9 | DAC0 OUT
SLIT WHEEL | SLTWHL_HS | RED | I+ | 1 | 1 |  |  |  |  | 2 | 1 | 9 | 9 | JH2-9 | GG | GG |  | GG | GG | JH2-9 | 9 | 9 | P1-C6 | P1-C6 | CH3_I+ |  | CH7_OUT | JP2-13 | 13 | ROTATR_POS | 13 | JH1-13 | Z | Z |  | ROTATR_POS | Z | Z | JH1-13 | 13 | ROTATR_POS | CH6 HI | 31 |  |  | 31 |  |  |  |  | 28 |  | L.L. GND | [8 empty] | KEYBD |  |  |  | KEYBD
HOME | SLTWHL_HS | BLACK | I- | 2 | 2 |  |  |  |  | 1 | 2 | 10 | 10 | JH2-10 | j | j |  | j | j | JH2-10 | 10 | 10 | P1-A6 | P1-A6 | AGND |  | AGND | JP2-14 | 14 | AGND | 14 | JH1-14 | D | D |  | AGND | D | D | JH1-14 | 14 | AGND | CH6LO/CH14HI |  | 12 |  |  | 12 |  |  |  |  | 10 | N.C. | [8 empty] | MOUSE |  |  |  | MOUSE
POSITION | SLTWHL_HS | GREEN | V+ | 3 | 3 |  |  |  |  | 4 | 3 | 11 | 11 | JH2-11 | P | P |  | P | P | JH2-11 | 11 | 11 | P1-C7 | P1-C7 | CH3_VIN+ |  | CH8_OUT | JP2-15 | 15 | GRTTRT_POS | 15 | JH1-15 | E | E |  | GRTTRT_POS | E | E | JH1-15 | 15 | GRTTRT_POS | CH7 HI | 30 |  |  | 30 |  |  |  |  | 29 |  | L.L. GND | [8 empty] | CONN |  |  |  | CONN
HALL EFFECT | SLTWHL_HS | YELLOW | V- | 4 | 4 |  |  |  |  | 3 | 4 | 12 | 12 | JH2-12 | R | R |  | R | R | JH2-12 | 12 | 12 | P1-A7 | P1-A7 | CH3_VIN- |  | AGND | JP2-16 | 16 | AGND | 16 | JH1-16 | s | s |  | AGND | s | s | JH1-16 | 16 | AGND | CH7LO/CH15HI |  | 11 |  |  | 11 |  |  |  |  | 11 | CH7LO/CH15HI | [8 empty] | VGA |  |  |  | VGA
[51 empty] | 30 |  | CH7 HI | [8 empty] | CONN |  |  |  | CONN
[52 empty] | 12 | CH6LO/CH14HI
GUIDER | GFLWHL_HS | RED | I+ | 1 | 1 |  |  |  |  | 2 | 1 | 13 | 13 | JH2-13 | i | i |  | i | i | JH2-13 | 13 | 13 | P1-C8 | P1-C8 | CH4_I+ | [25 empty] | 31 |  | CH6 HI | [9 empty] | PC POWER |  |  | POWER
FILTER WHL | GFLWHL_HS | BLACK | I- | 2 | 2 |  |  |  |  | 1 | 2 | 14 | 14 | JH2-14 | y | y |  | y | y | JH2-14 | 14 | 14 | P1-A8 | P1-A8 | AGND | [26 empty] | 13 | CH5LO/CH13HI | [9 empty] | SUPPLY |  |  | SLTWHLITCH
DETENT | GFLWHL_HS | GREEN | V+ | 3 | 3 |  |  |  |  | 4 | 3 | 15 | 15 | JH2-15 | N | N |  | N | N | JH2-15 | 15 | 15 | P1-C9 | P1-C9 | CH4_VIN+ | [25 empty] | 32 |  | CH5 HI
HALL EFFECT | GFLWHL_HS | YELLOW | V- | 4 | 4 |  |  |  |  | 3 | 4 | 16 | 16 | JH2-16 | h | h |  | h | h | JH2-16 | 16 | 16 | P1-A9 | P1-A9 | CH4_VIN- | [26 empty] | 14 | CH4LO/CH12HI
[10 empty] | 16PIN HDR | 16PIN HDR | [39 empty] | 33 |  | CH4 HI | [6 empty] | 3COM 3C590-COMBO
[52 empty] | 15 | CH3LO/CH11HI | [6 empty] | ETHERLINK III |  | 10BASET |  |  | 10BASET | 10BASET
SPEC | SFOCUS_HS | RED | I+ | 1 | 1 |  |  |  |  | 14 | 13 | 17 | 17 | JH2-17 | x | x |  | x | x | JH2-17 | 17 | 17 | P1-C10 | P1-C10 | CH5_I+ | [25 empty] | 34 |  | CH3 HI | [6 empty] | BOARD |  | ETHERNET |  |  | ETHERNET | ETHERNET
ARRAY | SFOCUS_HS | BLACK | I- | 2 | 2 |  |  |  |  | 13 | 14 | 18 | 18 | JH2-18 | FF | FF |  | FF | FF | JH2-18 | 18 | 18 | P1-A10 | P1-A10 | AGND | [26 empty] | 16 | CH2LO/CH10HI | [8 empty] | RJ45 |  |  | RJ45 | RJ45 FEED-
FOCUS | SFOCUS_HS | GREEN | V+ | 3 | 3 |  |  |  |  | 16 | 15 | 19 | 19 | JH2-19 | M | M |  | M | M | JH2-19 | 19 | 19 | P1-C11 | P1-C11 | CH5_VIN+ | [25 empty] | 35 |  | CH2 HI | [8 empty] | CABLE |  |  | CABLE | THROUGH
HALL EFFECT | SFOCUS_HS | YELLOW | V- | 4 | 4 |  |  |  |  | 15 | 16 | 20 | 20 | JH2-20 | g | g |  | g | g | JH2-20 | 20 | 20 | P1-A11 | P1-A11 | CH5_VIN- | [26 empty] | 17 | CH1LO/CH9HI | [8 empty] | CONN |  |  | CONN | CONN
[51 empty] | 36 |  | CH1 HI | [7 empty] | TD+ | 1 |  | TD+ | 1 | 1
[52 empty] | 18 | CH0LO/CH8HI | [7 empty] | TD- | 2 |  | TD- | 2 | 2
SPARE | FPM_HS | RED | I+ | 1 | 1 |  |  |  |  | 2 | 1 | 21 | 21 | JH2-21 | L | L |  | L | L | JH2-21 | 21 | 21 | P1-C12 | P1-C12 | CH6_I+ | [25 empty] | 37 |  | CH0 HI | [7 empty] | RD+ | 3 |  | RD+ | 3 | 3
MASK | FPM_HS | BLACK | I- | 2 | 2 |  |  |  |  | 1 | 2 | 22 | 22 | JH2-22 | w | w |  | w | w | JH2-22 | 22 | 22 | P1-A12 | P1-A12 | AGND | [26 empty] | 19 | L.L GND | [7 empty] | RD- | 6 |  | RD- | 6 | 6
POSITION | FPM_HS | GREEN | V+ | 3 | 3 |  |  |  |  | 4 | 3 | 23 | 23 | JH2-23 | f | f |  | f | f | JH2-23 | 23 | 23 | P1-C13 | P1-C13 | CH6_VIN+
HALL EFFECT | FPM_HS | YELLOW | V- | 4 | 4 |  |  |  |  | 3 | 4 | 24 | 24 | JH2-24 | K | K |  | K | K | JH2-24 | 24 | 24 | P1-A13 | P1-A13 | CH6_VIN- |  | HALL EFFECT | P2 | P2 |  | JH2 | [8 empty] | JH2
[5 empty] | Note Chng |  | 16P HDR | 16P HDR | [18 empty] | PREAMP | 5 PIN | 5 PIN |  | 30 PIN | [8 empty] | 30 PIN
[5 empty] | **** |  | FEEDTHRU | FEEDTHRU | [18 empty] | LOCAL | POWER | POWER |  | IDC | [8 empty] | IDC
ROTATOR | ROTATR_HS | RED | I+ | 1 | 2 |  | 6 | 5 |  | 2 | 1 | 25 | 25 | JH2-25 | v | v |  | v | v | JH2-25 | 25 | 25 | P1-C14 | P1-C14 | CH7_I+ |  | NAME | CONN | CONN |  | CONN | [8 empty] | CONN
HOME | ROTATR_HS | BLACK | I- | 2 | 1 |  | 5 | 6 |  | 1 | 2 | 26 | 26 | JH2-26 | e | e |  | e | e | JH2-26 | 26 | 26 | P1-A14 | P1-A14 | AGND
POSITION | ROTATR_HS | GREEN | V+ | 3 | 4 |  | 8 | 7 |  | 4 | 3 | 27 | 27 | JH2-27 | EE | EE |  | EE | EE | JH2-27 | 27 | 27 | P1-C15 | P1-C15 | CH7_VIN+ |  | +5V | P2-1 | 1 | PC+5V | 1 | JH2-1 | n | n |  | PC+5V | n | n | JH2-1 | 1
HALL EFFECT | ROTATR_HS | YELLOW | V- | 4 | 3 |  | 7 | 8 |  | 3 | 4 | 28 | 28 | JH2-28 | HH | HH |  | HH | HH | JH2-28 | 28 | 28 | P1-A15 | P1-A15 | CH7_VIN- |  | DGND | P2-2 | 2 | PC_GND | 2 | JH2-2 | U | U |  | PC_GND | U | U | JH2-2 | 2
[13 empty] | JH1 | [13 empty] | +15V | P2-3 | 3 | PC+12V | 3 | JH2-3 | T | T |  | PC+12V | T | T | JH2-3 | 3 | [23 empty] | OMS PC58-8S
[13 empty] | 30P IDC | [13 empty] | AGND | P2-4 |  | ? | [5 empty] | ? | [25 empty] | AMP749621-9 |  | STEPPER MOTOR
GRATING | GRTTRT_HS | RED | I+ | 1 | 1 |  |  |  |  | 2 | 1 | 1 | 1 | JH1-1 | p | p |  | p | p | JH1-1 | 1 | 1 | P1-C16 | P1-C16 | CH8_I+ |  | -15V | P2-5 | 5 | PC-12V | 4 | JH2-4 | AA | AA |  | PC-12V | AA | AA | JH2-4 | 4 | [21 empty] | OMS PC58-8 | OMS PC58-8 | CONTROLLER
TURRET | GRTTRT_HS | BLACK | I- | 2 | 2 |  |  |  |  | 1 | 2 | 2 | 2 | JH1-2 | BB | BB |  | BB | BB | JH1-2 | 2 | 2 | P1-A16 | P1-A16 | AGND | [37 empty] | MTR CONTL
HOME POSITION | GRTTRT_HS | GREEN | V+ | 3 | 3 |  |  |  |  | 4 | 3 | 3 | 3 | JH1-3 | q | q |  | q | q | JH1-3 | 3 | 3 | P1-C17 | P1-C17 | CH8_VIN+ | [36 empty] | CABLE | 100
HALL EFFECT | GRTTRT_HS | YELLOW | V- | 4 | 4 |  |  |  |  | 3 | 4 | 4 | 4 | JH1-4 | V | V |  | V | V | JH1-4 | 4 | 4 | P1-A17 | P1-A17 | CH8_VIN- | [36 empty] | 100 CONN | CONN
[14 empty] | JH1-28,29 | J | J |  | J | J | JH1-28,29 | [42 empty] | J93
[23 empty] | P1-C18 | P1-C18 | CH1_CMP+ | DICTRT_HOME |  |  |  |  | 5 | JH2-5 | z | z |  | DICTRT_HOME | z | z | JH2-5 | 5 | DICTRT_HOME | [18 empty] | DICTRT_HOME |  | 14 | 14 | XHOME
[23 empty] | P1-A18 | P1-A18 | DGND | DGND |  |  |  |  | 6 | JH2-6 | m | m |  | DGND | m | m | JH2-6 | 6 | DGND | [18 empty] | DGND |  | 58 | 58 | GND
[23 empty] | P1-C19 | P1-C19 | CH1_CMP- | DICTRT_DETENT |  |  |  |  | 7 | JH2-7 | k | k |  | DICTRT_DETENT | k | k | JH2-7 | 7 | DICTRT_DETENT | [18 empty] | DICTRT_DETENT |  | 1 | 1 | USERI/O 0
[23 empty] | P1-A19 | P1-A19 | DGND | DGND |  |  |  |  | 8 | JH2-8 | S | S |  | DGND | S | S | JH2-8 | 8 | DGND | [18 empty] | DGND |  | 60 | 60 | GND
[23 empty] | P1-C20 | P1-C20 | CH2_CMP+ | OSFLTR_HOME |  |  |  |  | 9 | JH2-9 | GG | GG |  | OSFLTR_HOME | GG | GG | JH2-9 | 9 | OSFLTR_HOME | [18 empty] | OSFLTR_HOME |  | 19 | 19 | YHOME
[23 empty] | P1-A20 | P1-A20 | DGND | DGND |  |  |  |  | 10 | JH2-10 | j | j |  | DGND | j | j | JH2-10 | 10 | DGND | [18 empty] | DGND |  | 70 | 70 | GND
[23 empty] | P1-C21 | P1-C21 | CH2_CMP- | OSFLTR_DETENT |  |  |  |  | 11 | JH2-11 | P | P |  | OSFLTR_DETENT | P | P | JH2-11 | 11 | OSFLTR_DETENT | [18 empty] | OSFLTR_DETENT |  | 52 | 52 | USERI/O 1
[23 empty] | P1-A21 | P1-A21 | DGND | DGND |  |  |  |  | 12 | JH2-12 | R | R |  | DGND | R | R | JH2-12 | 12 | DGND | [18 empty] | DGND |  | 81 | 81 | GND
[23 empty] | P1-C22 | P1-C22 | CH3_CMP+ | SLTWHL_HOME |  |  |  |  | 13 | JH2-13 | i | i |  | SLTWHL_HOME | i | i | JH2-13 | 13 | SLTWHL_HOME | [18 empty] | SLTWHL_HOME |  | 25 | 25 | ZHOME
[23 empty] | P1-A22 | P1-A22 | N.C. | DGND |  |  |  |  | 14 | JH2-14 | y | y |  | DGND | y | y | JH2-14 | 14 | DGND | [18 empty] | DGND |  | 91 | 91 | GND
[23 empty] | P1-C23 | P1-C23 | CH3_CMP- | SLTWHL_DETENT |  |  |  |  | 15 | JH2-15 | N | N |  | SLTWHL_DETENT | N | N | JH2-15 | 15 | SLTWHL_DETENT | [18 empty] | SLTWHL_DETENT |  | 2 | 2 | USERI/O 2
SERVO MOTORS | [22 empty] | P1-A23 | P1-A23 | DGND | DGND |  |  |  |  | 16 | JH2-16 | h | h |  | DGND | h | h | JH2-16 | 16 | DGND | [18 empty] | DGND
[13 empty] | ROTATOR | [9 empty] | P1-C24 | P1-C24 | CH4_CMP+ | GFLWHL_HOME |  |  |  |  | 17 | JH2-17 | x | x |  | GFLWHL_HOME | x | x | JH2-17 | 17 | GFLWHL_HOME | [18 empty] | GFLWHL_HOME |  | 29 | 29 | THOME
[13 empty] | ANMIMATICS | [9 empty] | P1-A24 | P1-A24 | DGND | DGND |  |  |  |  | 18 | JH2-18 | FF | FF |  | DGND | FF | FF | JH2-18 | 18 | DGND | [18 empty] | DGND
[13 empty] | SERVO | [9 empty] | P1-C25 | P1-C25 | CH4_CMP- | GFLWHL_DETENT |  |  |  |  | 19 | JH2-19 | M | M |  | GFLWHL_DETENT | M | M | JH2-19 | 19 | GFLWHL_DETENT | [18 empty] | GFLWHL_DETENT |  | 53 | 53 | USERI/O 3
AMPS= | 6 | LENGRTTRTH= | 15 | [9 empty] | MOTOR | [9 empty] | P1-A25 | P1-A25 | DGND | DGND |  |  |  |  | 20 | JH2-20 | g | g |  | DGND | g | g | JH2-20 | 20 | DGND | [18 empty] | DGND
CABLE | [9 empty] | GRATING | [12 empty] | P1-C26 | P1-C26 | CH5_CMP+ | SFOCUS+ |  |  |  |  | 21 | JH2-21 | L | L |  | SFOCUS+ | L | L | JH2-21 | 21 | SFOCUS+ | [18 empty] | SFOCUS+ |  | 84 | 84 | ULIMIT+
GAUGE | ohms/ft | 48V drop | [7 empty] | ANMIMATICS | [12 empty] | P1-A26 | P1-A26 | DGND | DGND |  |  |  |  | 22 | JH2-22 | w | w |  | DGND | w | w | JH2-22 | 22 | DGND | [18 empty] | DGND
AWG14 | 0.00252 | #REF! | [7 empty] | SERVO | [12 empty] | P1-C27 | P1-C27 | CH5_CMP- | SFOCUS- |  |  |  |  | 23 | JH2-23 | f | f |  | SFOCUS- | f | f | JH2-23 | 23 | SFOCUS- | [18 empty] | SFOCUS- |  | 85 | 85 | ULIMIT-
AWG16 | 0.00402 | #REF! | [7 empty] | MOTOR |  |  | CGRID 7PIN |  |  |  | ~2FT | [5 empty] | P1-A27 | P1-A27 | DGND | DGND |  |  |  |  | 24 | JH2-24 | K | K |  | DGND | K | K | JH2-24 | 24 | DGND | [18 empty] | DGND
AWG18 | 0.00638 | #REF! | [10 empty] | MOTOR | SERVO |  |  | WIRE | SERVO | SERVO |  |  |  | P1-C28 | P1-C28 | CH6_CMP+ | FPM+ |  |  |  |  | 25 | JH2-25 | v | v |  | FPM- | v | v | JH2-25 | 25 | FPM+ | [18 empty] | FPM- |  | 88 | 88 | VLIMIT+
AWG20 | 0.0101 | #REF! | [10 empty] | 7 PIN I/O | CABLE |  |  | COLOR/ | MOTOR | MOTRO |  |  |  | P1-A28 | P1-A28 | DGND | DGND |  |  |  |  | 26 | JH2-26 | e | e |  | DGND | e | e | JH2-26 | 26 | DGND | [18 empty] | DGND
AWG22 | 0.01614 | #REF! | [10 empty] | CONN | 7P CONN |  | ? | AWG | 14 PIN | 14 PIN |  |  |  | P1-C29 | P1-C29 | CH6_CMP- | FPM- |  |  |  |  | 27 | JH2-27 | EE | EE |  | FPM- | EE | EE | JH2-27 | 27 | FPM- | [18 empty] | FPM- |  | 89 | 89 | VLIMIT-
[10 empty] | CGRID 7PIN | [7 empty] | CONN | CONN |  |  |  | P1-A29 | P1-A29 | DGND | DGND |  |  |  |  | 28 | JH2-28 | HH | HH |  | DGND | HH | HH | JH2-28 | 28 | DGND | [18 empty] | DGND
[10 empty] | MOTOR | SERVO |  | 1 [ENCODER A] | 1 | ROTATR_HOME |  | BLK | 1 | 1 | 1 |  | 1 | P1-C30 | P1-C30 | CH7_CMP+ | ROTATR_HOME
CONNECTOR TYPES | [9 empty] | 7 PIN I/O | CABLE |  | 7 [GND] | 7 |  |  | BRN | 2 | 2 | 2 |  | 2 | P1-A30 | P1-A30 | DGND
COND | TYPE | CONT |  | PART NUM | [5 empty] | CONN | 7P CONN | [8 empty] | N.C. | N.C. | 3 | P1-C31 | P1-C31 | CH7_CMP-
100 | PLUG | PIN |  | AMP-749621-9 +749081-1 HD | [15 empty] | N.C. | N.C. | 4 | P1-A31 | P1-A31 | N.C. |  | CGRID 4PIN
[10 empty] | 1 [ENCODER A] | 1 | GRTTRT_HOME |  |  |  |  | BLK | 3 | 3 | 4 |  | 5 | P1-C32 | P1-C32 | CH8_CMP+ |  | 1 | 5V
[10 empty] | 7 [GND] | 7 | [5 empty] | BRN | 4 | 4 | N.C. | N.C. | 6 | P1-A32 | P1-A32 | -15V |  | 2 | FPM+ | [22 empty] | CYCLADES MUTIPORT SERIAL BOARD
41 | PLUG | SOCKET |  | MS3116-F2041S | [8 empty] | CD7WFW2000X | [13 empty] | 3 | FPM- |  |  | JH1 | [8 empty] | JH1 |  | PORT1
41 | RECT |  | HERMETIC | PT02-H20-41P | [8 empty] | MOTOR | [13 empty] | 4 | DGND |  |  | 30 PIN | [8 empty] | 30 PIN |  | CYCLOM-8YO | [5 empty] | CYCLOM-8YO |  | CYCLOM-8YO
26 | PLUG | SOCKET |  | HR-145-ND | [8 empty] | MAIN | [13 empty] | (*SEE BELOW*) |  |  |  | IDC | [8 empty] | IDC |  | CABLE | [5 empty] | CABLE |  | SERIAL BD
26 | RECT | PIN |  | HR152-ND | [8 empty] | CONN | [17 empty] | CONN | [8 empty] | CONN |  | DB25'S | [5 empty] | CONN |  | CONN
12 | PLUG | SOCKET |  | HR109-ND | [5 empty] | CD7WFW2000X
12 | RECT | PIN |  | HR207-ND | [5 empty] | MOTOR |  |  | 4 [RXD] | 4 | ROTATR | RXD_ROTATR | RED | 5 | 5 | [6 empty] | RXD_ROTATR |  |  |  |  | 21 | JH1-21 | G | G |  | RXD_ROTATR | G | G | JH1-21 | 21 | RXD_ROTATR | 2 | [5 empty] | ? |  | ? | TTYC0
12 | PLUG | PIN |  | HR107-ND | [5 empty] | MAIN |  |  | 3 [TXD] | 3 |  | TXD_ROTATR | ORG | 6 | 6 | [6 empty] | TXD_ROTATR |  |  |  |  | 22 | JH1-22 | c | c |  | TXD_ROTATR | c | c | JH1-22 | 22 | TXD_ROTATR | 3 | [5 empty] | ? |  | ? | TTYC0
12 | RECT | SOCKET |  | HR209-ND | [5 empty] | CONN |  |  | 5 [GND RS232] | 5 |  | DGND_ROTATR | YEL | 7 | 7 | [6 empty] | DGND_ROTATR |  |  |  |  | 23 | JH1-23 | u | u |  | DGND_ROTATR | u | u | JH1-23 | 23 | DGND_ROTATR | 7 | [5 empty] | ? |  | ? | TTYC0
10 | PLUG | SOCKET |  | HR108-ND | [37 empty] | PORT2 | [5 empty] | ? |  | ?
10 | RECT | PIN |  | HR106-ND | [5 empty] | 4 [RXD] | 4 | GRTTRT |  |  |  | RXD_GRTTRT | RED | 8 | 8 | [6 empty] | RXD_GRTTRT |  |  |  |  | 24 | JH1-24 | DD | DD |  | RXD_GRTTRT | DD | DD | JH1-24 | 24 | RXD_GRTTRT | 2 | [5 empty] | ? |  | ? | TTYC1
[10 empty] | 3 [TXD] | 3 |  |  |  |  | TXD_GRTTRT | ORG | 9 | 9 | [6 empty] | TXD_GRTTRT |  |  |  |  | 25 | JH1-25 | d | d |  | TXD_GRTTRT | d | d | JH1-25 | 25 | TXD_GRTTRT | 3 | [5 empty] | ? |  | ? | TTYC1
[10 empty] | 5 [GND RS232] | 5 |  |  |  |  | DGND_GRTTRT | YEL | 10 | 10 | [6 empty] | DGND_GRTTRT |  |  |  |  | 26 | JH1-26 | H | H |  | DGND_GRTTRT | H | H | JH1-26 | 26 | DGND_GRTTRT | 7 | [5 empty] | ? |  | ? | TTYC1
[31 empty] | N.C. | JH1-28,29 | J | J |  |  | J | J | JH1-28,29 | 27 | [7 empty] | ? |  | ?
SERVO
[31 empty] | JH2 |  | JBOX | JBOX | 15FT |  | JBOX | UTILITY PC |  | JH2 | [7 empty] | P2 |  | ANIMATICS
[31 empty] | 30 PIN | STEPPER | STEPPER | STEPPER | WIRE |  | STEPPER | STEPPER | STEPPER | 30 PIN | [7 empty] | POWER |  | PWR SUPPLY
[31 empty] | AMP | INTRF BD | 55 PIN | CABLE | COLOR/ |  | CABLE | CABLE | INTRF BD | AMP | [7 empty] | PLUG |  | PS42V6A
[31 empty] | CONN | PIN # | CONN | 55 CONN | AWG |  | 55 CONN | 55 CONN | PIN # | CONN | [7 empty] | CONN |  | 4P CONN
[10 empty] | 1A [20-48VDC] | 1A | [5 empty] | WHT | 11 | 11 | ROTATR+42V | [5 empty] | ROTATR+42V |  |  |  |  | 1 | JH2-1 | n | n |  | ROTATR+42V | n | n | JH2-1 | 1 | 42VDC | 22AWG |  |  |  | WHT |  | 1 |  | 1 | 42VDC
[10 empty] | 2A [GND] | 2A | [5 empty] | BLK | 12 | 12 | ROTATR_GND | [5 empty] | ROTATR_GND |  |  |  |  | 2 | JH2-2 | U | U |  | ROTATR_GND | U | U | JH2-2 | 2 | GND | 22AWG |  |  |  | BLK |  | 2 |  | 2 | GND
[31 empty] | 3 | JH2-3 | T | T |  | ROTATR+42V | T | T | JH2-3 | 3 | 42VDC | 22AWG |  |  |  | WHT
[31 empty] | 4 | JH2-4 | AA | AA |  | ROTATR_GND | AA | AA | JH2-4 | 4 | GND | 22AWG |  |  |  | BLK
[13 empty] | 1A [20-48VDC] | 1A |  |  | WHT | 13 | 13 | GRTTRT+42V | [5 empty] | GRTTRT+42V |  |  |  |  | 5 | JH2-5 | z | z |  | GRTTRT+42V | z | z | JH2-5 | 5 | 42VDC | 22AWG |  |  |  | WHT
[13 empty] | 2A [GND] | 2A |  |  | BLK | 14 | 14 | GRTTRT_GND | [5 empty] | GRTTRT_GND |  |  |  |  | 6 | JH2-6 | m | m |  | GRTTRT_GND | m | m | JH2-6 | 6 | GND | 22AWG |  |  |  | BLK
[31 empty] | 7 | JH2-7 | k | k |  | GRTTRT+42V | k | k | JH2-7 | 7 | 42VDC | 22AWG |  |  |  | WHT
[31 empty] | 8 | JH2-8 | S | S |  | GRTTRT_GND | S | S | JH2-8 | 8 | GND | 22AWG |  |  |  | BLK
STEPPER MOTORS | [7 empty] | M2-2232-S | [8 empty] | ~2FT |  | HR152-ND |  |  |  | CGRID 10PIN | CGRID 10PIN
[8 empty] | IMS | [7 empty] | LOCAL | WIRE | STEPPER | JBOX |  |  |  | MALE | FEMALE | [6 empty] | JH1
[8 empty] | STEPPER |  |  |  | WIRES | 6 PIN | STEPPER |  | SIGNAL | COLOR | MOTOR | STEPPER |  |  |  | 10 PIN | 10 PIN | [6 empty] | 30 PIN | [8 empty] | JH1 |  | IMS483-34P1 | LOCAL |  | 42VDC | [7 empty] | IMS483-34P1 |  |  |  | JH1
[8 empty] | MOTORS |  |  |  | COLOR(NAME) | STEPPER | 6 CONN |  | NAME | 22AWG | 26 PIN | 26 PIN |  |  |  | CGRID | CGRID | [6 empty] | IDC | [8 empty] | 30 PIN |  | DRIVER | SIGNAL |  | TERM | [7 empty] | DRIVER | CABLE |  |  | 30 PIN | PASSTHRU | PASSTHRU | PASSTHRU
[13 empty] | CONN |  |  |  |  | CABLE | CONN |  |  |  | CONN | CONN | [6 empty] | CONN | [8 empty] | AMP |  | P2 TERM | NAME |  | STRIP | [7 empty] | P1 34PIN HDR | CONN |  |  | IDC | INTRF BD | 55 PIN | CABLE
[40 empty] | CONN |  | STRIP | 14AWG | [6 empty] | STEPPER1 |  |  | NONSTANDARD |  |  |  | CONN | PIN # | CONN | 55 CONN
[8 empty] | DICHROIC | 2 | WHT+RED | 1 | YEL(A) | A | A | DICTRT | DICTRT_A+ | BLK | 1 | 1 | [5 empty] | DICTRT_A+ | [5 empty] | 1 | JH1-1 | p | p |  | DICTRT_A+ | p | p | JH1-1 | 1 | DICTRT_A+ | 8_TOP | PHASE A | [5 empty] | P2 | ANIMATICS |  |  | 4 | 4 | STEP | XSTEP | 2 | JH1-2 | BB | BB | XSTEP | 62 | 62 | XSTEP | DICHROIC
[9 empty] | 6 | WHT/YEL | 3 | RED(A!) | B | B |  | DICTRT_A- | BRN | 2 | 2 | [5 empty] | DICTRT_A- | [5 empty] | 2 | JH1-2 | BB | BB |  | DICTRT_A- | BB | BB | JH1-2 | 2 | DICTRT_A- | 7_TOP | PHASE A! | [5 empty] | POWER | PWR SUPPLY |  |  | 6 | 6 | DIR | XDIR | 4 | JH1-4 | V | V | XDIR | 12 | 12 | XDIR
[9 empty] | 7 | WHT/BLK | 4 | ORG(B) | C | C |  | DICTRT_B+ | RED | 3 | 3 | [5 empty] | DICTRT_B+ | [5 empty] | 3 | JH1-3 | q | q |  | DICTRT_B+ | q | q | JH1-3 | 3 | DICTRT_B+ | 6_TOP | PHASE B | [5 empty] | PLUG | PS42V6A |  |  | 8 | 8 | OPTO_PS | +5VDC | 6 | JH1-6 | W | W | +5VDC | 51 | 51 | +5VDC
[9 empty] | 8 | WHT/ORG | 5 | BLK(B!) | D | D |  | DICTRT_B- | ORG | 4 | 4 | [5 empty] | DICTRT_B- | [5 empty] | 4 | JH1-4 | V | V |  | DICTRT_B- | V | V | JH1-4 | 4 | DICTRT_B- | 5_TOP | PHASE B! | [5 empty] | CONN | 4P CONN |  |  | 10 | 10 | ENABLE | XAUX | 8 | JH1-8 | B | B | XAUX | 13 | 13 | XAUX
[42 empty] | 4_TOP | V+(42V) |  |  |  | 42VDC |  | 1 | 1 |  |  | 27 |  | GND
[42 empty] | 3_TOP | GND |  |  |  | GND |  | 2 | 2
[8 empty] | ORDER | 2 | WHT+RED | 1 | YEL(A) | A | A | OSFLTR | OSFLTR_A+ | BLK | 5 | 5 | [5 empty] | OSFLTR_A+ | [5 empty] | 5 | JH1-5 | A | A |  | OSFLTR_A+ | A | A | JH1-5 | 5 | OSFLTR_A+ | 8_TOP | PHASE A | [9 empty] | 4 | 4 | STEP | YSTEP | 10 | JH1-10 | C | C | YSTEP | 66 | 66 | YSTEP | ORDER
[8 empty] | SORTER | 6 | WHT/YEL | 3 | RED(A!) | B | B |  | OSFLTR_A- | BRN | 6 | 6 | [5 empty] | OSFLTR_A- | [5 empty] | 6 | JH1-6 | W | W |  | OSFLTR_A- | W | W | JH1-6 | 6 | OSFLTR_A- | 7_TOP | PHASE A! | [9 empty] | 6 | 6 | DIR | YDIR | 12 | JH1-12 | r | r | YDIR | 17 | 17 | YDIR | SORTER
[9 empty] | 7 | WHT/BLK | 4 | ORG(B) | C | C |  | OSFLTR_B+ | RED | 7 | 7 | [5 empty] | OSFLTR_B+ | [5 empty] | 7 | JH1-7 | X | X |  | OSFLTR_B+ | X | X | JH1-7 | 7 | OSFLTR_B+ | 6_TOP | PHASE B | [9 empty] | 8 | 8 | OPTO_PS | +5VDC | 14 | JH1-14 | D | D | +5VDC | 59 | 59 | +5VDC
[9 empty] | 8 | WHT/ORG | 5 | BLK(B!) | D | D |  | OSFLTR_B- | ORG | 8 | 8 | [5 empty] | OSFLTR_B- | [5 empty] | 8 | JH1-8 | B | B |  | OSFLTR_B- | B | B | JH1-8 | 8 | OSFLTR_B- | 5_TOP | PHASE B! | [9 empty] | 10 | 10 | ENABLE | YAUX | 16 | JH1-16 | s | s | YAUX | 18 | 18 | YAUX
[23 empty] | CGRID 10PIN | CGRID 10PIN | [17 empty] | 4_TOP | V+(42V) | [9 empty] | 27 |  | GND
[23 empty] | FEMALE | MALE | [17 empty] | 3_TOP | GND
[8 empty] | SLIT | 2 | WHT+RED | 1 | YEL(A) | A | A | SLTWHL | SLTWHL_A+ | BLK | 9 | 9 | [5 empty] | SLTWHL_A+ | [5 empty] | 9 | JH1-9 | Y | Y |  | SLTWHL_A+ | Y | Y | JH1-9 | 9 | SLTWHL_A+ | 8_TOP | PHASE A | [9 empty] | 4 | 4 | STEP | ZSTEP | 18 | JH1-18 | a | a | ZSTEP | 72 | 72 | ZSTEP | SLIT
[8 empty] | WHEEL | 6 | WHT/YEL | 3 | RED(A!) | B | B |  | SLTWHL_A- | BRN | 10 | 10 | [5 empty] | SLTWHL_A- | [5 empty] | 10 | JH1-10 | C | C |  | SLTWHL_A- | C | C | JH1-10 | 10 | SLTWHL_A- | 7_TOP | PHASE A! | [9 empty] | 6 | 6 | DIR | ZDIR | 20 | JH1-20 | t | t | ZDIR | 23 | 23 | ZDIR | WHEEL
[9 empty] | 7 | WHT/BLK | 4 | ORG(B) | C | C |  | SLTWHL_B+ | RED | 11 | 11 | [5 empty] | SLTWHL_B+ | [5 empty] | 11 | JH1-11 | CC | CC |  | SLTWHL_B+ | CC | CC | JH1-11 | 11 | SLTWHL_B+ | 6_TOP | PHASE B | [9 empty] | 8 | 8 | OPTO_PS | +5VDC | 22 | JH1-22 | c | c | +5VDC | 69 | 69 | +5VDC
[9 empty] | 8 | WHT/ORG | 5 | BLK(B!) | D | D |  | SLTWHL_B- | ORG | 12 | 12 | [5 empty] | SLTWHL_B- | [5 empty] | 12 | JH1-12 | r | r |  | SLTWHL_B- | r | r | JH1-12 | 12 | SLTWHL_B- | 5_TOP | PHASE B! | [9 empty] | 10 | 10 | ENABLE | ZAUX | 24 | JH1-24 | DD | DD | ZAUX | 24 | 24 | ZAUX
[42 empty] | 4_TOP | V+(42V) | [9 empty] | 27 |  |  |  | JH2
[42 empty] | 3_TOP | GND | [6 empty] | STEPPER2 | [6 empty] | 30P IDC
[8 empty] | GUIDER | 2 | WHT+RED | 1 | YEL(A) | A | A | GFLWHL | GFLWHL_A+ | BLK | 13 | 13 | [5 empty] | GFLWHL_A+ | [5 empty] | 13 | JH1-13 | Z | Z |  | GFLWHL_A+ | Z | Z | JH1-13 | 13 | GFLWHL_A+ | 8_BOTTOM | PHASE A | [5 empty] | P2 | ANIMATICS |  |  | 4 | 4 | STEP | TSTEP | 2 | JH2-2 | U | U | TSTEP | 76 | 76 | TSTEP | SLIT
[8 empty] | FILTER | 6 | WHT/YEL | 3 | RED(A!) | B | B |  | GFLWHL_A- | BRN | 14 | 14 | [5 empty] | GFLWHL_A- | [5 empty] | 14 | JH1-14 | D | D |  | GFLWHL_A- | D | D | JH1-14 | 14 | GFLWHL_A- | 7_BOTTOM | PHASE A! | [5 empty] | POWER | PWR SUPPLY |  |  | 6 | 6 | DIR | TDIR | 4 | JH2-4 | AA | AA | TDIR | 28 | 28 | TDIR | VIEWER
[8 empty] | WHEEL | 7 | WHT/BLK | 4 | ORG(B) | C | C |  | GFLWHL_B+ | RED | 15 | 15 | [5 empty] | GFLWHL_B+ | [5 empty] | 15 | JH1-15 | E | E |  | GFLWHL_B+ | E | E | JH1-15 | 15 | GFLWHL_B+ | 6_BOTTOM | PHASE B | [5 empty] | PLUG | PS42V6A |  |  | 8 | 8 | OPTO_PS | +5VDC | 6 | JH2-6 | m | m | +5VDC | 80 | 80 | +5VDC | FILTER
[9 empty] | 8 | WHT/ORG | 5 | BLK(B!) | D | D |  | GFLWHL_B- | ORG | 16 | 16 | [5 empty] | GFLWHL_B- | [5 empty] | 16 | JH1-16 | s | s |  | GFLWHL_B- | s | s | JH1-16 | 16 | GFLWHL_B- | 5_BOTTOM | PHASE B! | [5 empty] | CONN | 4P CONN |  |  | 10 | 10 | ENABLE | TAUX | 8 | JH2-8 | S | S | TAUX | 77 | 77 | TAUX | WHEEL
[23 empty] | CGRID 10PIN | CGRID 10PIN | [17 empty] | 4_BOTTOM | V+(42V) |  |  |  | 42VDC |  | 1 | 1 |  |  | 27 |  | GND
[23 empty] | FEMALE | MALE | [17 empty] | 3_BOTTOM | GND |  |  |  | GND |  | 2 | 2
[8 empty] | SPEC | 2 | WHT+RED | 1 | YEL(A) | A | A | SFOCUS | SFOCUS_A+ | BLK | 17 | 17 | [5 empty] | SFOCUS_A+ | [5 empty] | 17 | JH1-17 | F | F |  | SFOCUS_A+ | F | F | JH1-17 | 17 | SFOCUS_A+ | 8_BOTTOM | PHASE A | [9 empty] | 4 | 4 | STEP | USTEP | 10 | JH2-10 | j | j | USTEP | 83 | 83 | USTEP | SPEC
[8 empty] | ARRAY | 6 | WHT/YEL | 3 | RED(A!) | B | B |  | SFOCUS_A- | BRN | 18 | 18 | [5 empty] | SFOCUS_A- | [5 empty] | 18 | JH1-18 | a | a |  | SFOCUS_A- | a | a | JH1-18 | 18 | SFOCUS_A- | 7_BOTTOM | PHASE A! | [9 empty] | 6 | 6 | DIR | UDIR | 12 | JH2-12 | R | R | UDIR | 33 | 33 | UDIR | ARRAY
[8 empty] | FOCUS | 7 | WHT/BLK | 4 | ORG(B) | C | C |  | SFOCUS_B+ | RED | 19 | 19 | [5 empty] | SFOCUS_B+ | [5 empty] | 19 | JH1-19 | b | b |  | SFOCUS_B+ | b | b | JH1-19 | 19 | SFOCUS_B+ | 6_BOTTOM | PHASE B | [9 empty] | 8 | 8 | OPTO_PS | +5VDC | 14 | JH2-14 | y | y | +5VDC | 90 | 90 | +5VDC | FOCUS
[9 empty] | 8 | WHT/ORG | 5 | BLK(B!) | D | D |  | SFOCUS_B- | ORG | 20 | 20 | [5 empty] | SFOCUS_B- | [5 empty] | 20 | JH1-20 | t | t |  | SFOCUS_B- | t | t | JH1-20 | 20 | SFOCUS_B- | 5_BOTTOM | PHASE B! | [9 empty] | 10 | 10 | ENABLE | UAUX | 16 | JH2-16 | h | h | UAUX | 34 | 34 | UAUX
[42 empty] | 4_BOTTOM | V+(42V) | [9 empty] | 27 |  | GND
[8 empty] | 6PIN CGRID-M | 6PIN CGRID-F | [32 empty] | 3_BOTTOM | GND
CAL BOX |  |  | CAL |  |  | BLU | BLU | 1 | 1 | BLK(A) |  | (1)BLK | A | A | CBM | FPM_A+ | BLK | 21 | 21 | [5 empty] | FPM_A+ | [5 empty] | 21 | JH1-21 | G | G |  | FPM_A+ | G | G | JH1-21 | 21 | FPM_A+ | 8_BOTTOM | PHASE A | [9 empty] | 4 | 4 | STEP | VSTEP | 18 | JH2-18 | FF | FF | VSTEP | 87 | 87 | VSTEP | FOCAL
|  |  | BOX |  |  | WHT | WHT | 2 (NC) | 2 (NC) |  |  | (3)GRN | B | B |  | FPM_A- | BRN | 22 | 22 | [5 empty] | FPM_A- | [5 empty] | 22 | JH1-22 | c | c |  | FPM_A- | c | c | JH1-22 | 22 | FPM_A- | 7_BOTTOM | PHASE A! | [9 empty] | 6 | 6 | DIR | VDIR | 20 | JH2-20 | g | g | VDIR | 38 | 38 | VDIR | PLANE
|  |  | MIRROR |  |  | RED | RED | 3 | 3 | GRN(B) |  | (4)BRN | C | C |  | FPM_B+ | RED | 23 | 23 | [5 empty] | FPM_B+ | [5 empty] | 23 | JH1-23 | u | u |  | FPM_B+ | u | u | JH1-23 | 23 | FPM_B+ | 6_BOTTOM | PHASE B | [9 empty] | 8 | 8 | OPTO_PS | +5VDC | 22 | JH2-22 | w | w |  |  |  |  | MASK
|  |  | MOTOR |  |  | BLK | BLK | 4 | 4 | BRN(C) |  | (6)WHT | D | D |  | FPM_B- | ORG | 24 | 24 | [5 empty] | FPM_B- | [5 empty] | 24 | JH1-24 | DD | DD |  | FPM_B- | DD | DD | JH1-24 | 24 | FPM_B- | 5_BOTTOM | PHASE B! | [9 empty] | 10 | 10 | ENABLE | VAUX | 24 | JH2-24 | K | K | VAUX | 39 | 39 | VAUX
|  |  | MOTOR |  |  | WHT | WHT | 5 (NC) | 5 (NC) | [8 empty] | A | A | [11 empty] | 25 | JH1-25 | d | d |  |  | d | d | JH1-25 | 25 |  | 4_BOTTOM | V+(42V) | [9 empty] | 27 |  | GND
[6 empty] | GRN | GRN | 6 | 6 | WHT(D) | [7 empty] | B | B | [11 empty] | 26 | JH1-26 | H | H |  |  | H | H | JH1-26 | 26 |  | 3_BOTTOM | GND
[18 empty] | CALBOX | CALBOX | [11 empty] | 28,30 | JH1-28,30 | J | J |  |  | J | J | JH1-28,30 | 28,30 | [13 empty] | AMP 34P
[8 empty] | 4PIN CGRID-F | 4PIN CGRID-M |  |  |  | PT02A-1098S | PT06A-1098P |  |  |  | POS CBL | POS CBL | [6 empty] | CGRID 4PIN |  |  |  |  | JH2 | [8 empty] | JH2 | [13 empty] | CONN
|  | CAL BOX MIRROR |  |  | 5V | RED |  | 1 | 1 | BLK(A) | 5V | (1)BLK | A | A | CBPS | 5V | BLK | A | A | [5 empty] | 5V | 1 |  |  |  |  | 25 | JH2-25 | v | v |  | CBM_A+ | v | v | JH2-25 | 25 | CBM_A+ | 8_BOTTOM | PHASE A | [9 empty] | 4 | 4 | STEP | RSTEP | 25 | JH2-25 | v | v | RSTEP | 93 | 93 | RSTEP | CAL
|  | POSITION SENSORS |  |  | DGND | BLACK |  | 2 | 2 | RED(F) |  | (3)GRN | B | B |  | FPM+ | BRN | B | B | [5 empty] | FPM+ | 2 |  |  |  |  | 26 | JH2-26 | e | e |  | CBM_A- | e | e | JH2-26 | 26 | CBM_A- | 7_BOTTOM | PHASE A! | [9 empty] | 6 | 6 | DIR | RDIR | 26 | JH2-26 | e | e | RDIR | 44 | 44 | RDIR | BOX
|  | (NEAT STAGE) |  |  | FPM+ | BLUE |  | 3 | 3 | GRN(B) | FPM+ |  | C | C |  | FPM- | RED | C | C | [5 empty] | FPM- | 3 |  |  |  |  | 27 | JH2-27 | EE | EE |  | CBM_B+ | EE | EE | JH2-27 | 27 | CBM_B+ | 6_BOTTOM | PHASE B | [9 empty] | 8 | 8 | OPTO_PS | +5VDC | 27 | JH2-27 | EE | EE |  |  |  |  | MIRROR
[5 empty] | FPM- | BLACK |  | 4 | 4 | WHT(E) | FPM- |  | D | D |  | DGND | ORG | D | D | [5 empty] | DGND | 4 |  |  |  |  | 28 | JH2-28 | HH | HH |  | CBM_B- | HH | HH | JH2-28 | 28 | CBM_B- | 5_BOTTOM | PHASE B! | [9 empty] | 10 | 10 | ENABLE | RAUX | 28 | JH2-28 | HH | HH | RAUX | 45 | 45 | RAUX
[12 empty] | (4)WHT | E | E |  |  |  | E | E | [6 empty] | (*SEE ABOVE*) | [15 empty] | 4_BOTTOM | V+(42V) | [9 empty] | 27 |  | GND
[11 empty] | DGND | (2)RED | F | F |  |  |  | F | F | [22 empty] | 3_BOTTOM | GND
TEMPERATURE SENSORS |  |  |  | .1X1X10 | .1X1X10 |  | NON-STD | [40 empty] | DIN6-REAR PNL
FUNCTION |  |  |  | TEMP | TEMP |  | .1X2X5 | .1X2X5 |  | .1X1X8 | .1X1X8 |  | JH1 | 55 PIN | 55 PIN | TEMP | GLENAIR | [19 empty] | TEMP | TEMP | 55 PIN | [8 empty] | SPEC
| CRYO | WIRE | LOCAL | 10PIN | 10CON | ? INCH | 16PIN HDR | 16PIN CON | ? INCH | 8PIN | 8CON | ? INCH | 30 PIN | HERMETIC | HERMETIC | 55 PIN | 55COND | [17 empty] | GLENAIR |  | 55 PIN | 55 PIN | HERMETIC | [8 empty] | LAKESHORE
LAKESHORE | NUM/ | COLOR | SIGNAL | CGRID | CGRID | MANGANIN | FEED- | FEED- | MANGANIN | CGRID | CGRID | MANGANIN | IDC | INTRF BD | TEMP | CABLE | CABLE | [17 empty] | 55COND |  | CABLE |  | INTRF BD | [8 empty] | 330
DT471-BR | NAME | CODE | NAME | CONN | CONN | CABLE | [41 empty] | SENSOR
CONN
SPECTROGRAPH | SPEC_TMP | BLK/BLU | I+ | 3 | 3 |  | 1 | 2 |  | 1 | 1 | 1 | 1 | JH1-1 | p | p | [20 empty] | p | p | JH1-1 |  | SPEC_TMP | [6 empty] | 5 [I+] CHNL A
ARRAY | SPEC_TMP | YEL/CLR | V+ | 4 | 4 |  | 2 | 1 |  | 2 | 2 | 2 | 2 | JH1-2 | BB | BB | [20 empty] | BB | BB | JH1-2 |  | SPEC_TMP | [6 empty] | 4 [V+] CHNL A
TEMP | SPEC_TMP | RED | I- | 5 | 5 |  | 3 | 4 |  | 3 | 3 | 3 | 3 | JH1-3 | q | q | [20 empty] | q | q | JH1-3 |  | SPEC_TMP | [6 empty] | 1 [I-] CHNL A
SENSOR | SPEC_TMP | GRN | V- | 6 | 6 |  | 4 | 3 |  | 4 | 4 | 4 | 4 | JH1-4 | V | V | [20 empty] | V | V | JH1-4 |  | SPEC_TMP | [6 empty] | 2 [V-] CHNL A
[42 empty] | 330 |  |  |  | CK05 .1uF Cer Cap
FILTER
SPECTROGRAPH | SPEC_HTR |  | R3¦¦R4+ | 7 | 7 |  | 5 | 6 |  | 5 | 5 | 5 | 5 | JH1-5 | A | A | [20 empty] | A | A | JH1-5 |  | SPEC_HTR | HI | [5 empty] | HI
ARRAY | SPEC_HTR |  | R3¦¦R4- | 8 | 8 |  | 6 | 5 |  | 6 | 6 | 6 | 6 | JH1-6 | W | W | [20 empty] | W | W | JH1-6 |  | SPEC_HTR
HEATER | SPEC_HTR |  | R1¦¦R2+ | 9 | 9 |  | 7 | 8 |  | 7 | 7 | 7 | 7 | JH1-7 | X | X | [20 empty] | X | X | JH1-7 |  | SPEC_HTR | LO | [5 empty] | LO
RESISTORS | SPEC_HTR |  | R1¦¦R2- | 10 | 10 |  | 8 | 7 |  | 8 | 8 | 8 | 8 | JH1-8 | B | B | [20 empty] | B | B | JH1-8 |  | SPEC_HTR | GND | [5 empty] | GND
|  |  |  | 4PIN | 4PIN
|  |  |  | CGRID | CGRID
| STG2_TMP | BLK/BLU | I+ | 1 | 1 | [6 empty] | 9 | 9 | JH1-9 | Y | Y | [20 empty] | Y | Y | JH1-9 |  | STG2_TMP | [6 empty] | 5 [I+] CHNL B
CCC 2ND STG | STG2_TMP | YEL/CLR | V+ | 2 | 2 | [6 empty] | 10 | 10 | JH1-10 | C | C | [20 empty] | C | C | JH1-10 |  | STG2_TMP | [6 empty] | 4 [V+] CHNL B
TEMP | STG2_TMP | RED | I- | 3 | 3 | [6 empty] | 11 | 11 | JH1-11 | CC | CC | [20 empty] | CC | CC | JH1-11 |  | STG2_TMP | [6 empty] | 1 [I-] CHNL B
SENSOR | STG2_TMP | GRN | V- | 4 | 4 | [6 empty] | 12 | 12 | JH1-12 | r | r | [20 empty] | r | r | JH1-12 |  | STG2_TMP | [6 empty] | 2 [V-] CHNL B
|  |  |  | 10PIN | 10PIN |  | 16PIN | 16 PIN | [37 empty] | CK05 .1uF Cer Cap |  | GUIDER
|  |  |  | CGRID | CGRID |  | FD THRU | FD THRU | [39 empty] | LAKESHORE
GUIDER | GUID_TMP | BLK/BLU | I+ | 3 | 3 |  | 1 | 2 |  |  |  | 13 | 13 | JH1-13 | Z | Z | [20 empty] | Z | Z | JH1-13 |  | GUID_TMP | [6 empty] | 5 [I+] CHNL A
ARRAY | GUID_TMP | YEL/CLR | V+ | 4 | 4 |  | 2 | 1 |  |  |  | 14 | 14 | JH1-14 | D | D | [20 empty] | D | D | JH1-14 |  | GUID_TMP | [6 empty] | 4 [V+] CHNL A
TEMP | GUID_TMP | RED | I- | 5 | 5 |  | 3 | 4 |  |  |  | 15 | 15 | JH1-15 | E | E | [20 empty] | E | E | JH1-15 |  | GUID_TMP | [6 empty] | 1 [I-] CHNL A
SENSOR | GUID_TMP | GRN | V- | 6 | 6 |  | 4 | 3 |  |  |  | 16 | 16 | JH1-16 | s | s | [20 empty] | s | s | JH1-16 |  | GUID_TMP | [6 empty] | 2 [V-] CHNL A
[42 empty] | 330 |  |  |  | CK05 .1uF Cer Cap
FILTER
GUIDER | GUID_HTR |  | R3¦¦R4+ | 7 | 7 |  | 5 | 6 |  |  |  | 17 | 17 | JH1-17 | F | F | [20 empty] | F | F | JH1-17 |  | GUID_HTR | HI | [5 empty] | HI
ARRAY | GUID_HTR |  | R3¦¦R4- | 8 | 8 |  | 6 | 5 |  |  |  | 18 | 18 | JH1-18 | a | a | [20 empty] | a | a | JH1-18 |  | GUID_HTR
HEATER | GUID_HTR |  | R1¦¦R2+ | 9 | 9 |  | 7 | 8 |  |  |  | 19 | 19 | JH1-19 | b | b | [20 empty] | b | b | JH1-19 |  | GUID_HTR | LO | [5 empty] | LO
RESISTORS | GUID_HTR |  | R1¦¦R2- | 10 | 10 |  | 8 | 7 |  |  |  | 20 | 20 | JH1-20 | t | t | [20 empty] | t | t | JH1-20 |  | GUID_HTR | GND | [5 empty] | GND
|  |  |  | 4PIN | 4PIN
|  |  |  | CGRID | CGRID
| LN2_TMP | BLK/BLU | I+ | 1 | 1 | [6 empty] | 21 | 21 | JH1-21 | G | G | [20 empty] | G | G | JH1-21 |  | LN2_TMP | [6 empty] | 5 [I+] CHNL B
LN2 CAN | LN2_TMP | YEL/CLR | V+ | 2 | 2 | [6 empty] | 22 | 22 | JH1-22 | c | c | [20 empty] | c | c | JH1-22 |  | LN2_TMP | [6 empty] | 4 [V+] CHNL B
TEMP | LN2_TMP | RED | I- | 3 | 3 | [6 empty] | 23 | 23 | JH1-23 | u | u | [20 empty] | u | u | JH1-23 |  | LN2_TMP | [6 empty] | 1 [I-] CHNL B
SENSOR | LN2_TMP | GRN | V- | 4 | 4 | [6 empty] | 24 | 24 | JH1-24 | DD | DD | [20 empty] | DD | DD | JH1-24 |  | LN2_TMP | [6 empty] | 2 [V-] CHNL B
CK05 .1uF Cer Cap
UTILITY PC (LITTLEDOG)
Non standard connector pinouts in this section. Need to mark each connector with pin 1 and center key

## temp sensors
CRYOSTAT ASSEMBLIES | [5 empty] | .1X2X7 | .1X2X7 | [9 empty] | HR109-ND | DIN6-REAR PNL | BANANA |  | RJ45 | RJ45 |  | RJ45 | RJ45
FUNCTION | CRYO |  |  | MOLEX | 9 INCH | 18PIN HDR | 18PIN HDR | 22 INCH | 41 PIN | UTILITY | [6 empty] | TEMP | LAKESHORE | LAKESHORE |  | LAKESHORE | TEMP | WIRE
| NUMBER | WIRE | LOCAL | CGRID | COLD | COLD | COLD | COLD | HERMETIC | 41/26 PIN | WIRE |  |  |  |  | WIRE | 12 PIN | 330 | 330 |  | 330 | SERIAL | COLOR/ | LANTRONICS | LANTRONICS | LOCAL
| OR | COLOR | SIGNAL | 10 PIN | *(CONST)ANTAN | FEED- | FEED- | MANGANIN | CONN | CABLE | COLOR/ |  |  |  |  | COLOR/ | CABLE | SENSOR | FILTER |  | RS232 | COMM | AWG | SERVER | SERVER | SIGNAL
| NAME | CODE | NAME | CONN | CABLE | THROUGH | THROUGH | CABLE | PIN # | 41 CONN | AWG |  |  |  |  | AWG | CONN | CONN | CONN |  | CONN | CONN |  | CONN | PIN | NAME
|  |  |  | 1 | [20 empty] | 3 | TX+ (OUT)
|  |  |  | 2 | [20 empty] | 6 | RX+(IN)
ARRAY TEMP | 1 / ARRAY | BLK/BLU | I+ | 3 | *10mil CONST | 1 | 2 | 2 | A | A |  | 1 |  |  | 1 |  | 1 | 5 [I+] CHNL A | [6 empty] | 4 | TX-
SENSOR | 1 / ARRAY | YEL/CLR | V+ | 4 | *10mil CONST | 2 | 1 | 1 | B | B |  | 2 |  |  | 2 |  | 2 | 4 [V+] CHNL A | [6 empty] | 5 | RX-
| 1 / ARRAY | RED | I- | 5 | *10mil CONST | 3 | 4 | 4 | C | C |  | 3 |  |  | 3 |  | 3 | 1 [I-] CHNL A
| 1 / ARRAY | GRN | V- | 6 | *10mil CONST | 4 | 3 | 3 | D | D |  | 4 |  |  | 4 |  | 4 | 2 [V-] CHNL A
ARRAY | R3¦¦R4+ | N/A | R3¦¦R4+ | 7 | Cu | 15 | 16 | 16 | E | E |  | 5 |  |  | 5 |  | 5 | HI | HI
HEATER | R3¦¦R4- | N/A | R3¦¦R4- | 8 | Cu | 16 | 15 | 15 | F | F |  | 6 |  |  | 6 |  | 6
RESISTORS | R1¦¦R2+ | N/A | R1¦¦R2+ | 9 | Cu | 17 | 18 | 18 | G | G |  | 7 |  |  | 7 |  | 7
| R1¦¦R2- | N/A | R1¦¦R2- | 10 | Cu | 18 | 17 | 17 | H | H |  | 8 |  |  | 8 |  | 8 | LO | LO
[6 empty] | .1X2X7 | .1X2X7 |  |  |  |  | N.C. |  |  | N.C. |  | SHIELD | GND | GND
[6 empty] | 14PIN HDR | 14PIN HDR | 5 INCH
[6 empty] | COLD | COLD | COLD
[6 empty] | FEED- | FEED- | MANGANIN | [11 empty] | LAKESHORE 330
[6 empty] | THROUGH | THROUGH | CABLE | [11 empty] | TEMP CONTROLLER
COLD | 1 /CW3 | BLK/BLU | I+ | - | - | 1 | 2 | 2 | J | J |  | 9 | 9 | 9 | 9 |  | 9 | 5 [I+] CHNL B
STRUCTURE | 1 /CW3 | YEL/CLR | V+ | - | - | 2 | 1 | 1 | K | K |  | 10 | 10 | 10 | 10 |  | 10 | 4 [V+] CHNL B
TEMP SENSOR | 1 /CW3 | RED | I- | - | - | 3 | 4 | 4 | L | L |  | 11 | 11 | 11 | 11 |  | 11 | 1 [I-] CHNL B
| 1 /CW3 | GRN | V- | - | - | 4 | 3 | 3 | M | M |  | 12 | 12 | 12 | 12 |  | 12 | 2 [V-] CHNL B
[12 empty] | HR-145-ND | HR152-ND | HR152-ND | HR-145-ND |  | MULTIPLE | CGRID 7PIN
[12 empty] | UTILITY | JBOX | JBOX | MOTOR |  | UPPER FILTER | UPPER FILTER
[12 empty] | 41/26 PIN | UTILITY | MOTOR | 26 PIN/SPLIT | WIRE | 26 PIN/SPLIT | WHEEL MOTOR | MULTIPLE | CGRID 7PIN
[12 empty] | CABLE | 26 PIN | 26 PIN | CABLE | COLOR/ | CABLE | 7 PIN I/O | LOWER FILTER | LOWER FILTER | CGRID 7PIN | CD7WFW2000X
[12 empty] | 26 CONN | CONN | CONN | 26 CONN | AWG | 7 CONN | CONN | 26 PIN/SPLIT | WHEEL MOTOR | PUPIL IMAGER | PUPIL IMAGER
UPPER FILTER | FW1_NC | ORG | N.C. | - | - | 9 | 10 | 10 | N | N |  | 13 | 13 | 1 | 1 |  |  |  | 7 PIN I/O | 7 PIN I/O | 26 PIN/SPLIT | MOTOR
WHEEL | FW1_C | BLK | C. | - | - | 10 | 9 | 9 | P | P |  | 14 | 14 | 2 | 2 |  | - | 7 [GND] | CONN | CONN | CABLE | 7 PIN I/O
SWITCH | FW1_NO | YEL | N.O. | - | - | 11 | 12 | 12 | R | R |  | 15 | 15 | 3 | 3 |  | - | 1 [ENCODER A] |  |  | 7 CONN | CONN
LOWER FILTER | FW2_NC | VIO | N.O. | - | - | 12 | 11 | 11 | S | S |  | 16 | 16 | 4 | 4 |  |  |  | 1 [ENCODER A] | 1 [ENCODER A]
WHEEL | FW2_C | WHT | C. | - | - | 13 | 14 | 14 | T | T |  | 17 | 17 | 5 | 5 |  |  |  | 7 [GND] | 7 [GND]
SWITCH | FW2_NO | BLU | N.C. | - | - | 14 | 13 | 13 | U | U |  | 18 | 18 | 6 | 6
[6 empty] | 14PIN HDR | 14PIN HDR | 21 INCH
[6 empty] | COLD | COLD | COLD
[6 empty] | FEED- | FEED- | MANGANIN
[6 empty] | THROUGH | THROUGH | CABLE
PUPIL IMAGER | PIN_NC | VIO | N.C. | - | - | 1 | 2 | 2 | V | V |  | 19 | 19 | 7 | 7 | -
OUT (CW) | PIN_C | BLK | C. | - | - | 2 | 1 | 1 | W | W |  | 20 | 20 | 8 | 8 | - |  |  |  |  | 7 | 7 [GND]
SWITCH | PIN_NO | BLU | N.O. | - | - | 3 | 4 | 4 | X | X |  | 21 | 21 | 9 | 9 | - |  |  |  |  | 1 | 1 [ENCODER A]
PUPIL OPTICS | POUT_NC | GRN | N.C. | - | - | 12 | 11 | 11 | j | j |  | 22 | 22 | 10 | 10
IN (CCW) | POUT_C | WHT | C. | - | - | 13 | 14 | 14 | i | i |  | 23 | 23 | 11 | 11 | [5 empty] | 7 | 7 [GND]
SWITCH | POUT_NO | RED | N.O. | - | - | 14 | 13 | 13 | h | h |  | 24 | 24 | 12 | 12 | [5 empty] | 2 | 2 [ENCODER B]
CRYOSTAT ASSEMBLIES
[5 empty] | SERIAL MOTOR CONTROL |  |  |  | RJ45 | RJ45 |  | HR107-ND | HR209-ND |  |  |  | CD7WFW2000X | CD7WFW2000X | CD7WFW2000X | CD7WFW2000X | CD7WFW2000X | CD7WFW2000X
AMPS= | 10 | LENGTH= | 15 |  | FUNCTION |  |  |  |  | COMM |  | COMM | JBOX |  |  |  | UPPER FILTER | UPPER FILTER | LOWER FILTER | LOWER FILTER | PUPIL IMAGER | PUPIL IMAGER
CABLE | [5 empty] | NUMBER | LOCAL | LANTRONICS | LANTRONICS | 12 PIN | WIRE | 12 PIN | COMM |  |  |  | 26 PIN/SPLIT | WHEEL MOTOR | 26 PIN/SPLIT | WHEEL MOTOR | 26 PIN/SPLIT | MOTOR
GAUGE | ohms/ft | 48V drop |  |  |  | OR | SIGNAL | SERVER | SERVER | 15FT CABLE | COLOR/ | 15FT CABLE | 12 PIN |  |  |  | CABLE | MAIN | CABLE | MAIN | CABLE | MAIN
AWG14 | 0.00252 | 0.378 |  |  |  | NAME | NAME | PIN | CONN | CONN | AWG | CONN | CONN |  |  |  | MAIN CONN | CONN | MAIN CONN | CONN | MAIN CONN | CONN
AWG16 | 0.00402 | 4.02 |  |  | UPPER | FW1_TXD | TX+ (OUT) | 3 |  | 3 | ? | 1 | 1 | 13 | 13 |  | 4 [RXD] | 4 [RXD]
AWG18 | 0.00638 | 6.38 |  |  | FILTER WHEEL | FW1_RXD | RX+(IN) | 6 |  | 6 | ? | 2 | 2 | 14 | 14 |  | 3 [TXD] | 3 [TXD]
AWG20 | 0.0101 | 10.1 |  |  | SERIAL | GND | TX- | 4 |  | 4 | ? | 3 | 3 | 15 | 15 |  | 5 [GND RS232] | 5 [GND RS232]
AWG22 | 0.01614 | 16.14 |  |  | LINK | GND | RX- | 5 |  | 5 | ? | 4 | 4
[5 empty] | LOWER | FW2_TXD | TX+ (OUT) | 3 |  | 3 | ? | 5 | 5 | 16 | 16 |  |  |  | 4 [RXD] | 4 [RXD]
[5 empty] | FILTER WHEEL | FW2_RXD | RX+(IN) | 6 |  | 6 | ? | 6 | 6 | 17 | 17 |  |  |  | 3 [TXD] | 3 [TXD]
[5 empty] | SERIAL | GND | TX- | 4 |  | 4 | ? | 7 | 7 | 18 | 18 |  |  |  | 5 [GND RS232] | 5 [GND RS232]
[5 empty] | LINK | GND | RX- | 5 |  | 5 | ? | 8 | 8
[5 empty] | PUPIL IMAGER | FW2_TXD | TX+ (OUT) | 3 |  | 3 | ? | 9 | 9 | 19 | 19 | [5 empty] | 4 [RXD] | 4 [RXD]
[5 empty] | SERIAL | FW2_RXD | RX+(IN) | 6 |  | 6 | ? | 10 | 10 | 20 | 20 | [5 empty] | 3 [TXD] | 3 [TXD]
[5 empty] | LINK | GND |  | 4 |  | 4 | ? | 11 | 11 | 21 | 21 | [5 empty] | 5 [GND RS232] | 5 [GND RS232]
[6 empty] | GND | RX- | 5 |  | 5 | ? | 12 | 12
|  |  |  | MOTOR POWER SUPPLY |  | CONN: HR106-ND; CONN: HR108-ND; CONN: HR108-ND; CONN: HR106-ND
|  |  |  | FUNCTION |  | CONN: MTR PWR; CONN: MTR PWR; CONN: JBOX
[5 empty] | NUMBER | WIRE | ANIMATICS | ANIMATICS | MTR PWR | 10 PIN | WIRE | 10 PIN | MTR PWR
[5 empty] | OR | COLOR | POWER | POWER | 10 PIN | 15 FT CABLE | COLOR/ | 15 FT CABLE | 10 PIN
[5 empty] | NAME | CODE/AWG | SUPPLY | CONN | CONN | CONN | AWG | CONN | CONN
|  |  |  | MOTOR | 48VDC1 | RED | 1 |  | 1 | 1 | BLK | 1 | 1 | A | A |  | 1A [20-48VDC] | 1A [20-48VDC]
|  |  |  | POWER | GND1+2 | BLK | 2 |  | 2 | 2 | BRN | 2 | 2 | B | B |  | 2A [GND] | 2A [GND]
[5 empty] | 48VDC2 | ORG | 1 |  | 3 | 3 | RED | 3 | 3 | 23 | 23 |  |  |  | 1A [20-48VDC] | 1A [20-48VDC]
[5 empty] | GND1+2 | BLK | 2 |  | 4 | 4 | ORG | 4 | 4 | 24 | 24 |  |  |  | 2A [GND] | 2A [GND]
[9 empty] | 5 | 5 | YEL | 5 | 5 | 25 | 25 | [5 empty] | 1A [20-48VDC] | 1A [20-48VDC]
[9 empty] | 6 | 6 | GRN | 6 | 6 | 26 | 26 | [5 empty] | 2A [GND] | 2A [GND]
[10 empty] | N.C. | BRAID | SHIELD
[6 empty] | CONN: JBOX |  |  |  |  | UPPER FILTER |  | LOWER FILTER |  | PUPIL IMAGER
CONNECTOR TYPES | [17 empty] | WHEEL |  | WHEEL |  | MOTOR
COND | TYPE | CONT |  | PART NUM | [13 empty] | MOTOR |  | MOTOR
41 | PLUG | SOCKET |  | MS3116-F2041S
41 | RECT |  | HERMETIC | PT02-H20-41P
26 | PLUG | SOCKET |  | HR-145-ND
26 | RECT | PIN |  | HR152-ND
12 | PLUG | SOCKET |  | HR109-ND
12 | RECT | PIN |  | HR207-ND
12 | PLUG | PIN |  | HR107-ND
12 | RECT | SOCKET |  | HR209-ND
10 | PLUG | SOCKET |  | HR108-ND
10 | RECT | PIN |  | HR106-ND
C
N.C.
N.O.

## NOISEAMP
PREAMPLIFIER NOISE CALCULATION
Johnson Noise^2= 4kTR[BW]
k= 1.38E-13 J/K
T= Temp = 293K
BW= Bandwidth
[4kT]^1/2= 1.27E-1[R^1/2] nV/[Hz]^1/2
Rsource= | Nj [nV/[Hz}^1/2
1000 | 4.0161
For first stage opamp
Ntotal=[Njsourcev^2 + Ninputv^2 + (Ninputi * Requivalent)^2]^1/2
| Rsource | Njsource in nV/[Hz}^1/2
| 1000 | 4.0161
Ninputv= 3 nV/[Hz&1/2]
|  | 0.9 | for AD797
Ninputi= 0.4 pA/[Hz]^1/2
|  | 0.002 | for AD797
Requ = Rnon + Rinv¦¦Rfdbk
|  | Rnon input= | 50
|  | Rinv input=Rgain= | 1000
|  | Rfdbk= | 2000
|  | Requ= | 716.6667
Ntotal= | 4.3581 | in nV/Hz^1/2
| GAIN for 1st stage = 1+2(Rfdbk/Rcommon) |  |  | Compensation and BW limiting
|  | Rfdbk= | 2000 | Cfdbk= | 3.3e-11
|  | Rinv input=Rgain= | 1000 | Tc | 6.6e-08
|  | GAIN [1st stage] | 5 | 10 x Tc | 6.6e-07
|  |  |  | ~BW | 2411438.5317
| volts/[Hz]^1/2 | volts
N 1st stage= | 2.1791e-08 | 3.3838e-05
Second Stage
Ntotal=[Njsourcev^2 + Ninputv^2 + (Ninputi * Requivalent)^2]^1/2
| Rsource=Requ | Njsource in nV/[Hz}^1/2
| 300 | 2.1997
Ninputv= 3 nV/[Hz&1/2]
|  | 0.9 | for AD797
Ninputi= 0.4 pA/[Hz]^1/2
|  | 0.002 | for AD797
Requ = Rnon¦¦Roff + Rinv¦¦Rfdbk
|  | Rnon | 300
|  | Roff | 300
|  | Rinv | 300
|  | Rfdbk | 300
|  | Requ | 300
Ntotal= | 2.4513 | in nV/Hz^1/2
| GAIN for 2nd stage = 1+2(Rfdbk/Rinput) |  |  | Compensation and BW limiting
|  | Rfdbk | 300 | Cfdbk | 1.8e-10
|  | Rinput=Rinv | 300 | Tc | 5.4e-08
|  | GAIN [2nd stage] | 1 | 10 x Tc | 5.4e-07
|  |  |  | ~BW | 2947313.761
| volts/[Hz]^1/2 | volts
N 2nd stage= | 2.4513e-09 | 4.2083e-06
Offset stage
Ntotal=[Njsourcev^2 + Ninputv^2 + (Ninputi * Requivalent)^2]^1/2
| Requ | Njsource in nV/[Hz}^1/2
| 1177.6667 | 4.3583
Ninputv= 3 nV/[Hz&1/2]
|  | 0.9 | for AD797
Ninputi= 0.4 pA/[Hz]^1/2
|  | 0.002 | for AD797
Requ = Rnon + Rinv¦¦Rfdbk
|  | Rnon | 511
|  | Rinv | 2000
|  | Rfdbk | 1000
|  | Requ | 1177.6667
Ntotal= | 5.0351 | in nV/Hz^1/2
| GAIN for offset= -(Rfdbk/Rinv) |  |  | Compensation and BW limiting
|  | Rfdbk | 1000 | Cfdbk | 1e-06
|  | Rinv | 2000 | Tc | 0.001
|  | GAIN [offset] | 0.5 | 10 x Tc | 0.01
|  |  |  | ~BW | 159.1549
| volts/[Hz]^1/2 | volts
N offset= | 2.5175e-09 | 3.1761e-08
PREAMPLIFIER TOTAL
Ntotal = [ N1st(GAIN2nd)^2+N1st(GAIN2nd)^2 +Noff(GAIN2nd)^2 + N2nd^2]^1/2
volts/[Hz]^1/2
3.1016e-08
For Bandwidth | 2947313.761 | Total GAIN | 5
NOISE | 5.3248e-05 | volts
BESSEL POST AMPLIFIER FILTER
This is a noise calculation for the infrared preamplifier. It is presently based on the 256x256 InSb array preamp with its bandwidth opened up for faster pixel rates. Other opamps are being studied for the 1024x1024 application but no design decisions have been made yet.
The calculation is for voltage and current noise referenced to the instrumentation preamp input(s). The overall philosophy that was successfully used with 256x256 InSb, BIBs and CCDs was to limit the bandwidth of every stage of the preamplifier rather than just post filtering the output. This reduced the number of acrtive devices that were needed and helped the time domain response of the amplifier.
The array output equivalent resistance is assumed to be 1K ohms (a guess based on the 256x256 InSb array.
Voltage and current noise specs are for the AD797
The first stage bandwidth is held to a period that would let a small signal settle to 16 bits within 10 time constants. This time would be factored into the system timing. The large signal response is limited by the opamps slew response. This may be of concern in the raster reset/read mode. Typically the 1st stage has greater bandwidth than the 2nd stage.
The first stage noise number is given multiplied by its gain (in V/Hz^1/2) and in volts (over its bandwidth). It includes the source resistance.
The 10 time constants alloted here may need to be shortened, but typically the preamp is not seeing a step function from the array output (at least during Fowler readouts).
The Analogic ADCs have a 1lsb of 76uV for a plus and minus 2.5V input range. The NOISE number is the square root of sum the sum of squares of the noise contributions from each stage over the last stage bandwidth.
Present plans call for a Bessel or Gaussian filter to be used after the preamplifier. This will be located on the preamplifier board. The only filter located on the ADC board will be a passive RC filter at the input whose function is to provide for input charging of the capacitance of the ADC unbuffered input.

## Calibration Box
16 Pin Bendix
Pin # | Signal Name | Function | Color
A | V_QTH + | Quartz Tungsten Halogen Lamp +V | BLK
B | V_QTH - | Quartz Tungsten Halogen Lamp -V | BLK
C | V_INC + | 0.1 Watt Incandescent Lamp +V | ORG
D | V_INC - | 0.1 Watt Incandescent Lamp -V | ORG
E | V_BB + | Blackbody +V | WHT
F | V_BB - | Blackbody -V | WHT
G
H | V_ARG_L | Argon Discharge Tube (Oriel) 110VAC - L | BLK
J | V_ARG_N | Argon Discharge Tube (Oriel) 110VAC - N | WHT
K | V_ARG_G | Argon Discharge Tube (Oriel) 110VAC - GND | GRN
L
M
N
P
R
S

## insurance
Bigdog
10238 | Ultrasparc 2I Processor+accs
1300 | Odetics Timecode
700 | Disk
699 | Monitor/SONY 17PS
231 | keyboard & cables
3000 | VME backplane and power supply
16168
Guidedog
5427.25 | Themis SPARC 5/64-32-110
788 | Graphics TGX120,etc
200 | RAM
1300 | Odetics Timecode
699 | Monitor/SONY 17PS
700 | Disk
231 | keyboard & cables
3000 | VME backplane and power supply
12345.25
DSP boards
24165 | Ixthos DSP bd + memory
49581.5 | Ixthos/ qty 2 IXZ4444 DSP boards
2278 | Ixthos accessories
51859.5
Littledog
747 | Advantech/IPC-6010PC,PCA-6149
409 | Indus. Comp./ADIO 1600
285 | Chip merchant/16+8 MB 60ns
802 | PromoX 486CPU HD etc
2119 | Oregon Micro PC58-8S etc
2000 | chassis and power supply
800 | stepper motor dirve modules
7162
Temperature Control
2600 | Lake shore 330
2600 | Lake shore 331
2500 | Lakeshore 208
500 | filters
600 | chassis
8800
Cryostat mounted assys
3000 | Steppers
3200 | Servo Motors
80000 | Array Electronics
4000 | Hall Effect Preamp
90200
186534.75